## Slide 1: Qualitative and Quantitative Molecular QC
March 2016
Theresa Joy, Microbiology Supervisor
Julie West, PhD
Laboratory Technical Specialist

## Slide 2: Objectives
Refer to Molecular QC SOPs in the Microbiology Department at ATL VAMC
Review QC procedures/correct documentation for Molecular “New Shipments and New Lot Numbers”
Review proper frequency of Molecular QC testing

## Slide 3: VHA MISSION
Strategic Plan 2013-2018:
“Honor America’s Veterans by providing exceptional health care that improves their health and well-being.”

## Slide 4: Guidance
CLIA
CAP checklists (GEN, MOLECULAR, MICRO)
VAMC mandates; Chief of Pathology may add additional procedures as warranted
(Pathology holds all lab CLIA certificates)

## Slide 5: SOPs
Overarching SOPs:
QUALITY CONTROL and QUALITY IMPROVEMENT
MI-001-12
(contains definitions, frequency of QC/PMs, and expected results of molecular QC platforms)
LAB GEN (contains specific information on corrective actions – occurrence reporting, Check Point use, and other general laboratory)

## Slide 6: CHANGE
“The Only Thing That Is Constant Is Change -” ― Heraclitus (Greek philosopher, 500 B.C.)
NEW CAP Individualized QC PLAN (IQCP) requirements … refer to CAP letter.

## Slide 7: IQCP in Molecular… in our ATL VAMC Laboratory
Current procedures do not change
IQCP plans define the reasons for our current frequency of QC (because CLIA requires more frequent QC…)
Used for microbiology systems (Susceptibility, ID, rapid kits, and MOLECULAR -
Which molecular tests?

## Slide 8: Definitions

## Slide 9: IN THIS LAB, much of the QC required is dependent on the type of patient sample matrix and the test platform. The following is to be used as a basic guide as testing applies to Molecular platforms:
Molecular Tests – QUALITATIVE (AFFIRM and CEPHEID analyzers) (BD Affirm – multiplex test, MRSA-Cepheid, CDIF-Cepheid, Multiplex FLU-Cepheid) | 30-day external QC on lot in use Internal QC with each patient test QC each new lot/new shipment using external QC PRIOR to placing a lot in use (note: cannot use a surrogate patient for Affirm or MRSA; may use a surrogate patient specimen for CDIF); test one kit only (of the batch). Must also be performed after instrument software update or major maintenance.
Molecular Tests – QUALITATIVE (ROCHE CT/NG platform: 4800) | External QC is always on board with each run. No need for 30-day QC. QC each new lot/new shipment of detection reagents with surrogate patient specimens analyzed using a prior lot number (positive for CT/NG, negative for CT/NG). Internal QC is automatically on board with each patient test
Molecular Tests – QUANTITATIVE (ROCHE HIV PCR viral load) | 30-day external SURROGATE QC on lot in use QC each New lot/new shipment using external (saved, frozen) surrogate QC and a “TND” surrogate tested previously with the last reagent lot in use (test one run of one kit only of the lot/shipment) Internal QC automatic with each patient test External QC on board with each PCR run AMR/LINEARITY VERIFICATION every 6 months (using CAP linearity kit)

## Slide 10: New shipments – What to do?
Sign for shipment AFTER checking to be sure it is for our department
Sign-in product name (actual number of items received), vendor name, PO#, date/your initials
Go directly to QC CLIPBOARD to log-in the items that require QC. HIV and CT/NG items are logged in on forms in the respective sections.
IMPORTANT!! Pull the pack slips for supervisor review. (Need these for invoicing the purchase orders.)

## Slide 11: Linearity / CAL-ver
6 month Calibration Verification (linearities) on all quantitative systems:
HIV PCR
EIA plate reader
Quantiferon plate reader
How? Use HIV SOP and CAP survey instructions. Be sure to go to I-drive, CAP folder to enter linearity into the designated Excel chart – we want to be sure we are linear (no clerical errors) before sending results to CAP.

## Slide 12: Pause to Review
What is a surrogate control?
What does MULTIPLEX mean?
What is difference between Qualitative molecular testing and Quantitative molecular testing?

## Slide 13: Examples of Occurrences
Cepheid 30 day QC: No QC report printouts from instrument in QC book.
Wrong lot of Affirm kit in use; must match the QC form for “lot in use”
CT/NG: Remember to document the “lot in use” on the QC form
HIV/PCR: Failure to run all external and surrogate (Acrometrix) controls after an instrument repair, or during the 30 days.

## Slide 14: QUALITATIVE MOLECULAR
What do we look for when reviewing QC?
Cepheid (FLU, CDIF, MRSA)
Affirm (Vaginitis panel)
Roche 4800 (CT/NG)
Pull out the QC books – let’s review!

## Slide 15: Quantitative Molecular
Roche Ampliprep/Taqman
HIV PCR (v. 2)
(maybe more tests coming – HCV?)
Pull out the QC books – let’s review!

## Slide 16: HIV QC - During the month: (Review logs and calculations)
Wash Reagent Buffer log
Daily Run QC
Cross-check new lot, new shipment, post-repairs
New lot HIV KIT log
Other logs:
Running SD
New Lot Calculations
System software upgrade cross-check
Monthly Review of HIV QC cover sheet

## Slide 17: HIV: End of month QC paperwork…from the SOP:
Plot points, and Print Lot QC detail charts for CTM(-C) v2.0, LPC(L+C) v2.0, and HPC(H+C) v2.0 controls.
Calculate manufacturer means for Low and High Control (LPC, HPC) values, new kit lot numbers. According to Dr. Heaney (Roche Diagnostics, 2010a):
Molecular results are not always linear. Westgard Rules do not apply when calculating folded manufacturer means.
The ranges for the controls are determined around the mean as follows:
LPC: -4/+3 fold of mean, so this means the lower end of the range is 4
fold less than the mean and the upper end of the range is 3 fold higher
than the mean
HPC: -3/+3 fold of mean, so this means the lower end of the range is 3
fold less than the mean and the upper end of the range is 3 fold higher
than the mean

## Slide 18: Continued…
Plot LPC and HPC log values on corresponding QC LJ charts for each reagent lot number (using Excel or other chart software method).
After several log values are obtained (CLSI, 2003), calculate lab mean of the log value. Calculate the 1SD, 2SD, and 3SD for LPC and HPC values…
…using a running SD (obtained from at least three different shipments/lot numbers over time for each control SD). Recalculate when necessary, adding new SD data to the running SD calculations. Plot the resulting values. (Once established, no need to continue to calculate running SDs between new lots, unless you think your range is too wide or too narrow.) More on this later…

## Slide 19: Continued…
Review charts for signs of shifts or trends. Use 3s and 10x Westgard rules for guidance. Roche (2010b) recommends accepting all QC that falls within the lab’s own 3SD pre-determined range.
A shift or trend may be occurring if more than 10 consecutive points are on one side of the lab’s mean.

## Slide 20: Continued…
Compare lab mean with manufacturer mean for LPC and HPC. Variation is allowed if it is within 0.3 log (CLSI, 2003).
Verify that all the values are within the manufacturer’s (calculated) folded ranges.
Monthly review is to be performed by supervisor or designee. Take corrective action by notifying Roche technical support and document findings, if warranted. (Note: When changing lots, lock-in the previous parameters and then change SD after several new points have been established.)

## Slide 21: “Running SD” and Establishing historical CV (CVh) … FYI (for your information)…
CVh is the historical instrument precision for that method measured as a percent (%CV = SD/Mean X 100). It can be used to establish Standard Deviation (SD), QC ranges and to monitor your instrument precision.
The CVh is best established over several lots (3 or more) and can either be the average CV or more preferable, the highest acceptable CV over that period. CVs obtained during times of instrument malfunction or significantly higher than the average CV should not be used in calculations. If you have no previous CV data, then use the current CV as the CVh basis and modify this value as you move on to subsequent new lots. This basis value will need to be increased slightly (~1-2%) in order to establish QC ranges that are not too narrow.
Michael, K. (2006). Chemistry - Guideline for Establishing QC Ranges and use of CV.doc

## Slide 22: Surrogates…
External surrogate controls are run for each new lot number or shipment of test materials; after major system maintenance; and after software upgrades; or, every 30 days, whichever is more frequent.
For new reagent lot verification: New lots and/or shipments are validated before or concurrent with use for patient testing. For Ampliprep/Taqman HIV PCR, cross-checking of lots consists of testing three levels of known surrogate controls: a) weak positive, b) mid-to-high range positive, and c) log <1.30 (target not detected) patient obtained from a previous run (using previous lot number of reagent if performing a cross-check).
For external control testing after maintenance, software upgrades, or every 30 days: Perform testing with external QC reagents included within the kit currently in-use, and add two levels of stock surrogate controls.

## Slide 23: Continued…surrogates
The stock surrogate controls (weak positive and mid-to-high range positive) may be obtained from a third party vendor, or may be made in-house using neat plasma obtained from a known positive patient using the following procedure (see SOP…)
Record results of surrogate external QC on QC forms. Include kit lot number. Report in log values.
Plot values on QC LJ charts, and review for shifts and trends between kits.
Compare the lab mean with expected mean. Inter-assay variation should remain within 0.2 log of the calculated surrogate mean. Verify the values are within 3SD of the surrogate calculated laboratory mean. Check for violation of 10x rule.(CLSI, 2003).

## Slide 24: Continued… surrogates
Quality control statistical charts are created monthly to monitor trends over time and define analytic imprecision. Monthly review of charts to be performed by supervisor or designee. Print and review charts (all levels of external surrogate controls, both manufacturer and patient controls) for any shifts or trends over time or between reagent lot numbers. Also take care to review charts for changes after system or software maintenance. Take corrective action (and document) if warranted.
What is a shift? A shift is when the QC values move suddenly upward or downward from the mean and continue the same way mathematically changing the mean. This is usually seen when new reagents or quality control material has been used, when there is an issue with the instrument such as a change in internal temperature …
What is a trend? A trend is when the QC values slowly move up or down from the mean and continue moving the same direction over time. This can be the case of a calibration that is failing or reagent stability issues. For Taqman: Usually, the bulb is at end of its life.

## Slide 25: Time to look at the LJ Charts…
Pull out the HIV QC notebook – let’s review!
Examples

## Slide 26: CAP surveys – and QC
CAP survey participation is considered a part of the laboratory’s QC/QA program. Review CAP survey results and placement with regard to peer group participation. Take corrective action if lab results are out of range (flagged) when compared to peer group.
Again, look for shifts or trends.
Pull out the CAP HIV survey notebook – let’s review!

## Slide 27: Open Discussion
Any QC Questions?
*Report all concerns, suggestions, occurrences to the microbiology supervisor.
EVERYONE is responsible for QC! Look at each item with an INSPECTOR’s eye…

## Slide 28
Thank you