Qualigen Prostate Specific Antigen (PSA) Competency Assessment

- Specimens for PSA must sit until _________ before you can centrifuge the sample for testing.

- You run quality control (QC) for level 1 and level 2.  The value for level 1 is 2.6 (QC Range is 1.2 – 4.0).  The value for level 2 is 19.2 (QC range is 5-15).  What would be your next step?

- Nothing, patients can be tested

- Repeat QC level

- Recalibrate, repeat both QC

- Repeat both levels of QC

- True or False - Using the sample dispenser is optional; you can load the fast pack manually.

- What is the upper limit of linearity for the Qualigen PSA? ____________

What is the lower limit of linearity for the Qualigen PSA? ____________

- True or False – It is acceptable to prop open the sample chamber with a pipette tip so that you can manually dispense the sample.

- You must calibrate the Qualigen PSA when:  (circle all that apply)

- You come in Monday to find the refrigerator with your PSA reagents at 15 degrees Celcius, normal range is 2-8.

- You have a new lot of fast pack PSA

- It has been 32 days since last calibration

- After the analyzer has had service performed

- The pipette on the Qualigen must be replaced every ___________

- Before you run the quality control for PSA you must first check _________________

- What must you do after the last sample has been run for the day? (circle all that apply)

- Clean analyzer

- Take a break

- Check off maintenance log

- What is the volume of sample that is used for the PSA test?

- 50 microliters

- 50 deciliters

- 100 microliters

- 100 kiloliters