| Flag | Probable sample cause |  | Action |
| --- | --- | --- | --- |
| WL | Incomplete lysing of red blood cells, presence of nucleated red blood cells, increase of large platelets, platelet aggregation or agglutination, precipitation of fibrin, presence of proteins or lipids. |  | Check for clots. Clot present, do not report, redraw. No clot, Re-analyze. Flag still present, send to Main Concord Hospital Lab. |
| RL | Presence of fragmented red blood cells, increase of large platelets, platelet aggregation or agglutination, presence of micro-erythrocytes. |  | Check for clots. Clot present, do not report, redraw. No clot, Re-anlyze. Flag still present, send to Main Concord Hospital Lab. |
| PL | Effects of cryoglobulins, fragmented red blood cells, or cellular fragments of white blood cells. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| WU | Incomplete lysing of red blood cells, presence of immature white blood cells, white blood cell aggregation, platelet satellite phenomenon, etc. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| RU | Effects of cold agglutinin, inclusion of white blood cells. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| DW (RBC) | Significant anisocytosis, etc |  | Send to Main Concord Hospital Lab. |
| DW (PLT) | Inclusion of fragmented red blod cells, non-uniformity in size of platelets, effects of cryoglobulins, etc |  | Send to Main Concord Hospital Lab. |
| MP (RBC) | Effects of anemia treatment or blood transfusion causing the presence of cells of multiple sizes. |  | Send to Main Concord Hospital Lab. |
| MP (PLT) | Platelet aggregation, low platelet count. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| T1 | Presence of CML or other immature granulocytes, incomplete lysing of red blood cells, etc., causing the first two WBC populations in the WBC-Histogram not to be separated. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| T2 | Presence of CML or other immature granulocytes, aged sample, incomplete lysing of red blod cells, etc., causing the last two WBC populations in the WBC-Histogram not to be separated. |  | Re-analyze. Flag still present, do not repot, send to Main Concord Hospital Lab. |
| F1, F2, F3 | Presence of CML or other immature granulocytes, sample with high values for monocytes, eosinophils, and basophils, incomplete lysing of red blood cells, aged sample, etc. |  | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |
| AG |  | Presence of nucleated red blood cells, increase of large platelets, platelet aggregation or agglutination, precipitation of fibrin, presence of proteins or lipids, etc. | Re-analyze. Flag still present, do not report, send to Main Concord Hospital Lab. |

| Parameter | Unit | Repeat Criteria | Critical Range | Reportable Range |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| WBC | x103/uL | <2.0 >40.0 | <2.0 >40.0 | 1.1 – 90.00 |  |  |  |
| RBC | x106/uL | NA | NA | 1.00 – 6.60 |  |  |  |
| HGB | g/dL | NA | <7 >19 | 3.0 – 20.0 |  |  |  |
| HCT | % | <20 >54 | <20 >58 | 10.0 - 60.0 |  |  |  |
| PLT | x103/uL | <50 >700 | <50 >850 | 10 - 850 |  |  |  |
| Reportable Ranges have changed ** Rerun any specimen that meets or exceeds the REPEAT CRITERIA. |  |  |  |  |  |  |  |
| ** Indicate on the printout and in the nursing notes that the specimen was repeated (RPTD). |  |  |  |  |  |  |  |
| (!) Indicates the result exceeds the lower or upper limit of the reportable range. |  |  |  |  |  |  |  |
| * | Send the specimen to the Hospital Lab. |  |  |  |  |  |  |

pocH-100i Criteria / Histogram Error Flags / Action to be Taken