## Slide 1: BinaxNOW™
COVID-19 Ag Card
Overview
Emergency Use Authorization Granted by FDA

## Slide 2: Intended Use: Key Points
Proprietary and confidential — do not distribute
The BinaxNOW™ COVID-19 Ag Card is intended for the qualitative detection of nucleocapsid protein antigen from SARS-CoV-2 in nasal swabs from individuals suspected of COVID-19 by their healthcare provider within the first seven days of symptom onset.
Antigen is generally detectable in nasal swabs during the acute phase of infection.
Negative results from patients with symptom onset beyond seven days, should be treated as presumptive and confirmation with a molecular assay if necessary, for patient management, may be performed.

## Slide 3: Reagent and Materials
Proprietary and confidential — do not distribute
Materials Provided:
40 Test Cards
Extraction Reagent
Patient Collection Nasal Swabs
Positive Control Swab
Blank Nasal Swab for Negative Control
Product Insert
Procedure Card
Healthcare Provider & Patient COVID-19 Fact Sheets
Materials Required but not Provided:
Clock, timer or stopwatch
Optional Mate
Plastic Transport Tube
Storage & Stability:
Store kit at 2-30°C
Ensure all test components are at room temperature before use
Stable until the expiration date marked on the outer packaging

## Slide 4: Quality Control
Proprietary and confidential — do not distribute
Internal Procedural Controls:
BinaxNOW™ COVID-19 Ag Card has built-in procedural controls
In an untested card there will be a blue line at the Control Line position
In a valid, tested device, the blue line washes away and a pink/purple line appears, confirming that the sample has flowed through the test strip and the reagents are working
Note: If the blue line is not present at the Control Line position prior to running the test, do not use and discard

## Slide 5: When is Quality Control Required?
Proprietary and confidential — do not distribute
External Positive & Negative Controls:
CH requires the use of positive and negative controls to ensure that test reagents are working properly and that the test is correctly performed
BinaxNOW™ COVID-19 Ag Card kits contain a positive control swab and sterile swabs that can be used as a negative control
Required Frequency:
New shipments received
New Lot received
Untrained operators

## Slide 6: Nasal Swab Sample Collection
Insert the nasal swab into the nostril exhibiting the most drainage or congestion
Using gentle rotation, push the swab until resistance is met
  At the level of the nasal turbinates
  Less than one inch into nostril
Rotate the swab 5 times or more against the nasal wall
Slowly remove the swab
Using the same swab, repeat sample collection in the other nostril
Only the swab provided in the kit is to be used for nasal swab collection

## Slide 7: Specimen Transport & Storage
For best performance, direct nasal swabs should be tested as soon as possible after collection.
If immediate testing is not possible:
Do not return the nasal swab to the original paper packaging
To avoid contamination and preserve sample integrity, place the nasal swab in a clean, unused, tightly capped plastic tube & label with the patient information
The sample is stable in the plastic tube at room temperature (15-30°C) for up to one (1) hour prior to testing
If greater than one (1) hour delay occurs, dispose of sample & collect new sample

## Slide 8: BinaxNOW™ COVID-19 Ag Card Test Procedure Overview
Proprietary and confidential — do not distribute

## Slide 9: BinaxNOW™ COVID-19 Ag Card Overview
Exterior View
Interior View

## Slide 10: BinaxNOW™ COVID-19 Ag Card Procedure
Proprietary and confidential — do not distribute
Key Points
Prior to Beginning a Test Remember to:
Ensure all components are at room temperature
Open test card j ust prior to use
Lay test card flat for use
Ensure Blue Control Line is present

## Slide 11: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 1:
Hold extraction reagent bottle vertically to ensure adequate volume
Add 6 drops for a Patient test
Add 8 drops for a Quality Control test

## Slide 12: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 2:
Insert swab into the bottom hole of the test card
Firmly push swab upwards until the swab tip is visible in the top hole

## Slide 13: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 3:
Rotate swab shaft 3 times CLOCKWISE (to the right).
False negative results can occur if the sample swab is not rotated (twirled) prior to closing the card

## Slide 14: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 4:
Peel off adhesive liner from the right edge of the test card
Close and securely seal the card
To ensure proper test performance read results at 15 minutes and not before
Results should not be read after 30 minutes

## Slide 15: Result Interpretation
Proprietary and confidential — do not distribute
Negative
One pink/purple colored line in the top half of the window, in the Control Line position
Indicates a negative result & no antigen detected
Positive
Two pink/purple colored lines in both the Control & Sample Line position
Indicates COVID-19 antigen was detected
Any visible pink/purple line is positive

## Slide 16: Result Interpretation
Proprietary and confidential — do not distribute
Invalid
The assay is invalid & should be repeated if:
Only the Sample Line is present
No lines are present
Blue Control Line remains blue
Test should be discarded as Biohazard waste

## Slide 17: Patient Workflow
Proprietary and confidential — do not distribute
Step 1: Add the Extraction Reagent
Step 2: Insert sample into bottom hole and push firmly upwards so tip is visible in top hole
Step 3: Rotate (twirl) swab shaft 3 times clockwise (to the right)
Step 4: Peel off adhesive liner, close and securely seal card. Read results in the window 15 minutes after closing the card
Collect Nasal Swab
Collect Nasal Swab, Place in labeled transport tube, transport to testing area with in 1 hour of collection