Intended use

The ARCHITECT MYOGLOBIN assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of  myoglobin  in human serum and  plasma on the ARCHITECT i System. Myoglobin values are used to assist in the diagnosis of myocardial infarction (MI).

Summary

Myoglobin is a tightly folded, globular heme-protein located in the cytoplasm of both skeletal and cardiac muscle cells. Its function is to store and supply oxygen to muscle cells. The molecular weight of myoglobin is approximately 17,800 daltons.1,2 The relatively low molecular weight and the location of storage accounts for the rapid release from damaged muscle cells and earlier rises in concentration measured above baseline in blood as compared to other cardiac markers.1,3,4

In ischemic heart disease, such as myocardial infarction (MI), a temporal pattern of increased release of myoglobin into the blood stream is observed. The serum or plasma myoglobin level will start to show

an increase between 2-4 hours after an MI has occurred, peaking at approximately 8-10 hours, and returning to baseline after 24 hours. Measurement of myoglobin between 2-12 hours after an MI can be a

good adjunct to electrocardiography (ECG) in improving the efficiency of early diagnosis of MI.1,5,6 Monitoring the myoglobin levels can also help in evaluating the success of thrombolytic therapy.6,7

Since myoglobin is present in both cardiac and skeletal muscle, any damage to either of these muscle types results in its release into the blood stream. Elevated serum levels of myoglobin have been observed

under the following conditions: skeletal muscle damage, skeletal muscle or neuromuscular disorders, cardiac bypass surgery, renal failure, strenuous exercise.2,5 Therefore, serum myoglobin levels should be used in conjunction with other aspects of the patient assessment to aid in the diagnosis of an MI. Myoglobin may also rise moderately above the reference range in chronic ischemic heart disease (i.e. unstable angina).2 For diagnostic purposes, the ARCHITECT STAT Myoglobin assay results

should be used in conjunction with other data; e.g., other clinical testing, ECG, symptoms, clinical observations.

## Method

The ARCHITECT STAT Myoglobin assay is a two-step immunoassay to determine the presence of myoglobin in human serum and plasma using CMIA technology with flexible assay protocols, referred to as Chemiflex.

## Principle

In the first step, sample and anti-myoglobin coated paramagnetic microparticles are combined and incubated. Myoglobin present in the sample binds to the anti-myoglobin coated microparticles. After washing, anti-myoglobin acridinium-labeled conjugate is added in the second step. Following another incubation and wash, pre-trigger and trigger solutions are added to the reaction mixture. The resulting chemiluminescent reaction is measured as relative light units (RLUs). A direct relationship exists between the amount of myoglobin in the sample and the RLUs detected by the ARCHITECT i* System optics.

For additional information on system and assay technology, refer to the ARCHITECT System Operations Manual, Section 3.

## Specimen collection and handling

1.  The following specimen collection tubes may be used for the ARCHITECT STAT Troponin-I assay.

Serum:	Glass: 	Will not be used

Plastic:  Serum separator tubes with clot activators

Plasma

Glass:	 Will not be used

Plastic:  Lithium Heparin

Plasma separator tubes with lithium heparin

Sodium Heparin

Other anticoagulants have not been validated for use with the ARCHITECT STAT Myoglobin assay.

2.  Refer to the specimen collection tube manufacturer's instructions as well as these package insert instructions for specimen collection and preparation for analysis.  Each laboratory should determine the acceptability of its own specimen collection and processing procedures.

3.  Inadequate centrifugation of the specimen may cause an erroneous result.

4.  For serum specimens, ensure that complete clot formation has taken place prior to centrifugation.  Some specimens, especially those from patients receiving anticoagulant or thrombolytic therapy, may exhibit increased clotting times.  If the specimen is centrifuged before complete clot formation, the presence of fibrin may cause erroneous results.

5.  If a lipid layer forms on the specimen surface, avoid the lipid layer when withdrawing the specimen.

6.  If testing will be delayed more than 8 hours, remove the plasma or serum from the cells, clot, or gel.  Specimens removed from the cells, clot, or gel may be stored up to 72 hours at 2-8 degrees C. or stored frozen (-10 degrees C. or colder) prior to being tested.

7.  Specimens can be stored up to 30 days frozen at -10 degrees C. or colder.

8.  Thaw frozen specimens and mix thoroughly by LOW speed vortexing or by gently inverting, then centrifuge at 2,500-3,000 x g for 10 minutes prior to use to remove particulate matter and to ensure consistency in the results.  Thaw specimens only once.

9.  Ensure specimens are free of fibrin, red blood cells, and other particulate matter.

10.  Inspect all samples for bubbles.  Remove bubbles prior to analysis.

11.  When serial specimens are being evaluated, use the same specimen type throughout the evaluation.

12.  The ARCHITECT i System does not provide the capability to verify specimen type.  It is the responsibility of the operator to verify the correct specimen type is used in the ARCHITECT STAT Myoglobin assay.

13.  Use caution when handling patient specimens to prevent cross contamination.  Use of disposable pipettes or pipette tips is recommended.

14.  Test all samples (patient specimens, controls, and calibrators) within 3 hours of being placed on board the ARCHITECT i System.  Refer to the ARCHITECT System Operations Manual, Section 5, for a more detailed discussion of on-board sample storage constraints.

15.  Performance has not been established using cadaver specimens or body fluids other than human plasma or serum.

16.  Do not use samples with obvious microbial contamination.

17.  Do not use heat-inactivated specimens.

18.  When shipping specimens, package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.  Prior to shipping, remove the plasma or serum specimen from the cells, clot, or gel.  Ship specimens frozen on dry ice.

Materials and Equipment Required

Provided

- 2K43 ARCHITECT STAT Myoglobin Reagent Kit

Materials Required but not Provided

- ARCHITECT i System

3K51 ARCHITECT i	-US- Addition B

3k53 ARCHITECT i	-WW (excluding US)- Addition B

- 2K43-01 ARCHITECT STAT Myoglobin Calibrators

- 2K43-10 ARCHITECT STAT Myoglobin Controls

ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- Pipettes or pipette tips (optional) to deliver the volumes spedified on the patient or control order screen.

For information on materials required for maintenance procedures, refer to the ARCHITECT System Operations Manual, Section 9.

## Reagents – working solutions

Reagent Kit, 100 Tests/500 Tests

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT i Systems. Please contact your local distributor.

ARCHITECT STAT MYOGLOBIN Reagent Kit (2K43)

1.  Microparticles:  1 or 4 Bottles (6.6 mL/27.0 mL) Anti-myoglobin (mouse, monoclonal) coated microparticles in TRIS buffer with protein (bovine ) stabilizers.  Preservative:  antimicrobial agents.

2.  Conjugate:  1 or 4 Bottles (5.9 mL/26.3 mL) Anti-Myoglobin (mouse, monoclonal) acridinium labeled conjugate in MES buffer with protein (bovine) stabilizer.  Preservative:  ProClin 300.

3.  Assay Diluent:  1 or 4 Bottles (10.0 mL/50.9 mL) Myoglobin Specimen Diluent, containing protein (bovine ) stabilizer in TRIS buffer.  Preservative:  Sodium Azide.

Other Reagents

ARCHITECT i Pre-Trigger Solution

Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT i Trigger Solution

Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT i Wash Buffer

Wash Buffer containing phosphate buffered saline solution. Preservative: antimicrobial agents.

## Storage and stability

| 1. The ARCHITECT STAT Myoglobin Reagent Kit must be stored at 2-8 degrees C. in an upright position and may be used immediately after removal from 2-8 degrees C. storage. 2. When stored and handled as directed, reagents are stable until the expiration date. 3. The ARCHITECT STAT Myoglobin Reagent Kit may be stored on-board the ARCHITECT i System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking on-board time, refer to the ARCHITECT System Operations Manual, Section 5. 4. Reagents may be stored on or off the ARCHITECT i System. If reagents are removed from the system, store them at 2-8 degrees C. (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If the microparticle bottle does not remain upright (with a septum installed) while in 2-8 degrees C. storage off the system, the reagent kit must be discarded. After reagents are removed from the system, one must initiate a scan to update the on-board stability timer. INDICATIONS OF REAGENT DETERIORATION When a control value is out of the specified range, it may indicate deterioration of the reagents or errors in technique. Associated test results may be invalid and may require retesting. Assay recalibration may be necessary. For troubleshooting information, refer to the ARCHiTECT System Operations Manual, Section 10. |  |
| --- | --- |

## Calibration

1.  To perform an ARCHITECT STAT Myoglobin calibration, test the Calibrators A, B, C, D, E, and F in duplicate.  A single sample of each ARCHITECT STAT Myoglobin Control level must be tested to evaluate the assay calibration.  Ensure that assay control values are within the concentration ranges specified in the control package insert.  Calibrators should be priority loaded.

2.  Calibration range:  0.00 – 1200.00 ng/mL.

3.  Once an ARCHITECT STAT Myoglobin calibration is accepted and stored, all subsequent samples may be tested without further calibration unless:

a.  A reagent kit with a new lot number is used

b.  Controls are out of range

4.  For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Quality control

The recommended control requirement for the ARCHITECT STAT Myoglobin assay is that a single sample of each control level be tested once every 24 hours each day of use. If the quality control procedures in your laboratory require more frequent use of controls to verify test results, follow your laboratory-specific procedures.

The ARCHITECT STAT Myoglobin control values must be within the acceptable ranges specified in the control package insert. If a control is out of the specified range, the associated test results are invalid and samples must be retested. Recalibration may be indicated.

For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Preparation of Working Solutions

Handling Precautions

•   Do not use reagent kits beyond the expiration date.

•   Do not pool reagents within a kit or between reagent lots.

•  Prior to loading the ARCHITECT STAT Myoglobin Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this package insert.

•   Septums MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

•   To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

•   Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

- Over time, residual liquids may dry on the septum surface. These are typically dried salts, which have no effect on assay efficacy.

- For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay Procedure

INSTRUMENT PROCEDURE

1.  The ARCHITECT STAT Myoglobin assay file must be installed on the ARCHITECT i System with STAT protocol capability from the ARCHITECT i Assay CD-ROM Addition B prior to performing the assay.  For detailed information on assay file installation and on viewing and editing assay parameters, refer to the ARCHITECT System Operations Manual, Section 2.

2.  For information of printing assay parameters, refer to the ARCHITECT System Operations Manual, Section 5.

3.  For a detailed description of system procedures, refer to the ARCHITECT System Operations Manual.

4.  The default result unit for the ARCHITECT STAT Myoglobin assay is ng/mL.  An alternate result unit, ug/L, may be selected for reporting results by editing assay parameter "Result concentration units", to ug/L.  The conversion factor used by the system is 1.0.

ASSAY PROCEDURE

1.  Before loading the ARCHITECT STAT Myoglobin Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment:

a.  Invert the microparticle bottle 30 times.

b.  Visually inspect the bottle to ensure microparticles are resuspended.  If microparticles are still adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

c.  If the microparticles do not resuspend, DO NOT USE.  Contact your local Abbott Laboratories representative.

d.  Once the microparticles have been resuspended, remove and discard the cap.  Wearing clean gloves, remove a septum from the bag.  Squeeze the septum in half to confirm that the slits are open.  Carefully snap the septum onto the top of the bottle.

2.  Order calibration, if necessary

a.  For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

3.  Order tests.

a.  For information on ordering patient specimens and controls, refer to the ARCHITECT System Operations Manual, Section 5.

4.  Load the ARCHITECT STAT Myoglobin Reagent Kit on the ARCHITECT i System with STAT protocol capability.

a.  Verify that all necessary assay reagents are present.  Ensure that septa are present on all reagent bottles.

5.  The minimum sample volume is calculated by the system and is printed on the Orderlist report.  No more than 10 replicates may be sampled from the same sample cup.  To minimize the effects of evaporation, verify adequate sample cup volume is present prior to running the test.

a.  Priority:  80 uL for the first ARCHITECT STAT Myoglobin test plus 30 uL for each additional ARCHITECT STAT Troponin-I test from the same sample cup.

b.  </= 3 hours on board:  150 uL for the first ARCHITECT STAT Myoglobin test plus 30 uL for each additional ARCHITECT STAT Troponin-I test from the same sample cup.

c.  If using primary or aliquot tubes, use the sample gauge to ensure sufficient patient specimen is present.

6.  Prepare calibrators and controls.

a.  ARCHITECT STAT Myoglobin Calibrators and Controls should be mixed according to instructions in their respective package inserts.

b.  To obtain the recommended volume requirements for the ARCHITECT STAT Myoglobin Calibrators, hold the bottles vertically and dispense 8 drops of each calibrator into each respective sample cup.  Dispense 150 uL of each control into each respective sample cup.

7.  Load samples.

a.  For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

8.  Press RUN.  The system performs the following functions:

a.  Moves the sample to the aspiration point.

b.  Loads a reaction vessel (RV) into the process path.

c.  Aspirates and transfers sample into the RV.

d.  Advances the RV one position and transfers assay diluent and microparticles into the RV.

e.  Mixes, incubates, and washes the reaction mixture.

f.  Adds conjugate to the RV.

g.  Mixes, incubates, and washes the reaction mixture.

h.  Adds pre-trigger and trigger solutions.

i.  Measures chemiluminescent emission to determine the quantity of troponin-I in the sample.

j.  Aspirates contents of RV to liquid waste and unloads RV to solid waste.

k.  Calculates the result.

9.  For optimal performance, it is important to follow the routine maintenance procedures defined in the ARCHITECT System Operations Manual, Section 9.  If your laboratory requires more frequent maintenance, follow those procedures.

#### Interpretation: reporting results

Calculation

- The ARCHITECT STAT Myoglobin assay utilizes a 4 Parameter Logistic Curve Fit (4PLC,Y weighted) data reduction method to generate a calibration curve.

Flags

- Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

- Normal Range  Female 0.0 – 106.0 ng/mL

- Normal Range Male  0.0 – 155.0  ng/mL

Critical Values:  N/A

Any condition resulting in skeletal or cardiac muscle damage can potentially increase myoglobin levels.  Published studies have documented that these conditions include, but are not limited to, myocardial infarction (MI), unstable angina, skeletal muscle damage, skeletal or neuromuscular disorders, cardiac bypass surgery, renal failure, strenuous exercise.

Serial sampling is recommended to detect the temporal rise and fall of myoglobin levels characteristic of MI. Myoglobon levels should also be used in conjunction with other aspects of the patient assessment (ECG, symptoms, clinical observations) to aid in the diagnosis of an MI.

For diagnostic cutoff and additional information, refer to the SPECIFIC PERFORMANCE CHARACTERISTICS, Clinical Performance section in this package insert.

A reference range study was conducted based on guidance from National Committee for Clinical Laboratory Standards (NCCLS) Protocol C28-A2.  Apparently healthy individuals were evaluated in replicates of one using the ARCHITECT STAT Myoglobin assay.  The observed values are summarized in the following table.*

APPARENTLY HEALTHY POPULATION

Population:  Female

n:  160

Age Range:  18-62

99th Percentile (ng/mL):  106.0

Population:  Male

n:  159

Age Range:  18-63

99th Percentile (ng/mL):  154.9

Population:  TOTAL

n:  319

Age Range:  18-63

99th Percentile:  140.1

*Representative data; results in individual laboratories may vary from these data.

ALTERNATE RESULT UNITS

1.  The default result unit for the ARCHITECT STAT Myoglobin assay is ng/mL.  When the alternate result unit, ug/L, is selected, the conversion factor used by the system is 1.0.

2.  Conversion Formula:  (Concentration in ng/mL) x (1.0) = ug/L.

## Measuring Range:

- The measurement range for the ARCHITECT STAT Myoglobin assay is 0 ng/mL to 1200 ng/mL.

Dilutions

Specimens with an myoglobin value exceeding 1200.0 ng/mL are flagged with the code "> 1200.00" and may be diluted with the Automated Dilution Protocol or the Manual Dilution Procedure.

AUTOMATED DILUTION PROTOCOL

1.  If using the Automated Dilution Protocol, the system performs a 1:10 dilution of the specimen and automatically calculates the concentration of the specimen before dilution and reports the result.

2.  Specimens with a myoglobin value exceeding 12000.0 ng/mL are flagged with the code "> 12000.0" when run using the Automated Dilution Protocol.  These specimens may be diluted by following the Manual Dilution Procedure.

MANUAL DILUTION PROCEDURE

1.  Manual dilutions should be performed as follows:

a.  The suggested dilution for a myoglobin test is 1:20.

b.  Prior to diluting the specimen, dispense 10 drops of Myoglobin Calibrator A into a clean test tube for use in the next step.

c.  Add 10 uL of the patient specimen to 190 uL of ARCHITECT STAT Myoglobin Calibrator A.

d.  The operator must enter the dilution factor in the Patient or Control order screen.  The system will use this dilution factor to automatically calculate the concentration of the sample before dilution.  This will be the reported result.  The dilution should be performed so that the diluted result reads greater than 40 ng/mL or greater.

e.  For detailed information on ordering dilutions, refer to the ARCHITECT System Operations Manual, Section 5.

Precautions and Warnings

1.  For In Vitro Diagnostic Use.

This product requires the handling of human specimens.  It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens.  Biosafety Level 2 or other appropriate biosafety practices should be used for materials that contain or are suspected of containing infectious agents.

1.  The ARCHITECT STAT Myoglobin Assay Diluent and Conjugate contain a mixture of:  5-chloro-2-methyl-4-isothiazolin-3-one and 2-methyl-4-isothiazolin-3-one (3:1), which is a component of ProClin, and are classified per applicable European Community (EC) Directives as:  Irritant (Xi).  The following are the appropriate Risk (R) and Safety (S) phrases.

R43  May cause sensitization by skin contact.

S24  Avoid skin contact.

S35  This material and its container must be disposed of in a safe way.

S37  Wear suitable gloves.

S46  If swallowed, seek medical advice immediately and show this container or label.

2.  ARCHITECT i Trigger Solution contains sodium hydroxide (NaOH) and is classified per applicable European Community (EC) Directives as:  Irritant (Xi).  The following are the appropriate Risk (R) and Safety (S) phrases.

R41  Risk of serious damage to eyes.

S25  Avoid contact with eyes.

S26  In case of contact with eyes, rinse immediately with plenty of water and seek medical advice.

S35  This material and its container must be disposed of in a safe way.

S36/39  Wear suitable protective clothing and eye/face protection.

S46  If swallowed, seek medical advice immediately and show this container or label.

3.  Information for European customers:  For product not classified as dangerous per European Directive 1999/45/EC - Safety data sheet available for professional user on request.

4.  For a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8.

HANDLING PRECAUTIONS

1.  Do not use reagent kits beyond the expiration date.

2.  Do not pool reagents within a kit or between reagent kits.

3.  Prior to loading the ARCHITECT STAT Myoglobin Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.  For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this procedure.

4.  Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity.  Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in your package insert.

5.  To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

6.  Prior to placing the septum on an uncapped reagent bottle, squeeze the septum in half to confirm that the slits are open.  If the slits appear sealed, continue to gently squeeze the septum to open the slits.

7.  Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

8.  Over time, residual liquids may dry on the septum surface.  These are typically dried salts which have no effect on assay efficacy.

9.  For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT i System Operations Manual, Section 7.

## Limitations — interference

.  Cardiac Myoglobin levels can be increased in any condition resulting in cardiac cell damage.  For MI diagnostic purposes, the ARCHITECT STAT Troponin-I results should be used in conjunction with other information such as cardiac marker results (e.g., CK-MB and/or myoglobin), ECG, clinical observations and symptoms, etc.

2.  Specimens from patients who have received preparations of mouse monoclonal antibodies for diagnosis or therapy may contain human anti-mouse antibodies (HAMA).  Such specimens may show either falsely elevated or depressed values when tested with assay kits that employ mouse monoclonal antibodies.  Additional clinical or diagnostic information may be required to determine patient status.

3.  Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.  The presence of heterophilic antibodies in a patient specimen may cause anomalous values to be observed.  Additional information may be required for diagnosis.

4.  Although the ARCHITECT STAT Myoglobin assay is specifically designed to minimize the effects of HAMA and heterophilic antibodies, assay results that are not consistent with other clinical observations may require additional information for diagnosis.

Contacts:

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Alternative method

Send to sister facility

References

ABBOTT STAT Myoglobin package insert

Abbott Laboratories

Diagnostics Division

Abbott Park, Illinois 60064

PN 840174/R3

Revision 07/04

Effective date

Effective date for this procedure: 07/01/2014

Author

Compiled by Abbott Diagnostics

Written by:  	 Jessica Mercer, MLS (ASCP)