Intended use

The ARCHITECT Tacrolimus assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of tacrolimus in human whole blood on the ARCHITECT i System. The ARCHITECT Tacrolimus assay is to be used as an aid in the management of liver and kidney allograft patients receiving tacrolimus therapy.1,2

Summary

Tacrolimus is an immunosuppressive drug discovered in 1984 by the Fujisawa Pharmaceutical Co., Ltd. It has been shown to be effective for the treatment of organ rejection following transplantation. The results of clinical trials with liver3 and kidney4,5 have been published. Clinical studies are continuing for a variety of indications. Tacrolimus binds to a family of proteins termed FK506 (tacrolimus) binding proteins (FKBPs).6,7 The formation of a larger pentameric complex comprised of FKBP, tacrolimus, calmodulin and calcineurins A and B results in the inhibition of the phosphatase activity of calcineurin.8 The action of transcription factors requiring dephosphorylation for transport to the cell nucleus are thus inhibited leading to blockage of T-cell proliferation and function.

Tacrolimus may be administered IV or orally. Absorption from the gastrointestinal tract is variable and irregular.9 Pharmacokinetic studies with tacrolimus have shown that there are large inter- and intra-individual differences in its kinetics in organ transplant patients.10,11 Pharmacokinetic studies have also indicated that whole blood rather than plasma may serve as the more appropriate medium to describe the pharmacokinetic characteristics of tacrolimus. Tacrolimus is bound to proteins, mainly albumin and α-1-acid glycoprotein, and is highly bound to erythrocytes. The distribution of tacrolimus between whole blood and plasma depends on several factors such as hematocrit, temperature of separation of plasma, drug concentration, and plasma protein concentration. In a U.S. study, the ratio of whole blood concentration to plasma concentration ranged from 12 to 67 (mean 35).

Tacrolimus is extensively metabolized in the liver and small intestine microsomes utilizing the cytochrome P-450 enzymes.13 Nine different metabolites of tacrolimus have been identified; several of the metabolites have been found and tested in whole blood.14-18

The use of tacrolimus is associated with serious toxic side effects, primarily nephrotoxicity.19,20 At the present time it is not clear whether the nephrotoxicity of tacrolimus is the result of parent drug, metabolites, or a combination of both. Other adverse side effects include neurotoxicity, hypertension, insomnia, and nausea.21

## Method

The ARCHITECT Tacrolimus assay is a delayed one-step immunoassay for the quantitative determination of tacrolimus in human whole blood using CMIA technology with flexible assay protocols, referred to as Chemiflex.

## Principle

Prior to the initiation of the automated ARCHITECT sequence, a manual pretreatment step is performed in which the whole blood sample is extracted with a precipitation reagent and centrifuged. The supernatant is decanted into a Transplant Pretreatment Tube, which is placed onto the 	         ARCHITECT i System.

Sample, assay diluent, and anti-tacrolimus coated paramagnetic microparticles are combined to create a reaction mixture. Tacrolimus present in the sample binds to the anti-tacrolimus coated microparticles.

After a delay, tacrolimus acridinium-labeled conjugate is added to the reaction mixture. The tacrolimus on the acridinium-labeled conjugate competes for the available binding sites on the microparticles. Following

an incubation, the microparticles are washed, and pre-trigger and trigger solutions are added to the reaction mixture. The resulting chemiluminescent reaction is measured as relative light units (RLUs).

An indirect relationship exists between the amount of tacrolimus in the sample and the RLUs detected by the ARCHITECT i System optics.

For additional information on system and assay technology refer to the

ARCHITECT System Operations Manual, Section 3.

## Specimen collection and handling

1.  The following specimen collection tubes may be used for the ARCHITECT STAT Tacrolimus assay.

Whole Blood:	Glass: 	Will not be used

Plastic:  EDTA (lavender)

Other anticoagulants and serum/plasma have not been validated for use with the ARCHITECT Tacrolimus assay.

2.  Refer to the specimen collection tube manufacturer's instructions as well as these package insert instructions for specimen collection and preparation for analysis.  Each laboratory should determine the acceptability of its own specimen collection and processing procedures.

3.  DO NOT CENTRIFUGE. Centrifugation of the specimen prior to pretreatment may cause an erroneous result.

4.  Serum/Plasma samples can not be accepted for testing. Erroneous results will result when       the wrong specimen type is ran.

5.  If a lipid layer forms on the specimen surface, avoid the lipid layer when withdrawing the specimen.

6. Specimen are stable at 2-8°C for 7 days.  If testing will be delayed more than 7 days, store frozen (-10 degrees C. or colder) prior to being tested. Specimens must be mixed thouroughly upon thaswing prior to testing to ensure proper results.

7.  Specimens can be stored up to 6 months frozen at -10 degrees C. or colder and yield >90% recovery.

8.  Thaw frozen specimens and mix thoroughly by LOW speed vortexing or by gently inverting the specimen 10 times until completely mixed. Thaw specimens only once.

9.  Ensure specimens are free of fibrin, red blood cells, and other particulate matter.

10.  Inspect all samples for bubbles.  Remove bubbles prior to analysis.

11.  When serial specimens are being evaluated, use the same specimen type throughout the evaluation.

12.  The ARCHITECT i System does not provide the capability to verify specimen type.  It is the responsibility of the operator to verify the correct specimen type is used in the ARCHITECT Tacrolimus assay.

13.  Use caution when handling patient specimens to prevent cross contamination.  Use of disposable pipettes or pipette tips is recommended.

14.  Test all samples (patient specimens, controls, and calibrators) within 3 hours of being placed on board the ARCHITECT i System.  Refer to the ARCHITECT System Operations Manual, Section 5, for a more detailed discussion of on-board sample storage constraints.

15.  Performance has not been established using cadaver specimens or body fluids other than human plasma or serum.

16.  Do not use samples with obvious microbial contamination.

17.  Do not use heat-inactivated specimens.

18.  When shipping specimens, package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.  Ship specimens frozen on wet or dry ice.

Materials and Equipment Required

Provided

- 1L77 ARCHITECT Tacrolimus Reagent Kit

- 1L77-55 ARCHITECT Tacrolimus Whole Blood Precipitation Reagent

- 9527-40 XSYSTEMS Centrifuge Tubes

- 1P06 Transplant Treatment Tubes

Materials required but not provided

- ARCHITECT i System

- ARCHITECT I System, add

- ARCHITECT Tacrolimus Assay file, may be obtained from: www.abbottdiagnostics.com

3K50 ARCHITECT i	-US- Addition A

3k52 ARCHITECT i	-WW (excluding US)- Addition A

- 1L77-01 ARCHITECT Tacrolimus Calibrators

- 1P05-10 Abbott Immunosuppressant-MCC Controls or other commercial controls.

- Vortex Mixer

- Precision Micropipettes

- Pipette Tips

- 9528-02 XSYSTEMS Precision Dispenser, or equivalent

- 2.5 mL Comb tips, or equivalent, for precision dispenser

- 9257-26 XSYSTEMS Centrifuge or equivalent laboratory micro centrifuge.

ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- Pipettes or pipette tips (optional) to deliver the volumes specified on the patient or control order screen.

For information on materials required for maintenance procedures, refer to the ARCHITECT System Operations Manual, Section 9.

## Reagents – working solutions

Reagent Kit, 100 Tests/500 Tests

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT i Systems. Please contact your local distributor.

ARCHITECT Tacrolimus Reagent Kit (1L77)

1.  Microparticles:  1 Bottle (6.6 mL/27.0 mL) Anti-tacrolimus (mouse, monoclonal) coated microparticles in EDTA buffer with protein (bovine ) stabilizers. Minimum Concenration:0.09% solids,  Preservative:  sodium azide and ProClin 950.

2.  Conjugate:  1 Bottle (7.8 mL/16.0 mL) Tacrolimus acridinium-labeled conjugate in citrate buffer with protein (bovine) stabilizer. Minimum Concentration: 5.0 ng/mL,  Preservative: ProClin 300.

3.  Assay Diluent:  1 (8.9 mL/45.8 mL) Assay Diluent, containing MES buffer and sodium chloride. Preservative: ProClin 950 and ProClin 300.

Other Reagents

ARCHITECT i Pre-Trigger Solution

Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT i Trigger Solution

Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT i Wash Buffer

Wash Buffer containing phosphate buffered saline solution. Preservative: antimicrobial agents.

|  |  | ## Storage and stability 1. The ARCHITECT Tacrolimus Reagent Kit must be stored at 2-8 degrees C. in an upright position and may be used immediately after removal from 2-8 degrees C. storage. 2. When stored and handled as directed, reagents are stable until the expiration date. 3. The ARCHITECT Tacrolimus Reagent Kit may be stored on-board the ARCHITECT i System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking on-board time, refer to the ARCHITECT System Operations Manual, Section 5. 4. Reagents may be stored on or off the ARCHITECT i System. If reagents are removed from the system, store them at 2-8 degrees C. (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If the microparticle bottle does not remain upright (with a septum installed) while in 2-8 degrees C. storage off the system, the reagent kit must be discarded. After reagents are removed from the system, one must initiate a scan to update the on-board stability timer. INDICATIONS OF REAGENT DETERIORATION When a control value is out of the specified range, it may indicate deterioration of the reagents or errors in technique. Associated test results may be invalid and may require retesting. Assay recalibration may be necessary. For troubleshooting information, refer to the Architect System Operations Manual, Section 10 ## Manual Pretreatment Procedure - The ARCHITECT Tacrolimus Assay requires a manual pretreatment step for all whole blood patient specimens, ARCHITECT Tacrolimus Calibrators, and Abbott Immunosuppressant-MCC or other controls (Bio-Rad Whole Blood Immunosuppressant Controls). - Use only ARCHITECT Tacrolimus Whole Blood Precipitation Reagent (1L77-55). - Once the Manual Pretreatment Procedure has been initiated, all steps must be completed in immediate succession. NOTE: If specimen requires dilution, it must be diluted prior to the manual pretreatment step. Refer to Specimen Dilution Pretreatment section of package insert. |  |  |
| --- | --- | --- | --- | --- |
| Manual Pretreatment Procedure - Mix each sample (specimen, calibrator or control) thoroughly by slow inversion of the container 5-10 times; older whole blood specimens may take longer mixing times. Visual inspection is recommended to assure the specimen is adequately mixed. - Precision pipette 200ul of each sample into an XSYSTEMS Centrifuge Tube immediately after mixing. Use a different tube for each sample. - NOTE: A new pipette tip must be used each time a sample is aspirated - Dispense 200 uL of Tacrolimus Whole Blood Precipitation Reagent from the blue labeled bottle - Ensure that bubbles are not present in the pipete tips. - Add the 200 uL of Tacrolimus Whole Blood Precipitation Reagent to the Centrifuge Tube containing the 200 uL of sample; dispense the Reagent slowly along the wall of the tube. - Cap the tube and vortex immediately |  |  |  |  |
| - Vortex vigorously for 5-10 seconds using the maximum vortex setting. NOTE: Visually inspect that the mixture of the sample is uniform, smooth, and homogeneous. - Repeat the procedure until all specimens have been processed. - Centrifuge each specimen for a minimum of 4 minutes at >9500 x gRCF or 38,500 g-minutes. - Remove each tube and inspect for the presence of a well-formed pellet and clear supernatant. - Uncap each tube and decant (pour off ) the supernatant into the Transplant Pretreatment Tube, when the ARCHITECT I system is ready to accept sample - NOTE: DO NOT pipette supernatant off of pellet. - All pretreated samples (specimens, calibrators or controls) must be tested within 30 minutes of being decanted. |  |  |  |  |

## Calibration

1.  To perform an ARCHITECT Tacrolimus calibration, test the Calibrators A, B, C, D, E, and F in duplicate.  Only one pretreated sample of each ARCHITECT Tacrolimus Calibrator is required to perform a calibration on the ARCHITECT I System. This provides adequate volume to run each calibrator in duplicate. A single sample of each Tacrolimus control must be tested to evaluate the assay calibration.  Ensure that assay control values are within the concentration ranges specified in the control package insert.  Calibrators should be priority loaded. (The use of a sample cup inserted into the pretreatment tube will prevent aspiration errors due to low volume).

2.  Calibration range:  0.0 – 30.0 ng/mL.

3.  Once an ARCHITECT Tacrolimus calibration is accepted and stored, all subsequent samples may be tested without further calibration unless:

a. A reagent kit with a new lot number is used

b. Controls are out of range

4.  For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Quality control

The recommended control requirement for the ARCHITECT Tacrolimus assay is that a single sample of each control level be tested once every 24 hours each day of use. If the quality control procedures in your laboratory require more frequent use of controls to verify test results, follow your laboratory-specific procedures.

The ARCHITECT Tacrolimus control values must be within the acceptable ranges specified in the control package insert. If a control is out of the specified range, the associated test results are invalid and samples must be retested. Recalibration may be indicated.

For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Preparation of Working Solutions

Handling Precautions

•   Do not use reagent kits beyond the expiration date.

•   Do not pool reagents within a kit or between reagent lots.

• Prior to loading the ARCHITECT Tacrolimus Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this package insert.

•   Septum’s MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

•   To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

•   Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

- Over time, residual liquids may dry on the septum surface. These are typically dried salts, which have no effect on assay efficacy.

- For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay Procedure

INSTRUMENT PROCEDURE

1.  The ARCHITECT Tacrolimus assay file must be installed on the ARCHITECT i System with the ARCHITECT i Assay CD-ROM Addition A  prior to performing the assay.  For detailed information on assay file installation and on viewing and editing assay parameters, refer to the ARCHITECT System Operations Manual, Section 2.

2.  For information of printing assay parameters, refer to the ARCHITECT System Operations Manual, Section 5.

3.  For a detailed description of system procedures, refer to the ARCHITECT System Operations Manual.

4.  The default result unit for the ARCHITECT Tacrolimus assay is ng/mL.

ASSAY PROCEDURE

1.  Before loading the ARCHITECT Tacrolimus Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment:

a. Invert the microparticle bottle 30 times.

b. Visually inspect the bottle to ensure microparticles are resuspended.  If microparticles are still adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

c. If the microparticles do not resuspend, DO NOT USE.  Contact your local Abbott Laboratories representative.

d. Once the microparticles have been resuspended, remove and discard the cap.  Wearing clean gloves, remove a septum from the bag.  Squeeze the septum in half to confirm that the slits are open.  Carefully snap the septum onto the top of the bottle.

2.  Load the ARCHITECT Tacrolimus Reagent Kit on the ARCHITECT i System.

a. Verify that all necessary assay reagents are present.  Ensure that septa are present on all reagent bottles.

3.  Order calibration, if necessary

a. For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

4.  Order tests.

- For information on ordering patient specimens and controls, refer to the ARCHITECT System Operations Manual, Section 5.

- NO MORE than 3 replicates may be sampled from the same Transplant Pretreatment Tube.

- All pretreated samples (specimens, calibrators, or conrols)

must be  test within 30 minutes  of being decanted into the

Transplant Pretreatment Tubes and placed on the ARCHITECT

i system.

- All ARCHITECT Tacrolimus samples must be priority loaded. Priority loading of samples prevents evaporation that will impact the assay results. No more than 24 Tacrolimus samples may be loaded on the  i1000 system at the same time. For information on priority loading of samples, please refer to the ARCHITECT System Operations Manual, Section 5

- With the Transplant Pretreatment Tube, use the sample gauge to ensure sufficient patient specimen is present for the ARCHITECT Tacrolimus Assay.

5.  The minimum sample volume is calculated by the system and is printed on the Order list report.  No more than 3 replicates may be sampled from the same sample cup.  To minimize the effects of evaporation, verify adequate sample cup volume is present prior to running the test.

a. If using primary or aliquot tubes, use the sample gauge to ensure sufficient patient specimen is present.

6.  Prepare calibrators and controls.

a. Refer to the Manual Pretreatment Procedure Section.

7.  Load pretreated samples.

a. For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

8.  Press RUN.  The system performs the following functions:

a. Moves the sample to the aspiration point.

b. Loads a reaction vessel (RV) into the process path.

c. Aspirates and transfers sample into the RV.

d.  Advances the RV one position and transfers assay diluent and microparticles into the RV.

e. Mixes, incubates, and washes the reaction mixture.

f. Ads conjugate to the RV.

g. Mixes, incubates, and washes the reaction mixture.

h. Ads pre-trigger and trigger solutions.

i. Measures chemiluminescent emission to determine the quantity of troponin-I in the sample.

j. Aspirates contents of RV to liquid waste and unloads RV to solid waste.

k. Calculates the result.

9.  For optimal performance, it is important to follow the routine maintenance procedures defined in the ARCHITECT System Operations Manual, Section 9.  If your laboratory requires more frequent maintenance, follow those procedures.

#### Interpretation: reporting results

Calculation

- The ARCHITECT Tacrolimus assay uses a 4 Parameter Logistic Curve Fit (4PLC, Y-weighted) data reduction method to generate a calibration curve

Flags

- Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

- Therapeutic  Range

|  | Therapeutic Range: |
| --- | --- |
| Kidney transplant: | 0-​3 months post-​transplant: 7.0-​20.0 ng/mL 3 months and older: 5.0-​15.0 ng/mL |
| Heart transplant: | 0-​3 months post-​transplant: 10.0-​20.0 ng/mL 3 months and older: 5.0-​15.0 ng/mL |
| Liver transplant: | 1-​12 months post-​transplant: 5-​20 ng/mL |
| Toxic value: | Greater than 25 ng/mL |

Critical Values:  >25.0 ng/mL

All critical values will be called to the appropriate physician/nurse (RN) and documented using the 1CALL canned text comment.

For diagnostic cutoff and additional information, refer to the SPECIFIC PERFORMANCE CHARACTERISTICS, Clinical Performance section in this package insert.

CAUTION: No Firm therapeutic range exists for tacrolimus in whole blood. The complexity of the clinical state, individual differences in sensitivity to immunosuppressive and nephrotoxic effects of tacrolimus, co-administration of other immunosuppressant’s, type of transplant, time post-transplant and a number of other factors contribute to different requirements for optimal blood levels of tacrolimus. Therefore, individual tacrolimus values cannot be used as a sole indicator for making changes in treatment regimen and each patient should be thoroughly evaluated clinically before making changes in treatment regimen are made.

Therapeutic ranges vary according to the commercial test used, and therefore should be established for each commercial test. Values obtained with different assay methods cannot be used interchangeably due to differences in assay methods and cross-reactivity with metabolites, not should correction factors be applied. Therefore, consistent use of one assay for individual patients is recommended.

The Consensus Document has reported that the therapeutic range of tacrolimus is not clearly define, but target 12-hour trough whole blood concentrations are 5-20ng/mL early post-transplant. Higher concentrations are associated with an increase incidence of adverse effects. Twenty-four hour trough concentrations are 33-50% less than the corresponding 12-hour trough levels.

## Measuring Range:

- The measurement range for the ARCHITECT Tacrolimus assay is 2 ng/mL to 30 ng/mL.

Dilutions

Specimens with a Tacrolimus value exceeding 30.0 ng/mL are flagged with the code "> 30.0 ng/mL" and may be diluted with the Manual Dilution Procedure. The STAT Laboratory does not report diluted values and will report dilutions as >30.00 ng/mL.

MANUAL DILUTION PROCEDURE

1.  Manual dilutions should be performed as follows:

- The suggested dilution for a Tacrolimus test is 1:2.

- Specimens must be diluted before pretreatment.

c. Add 150 uL of the patient specimen to 150 uL of ARCHITECT Tacrolimus Calibrator A, then precede wih the Manual Pretreatment Procedure see Manual Pretreatment Section.

d.  The operator must enter the dilution factor in the Patient or Control order screen.  The system will use this dilution factor to automatically calculate the concentration of the sample before dilution.  This will be the reported result.  The dilution should be performed so that the diluted result reads greater than 3.0 ng/mL or greater.

e.  For detailed information on ordering dilutions, refer to the ARCHITECT System Operations Manual, Section 5.

Precautions and Warnings

1.  For In Vitro Diagnostic Use.

This product requires the handling of human specimens.  It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens.  Biosafety Level 2 or other appropriate biosafety practices should be used for materials that contain or are suspected of containing infectious agents.

1.  The microparticles, conjugate, and assay diluent contain methylisothiazolones, which are components of ProClin, and are classified per applicable European Community (EC) Directives as: Irritant (Xi). The following are the appropriate Risk (R) and Safety (S) phrases..

R43  May cause sensitization by skin contact.

S24  Avoid skin contact.

S35  This material and its container must be disposed of in a safe way.

S37  Wear suitable gloves.

S46  If swallowed, seek medical advice immediately and show this container or label.

2.  ARCHITECT i Trigger Solution contains sodium hydroxide (NaOH) and is classified per applicable European Community (EC) Directives as:  Irritant (Xi).  The following are the appropriate Risk (R) and Safety (S) phrases.

R41  Risk of serious damage to eyes.

S25  Avoid contact with eyes.

S26  In case of contact with eyes, rinse immediately with plenty of water and seek medical advice.

S35  This material and its container must be disposed of in a safe way.

S36/39  Wear suitable protective clothing and eye/face protection.

S46  If swallowed, seek medical advice immediately and show this container or label.

3.  Information for European customers:  For product not classified as dangerous per European Directive 1999/45/EC - Safety data sheet available for professional user on request.

4.  For a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8.

HANDLING PRECAUTIONS

1.  Do not use reagent kits beyond the expiration date.

2.  Do not pool reagents within a kit or between reagent kits.

3.  Prior to loading the ARCHITECT Tacrolimus Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.  For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this procedure.

4.  Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity.  Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in your package insert.

5.  To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

6.  Prior to placing the septum on an uncapped reagent bottle, squeeze the septum in half to confirm that the slits are open.  If the slits appear sealed, continue to gently squeeze the septum to open the slits.

7.  Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

8.  Over time, residual liquids may dry on the septum surface.  These are typically dried salts which have no effect on assay efficacy.

9.  For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT i System Operations Manual, Section 7.

## Limitations — interference

. For diagnostic purposes, results should be used in conjunction with other data; e.g., symptoms, results of other tests, clinical impressions,etc.

• If the tacrolimus results are inconsistent with clinical evidence, additional testing is suggested to confirm the result.

• The concentration of tacrolimus in a given specimen determined with assays from different manufacturers can vary due to differences in assay methods and reagent specificity.

• Immunoassays are nonspecific and cross react with metabolites. When elimination of tacrolimus is impaired (e.g. during cholestasis), tacrolimus metabolites may accumulate. The immunoassay

may overestimate the concentration of tacrolimus. In such cases, the use of a specific assay (e.g. Liquid Chromatography Mass Spectrometry/Mass Spectrometry [LC/MS/MS]) could be considered.

Refer to the SPECIFICITY section below for estimates of cross-reactivity of ARCHITECT Tacrolimus to some metabolites of tacrolimus. Refer to the METHOD COMPARISON section below for representative data comparing patient results from the ARCHITECT Tacrolimus assay to the IMx Tacrolimus II assay and an LC/MS/MS method.

• Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.27 Patients routinely exposed to animals or to animal serum products can be

prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

• Specimens from patients who have received preparations of mouse monoclonal antibodies for diagnosis or therapy may contain human anti-mouse antibodies (HAMA).28,29 Specimens containing HAMA may

produce anomalous values when tested with assay kits (such as ARCHITECT Tacrolimus) that employ mouse monoclonal antibodies.

Contacts:

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Alternative method

Send to ARUP or Driscoll (STAT Orders)

References

ABBOTT Tacrolimus package insert

Abbott Laboratories

Diagnostics Division

Abbott Park, Illinois 60064

PN 840174/R3

Revision 07/04

Effective date

Effective date for this procedure: 04/15/2015

Author

Compiled by Abbott Diagnostics

Written by:  	 Jessica Mercer, MLS (ASCP)