Intended use

The ARCHITECT CORE assay is a chemiluminescent microparticle immunoassay (CMIA) for the qualitative  detection of IgG and IgM antibodies to hepatitis B core antigen (anti-HBc) in human adult and

pediatric serum and plasma (dipotassium EDTA, lithium heparin, sodium heparin) and neonatal serum.

It is intended as an aid in the diagnosis of acute, chronic, or resolved hepatitis B virus (HBV) infection in conjunction with other laboratory results and clinical information.

Warning: Not intended for use in screening blood, plasma, or tissue donors. The effectiveness of ARCHITECT CORE for use in screening blood, plasma, or tissue donors has not been established.

Assay performance characteristics have not been established when the ARCHITECT CORE assay is used in conjunction with other manufacturers’ assays for specific hepatitis markers. Users are responsible for establishing their own performance characteristics.

Assay performance characteristics have not been established for immunocompromised or immunosuppressed patients. Users are responsible for establishing their own assay performance characteristics in these populations.

Summary

The ARCHITECT CORE assay utilizes microparticles coated with recombinant hepatitis B virus core antigen (rHBcAg) for the detection of anti-HBc antibodies. Anti-HBc antibody determinations can be used as an indicator of current or past HBV infection. Anti-HBc antibodies are found in serum shortly after the appearance of hepatitis B surface antigen (HBsAg) in acute HBV infections. They will persist after the

disappearance of HBsAg and before the appearance of detectable antibody to HBsAg (anti-HBs).1-7 In the absence of information about any other HBV markers, it must be considered that an individual with detectable levels of anti-HBc antibodies may be actively infected with HBV or that the infection may have resolved, leaving the person immune.8 Anti-HBc antibodies may be the only serological marker of HBV infection and potentially infectious blood.9-16

The presence of anti-HBc antibodies does not differentiate between acute or chronic hepatitis B infection.

## Method

The ARCHITECT CORE assay is a two-step immunoassay for the qualitative determination of anti-HBc antibodies in human serum and plasma using CMIA technology with flexible assay protocols, referred to as Chemiflex.

## Principle

1. Sample, assay diluent, speciment diluent, and rHBcAg coated paramagnetic microparticles are combined. The anti-HBc antibodies present in the sample bind to the rHBcAg coated microparticles.

2. After washing, anti-human IgG and IgM acridinium-labeled conjugate is added to create a reaction mixture.

3. Following another wash cycle, Pre-Trigger and Trigger Solutions are added to the reaction mixture.

4. The resulting chemiluminescent reaction is measured as relative light units (RLUs). There is a relationship between the presence of anti-HBc antibodies in the sample and the RLUs detected by the ARCHITECT iSystem optics.

The presence or absence of anti-HBc antibodies in the specimen is determined by comparing the chemiluminescent signal in the reaction to the cutoff signal determined from an active calibration.

For additional information on system and assay technology, refer to the ARCHITECT System Operations Manual, Section 3.

## Specimen collection and handling

Verified specimen types to be used with this assay:

|  | Collection Tubes |  |
| --- | --- | --- |
| Specimen Types | Glass | Plastic |
| Serum | Serum | Serum Serum separator |
| Plasma |  | Lithium heparin plasma separator Sodium heparin Dipotassium EDTA |

• Assay performance characteristics have not been established for any specimen matrices other than serum or plasma (dipotassium EDTA, lithium heparin, sodium heparin).

• Performance has not been established for the use of cadaveric specimens or the use of body fluids other than human serum and plasma.

• The instrument does not provide the capability to verify specimen type. It is the responsibility of the operator to verify that the correct specimen types are used in the assay.

Specimen Conditions

• Do not use specimens with the following conditions:

• heat-inactivated

• pooled

• grossly hemolyzed

• obvious microbial contamination

• For accurate results, serum and plasma specimens should be free of fibrin, red blood cells, and other particulate matter. Serum specimens from patients receiving anticoagulant or thrombolytic therapy may

contain fibrin due to incomplete clot formation.

• To prevent cross contamination, use of disposable pipettes or pipette tips is recommended.

Preparation for Analysis

• Follow the tube manufacturer’s processing instructions for collection tubes. Gravity separation is not sufficient for specimen preparation.

• Mix thawed specimens thoroughly by low speed vortexing or by inverting 10 times. Visually inspect the specimens. If layering or stratification is observed, continue mixing until specimens are visibly homogeneous.

If samples are not homogeneous, inconsistent results may be obtained.

• To ensure consistency in results, specimens must be transferred to a centrifuge tube and centrifuged at > 10,000 RCF (Relative Centrifugal Force) for 10 minutes before testing if

• they contain fibrin, red blood cells, or other particulate matter or

• they were frozen and thawed.

• Centrifuged specimens with a lipid layer on the top must be transferred to a sample cup or secondary tube. Care must be taken to transfer only the clarified specimen without the lipemic material.

• Transfer clarified specimen to a sample cup or secondary tube for testing.

• Inspect all specimens for bubbles. Remove bubbles with an applicator stick before analysis. Use a new applicator stick for eachspecimen to prevent cross contamination.

Specimen Storage

| Specimen Type | Storage Temperature | Maximum Storage Time |
| --- | --- | --- |
| Serum/Plasma | Room temperature (study performed at 23 to 30°C) | ≤ 3 days |
|  | 2-8°C | ≤ 7 days |
|  | -20°C or colder | -- |

Specimens may be stored on or off the clot, red blood cells, or separator gel.

Remove serum or plasma from the clot, red blood cells, or separator gel if stored longer than the maximum room temperature or 2-8°C storage time and store frozen at -20°C or colder.

Avoid more than 3 freeze/thaw cycles.

Specimen Shipping

• Package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.

• It is recommended that specimens be removed from the clot, red blood cells, or separator gel.

• Ship ambient, at 2-8°C (wet ice), or frozen (dry ice).

• Do not exceed the storage time limitations listed above.

Materials and Equipment Required

Materials Provided

6L22 ARCHITECT CORE Reagent Kit

Materials Required but not Provided

• ARCHITECT CORE Assay file obtained from the ARCHITECT iSystem e-Assay CD-ROM found on www.abbottdiagnostics.com.

• 6L22-01 ARCHITECT CORE Calibrator

• 6L22-10 ARCHITECT CORE Controls or other control material

• ARCHITECT Pre-Trigger Solution

• ARCHITECT Trigger Solution

• ARCHITECT Wash Buffer

• ARCHITECT Reaction Vessels

• ARCHITECT Sample Cups

• ARCHITECT Septum

• ARCHITECT Replacement Caps

• Pipettes or pipette tips (optional) to deliver the volumes specified on the patient or control order screen.

For information on materials required for maintenance procedures, refer to the ARCHITECT System Operations Manual, Section 9.

## Reagents – working solutions

Kit Contents

ARCHITECT CORE 6L22

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT iSystems. Please contact your local distributor.

| REF | 6L22-25 | 6L22-35 |
| --- | --- | --- |
|  | 100 | 500 |
| MICROPARTICLES | 1 x 6.60 mL | 1 x 27.00 mL |
| CONJUGATE | 1 x 11.00 mL | 1 x 28.82 mL |
| ASSAY DILUENT | 1 x 5.36 mL | 1 x 23.72 mL |
| SPECIMEN DILUENT | 1 x 5.36 mL | 1 x 23.72 mL |
| MICROPARTICLES | Hepatitis B core (E. coli, recombinant) antigen coated microparticles in TRIS buffer. Minimum concentration: 0.08% solids. Preservatives: ProClin 950 and sodium azide. |  |
| CONJUGATE | Anti-human (mouse, monoclonal) acridinium-labeled conjugate in MES buffer with protein stabilizers (bovine). Minimum concentration: 0.048 μg/mL. Preservatives: sodium alkyl paraben andsodium azide. |  |
| ASSAY DILUENT | Assay diluent containing protein stabilizers (mouse) in MOPSO buffer. Preservatives: ProClin 950 and sodium azide. |  |
| SPECIMEN DILUENT | Specimen diluent containing reductant in MOPSO buffer. |  |

Other Reagents

ARCHITECT Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT Wash Buffer containing phosphate buffered saline solution. Preservatives: antimicrobial agents.

## Storage and stability

| When stored and handled as directed, reagents are stable until the expiration date. Storage Temperature; Maximum Storage Time; Additional Storage Instructions. Unopened/Opened*; 2-8°C; Until expiration date; May be used immediately after removal from 2-8°C storage. Store in upright position.. On board; System Temperature; 30 Days; Discard after 30 days. For information on tracking onboard time, refer to the ARCHITECT System Operations Manual, Section 5.. * Reagents may be stored on or off the ARCHITECT iSystem. If reagents are removed from the system, store them at 2-8°C (with septums and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded. For information on unloading reagents, refer to the ARCHITECT System Operations Manual, Section 5. Indications of Reagent Deterioration When a control value is out of the specified range, it may indicate deterioration of the reagents or errors in technique. Associated test results are invalid, and samples must be retested. Assay recalibration may be necessary. For troubleshooting information, refer to the ARCHITECT System Operations Manual, Section 10. |  |
| --- | --- |

## Calibration

• Test Calibrator 1 in triplicate. The calibrator should be priority loaded.

• A single sample of each control level must be tested to evaluate the assay calibration.

• Order controls as described in the package insert.

• Ensure that assay control values are within the S/CO ranges specified in the control package insert.

• Once an ARCHITECT CORE calibration is accepted and stored, all subsequent samples may be tested without further calibration unless:

• A reagent kit with a new lot number is used or

• Controls are out of range.

## Quality control

The ARCHITECT CORE Controls are in a serum matrix made from recalcified plasma. The user should provide alternate control material for plasma when necessary.

The recommended control requirement for the ARCHITECT CORE assay is that a single sample of each control be tested once every 24 hours each day of use. Additional controls may be tested in accordance with

local, state, and/or federal regulations or accreditation requirements and your laboratory’s quality control policy.

Control values must be within the ranges specified in the control package insert. If a control result is out of its specified range, any test results generated since the last acceptable control results must be evaluated to determine if sample results may have been adversely affected. Adversely affected test results are invalid, and these samples must be retested. For troubleshooting information, refer to the ARCHITECT System Operations Manual, Section 10.21, 22

## Preparation of Working Solutions

Handling Precautions

• Do not use reagent kits beyond the expiration date.

• Do not pool reagents within a kit or between kits.

• Before loading the reagent kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that may have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section ofthis package insert.

• Septums MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

• To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

• Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

• Over time, residual liquids may dry on the septum surface. These are typically dried salts and have no effect on assay efficacy.

• When handling conjugate vials, change gloves that have contacted human serum or plasma, since introduction of antibody will result in a neutralized conjugate.

For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay Procedure

INSTRUMENT PROCEDURE

The ARCHITECT CORE assay file must be installed on the ARCHITECT iSystem from an ARCHITECT iSystem Assay CD-ROM prior to performing the assay.

For detailed information on assay file installation and viewing and editing assay parameters, refer to the ARCHITECT System Operations Manual, Section 2.

For information on printing assay parameters, refer to the ARCHITECT System Operations Manual, Section 5.

For a detailed description of system procedures, refer to the ARCHITECT System Operations Manual.

ASSAY PROCEDURE

• Before loading the reagent kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that may have settled during shipment. After the first time the microparticles have been loaded, no further mixing is required.

• Invert the microparticle bottle 30 times.

• Visually inspect the bottle to ensure microparticles are resuspended. If microparticles are still adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

• If the microparticles do not resuspend, DO NOT USE. Contact your local Abbott Laboratoriesrepresentative.

• Once the microparticles have been resuspended, discard the cap and place a septum on the bottle. For instructions about placing septums on bottles, refer to the Reagent Handling section of this package insert.

• Load the reagent kit on the ARCHITECT iSystem.

• Verify that all necessary reagents are present.

• Ensure that septums are present on all reagent bottles.

• Order calibration, if necessary.

• For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

• Order tests.

• If utilizing ARCHITECT system software version 5.0 or higher, refer to the ARCHITECT System Operations Manual, Section 5, for information on ordering patient specimens and controls.

• If utilizing an ARCHITECT system software version lower than 5.0, use the following  instructions to order patient specimens and controls:

• For information on ordering patient specimens and controls and for general operating procedures, refer to the ARCHITECT System Operations Manual, Section 5.

• Use the following instructions to order a negative control (nonreactive for anti-HBc antibodies):

• Order a negative control as a patient specimen, not as a Control.

• Manually verify the validity of the negative control every time it is run. Because the control is run as a patient specimen, a result will not be flagged by the ARCHITECT iSystem if it is outside the acceptable control range.

• To troubleshoot control values that fall outside the control range, refer to the ARCHITECT System Operations Manual, Section 10.

• Minimum sample cup volume is calculated by the system and printed on the Orderlist report. To minimize the effects of evaporation, verify adequate sample cup volume is present prior to running the test. Maximum number of replicates sampled from the same sample cup: 10

• Priority:

Sample volume for first test: 75 μL

Sample volume for each additional test from same sample cup: 25 μL

• ≤ 3 hours on board:

Sample volume for first test: 150 μL

Sample volume for each additional test from same sample cup: 25 μL

• > 3 hours on board: Replace with a fresh sample (patient specimens, controls, and calibrators).

• If using primary or aliquot tubes, use the sample gauge to ensure sufficient patient specimen is present.

• Prepare ARCHITECT CORE Calibrator and Controls.

• Mix calibrator(s) and controls by gentle inversion before use.

• Hold bottles vertically and dispense recommended volumes into each respective sample cup.

• Recommended volumes:

for each calibrator: 5 drops

for each control: 4 drops

• Load samples.

• For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

• Press RUN.

• For additional information on principles of operation, refer to the ARCHITECT System Operations Manual, Section 3.

• For optimal performance, it is important to perform routine maintenance as described in the ARCHITECT System Operations Manual, Section 9. Perform maintenance more frequently when required by laboratory procedures.

#### Interpretation: reporting results

Calculation

• The ARCHITECT iSystem calculates the cutoff RLU from the mean RLU of three replicates of Calibrator 1 and stores the result. The cutoff RLU is determined by multiplying the CORE Calibrator 1 mean RLU by 1.0.

Cutoff RLU = Calibrator 1 Mean RLU x 1.0

• The ARCHITECT iSystem calculates the S/CO result for each specimen and control as follows:

S/CO = Sample RLU/Cutoff RLU

Interpretation of Results

| Initial ARCHITECT CORE Results |  |  |  |
| --- | --- | --- | --- |
| Initial Result (S/CO) | Instrument Flag | Interpretation | Retest Procedure |
| < 0.80 | NONREACTIVE | Nonreactive | No retest required. |
| 0.80 to < 1.21 | GRAYZONE | Grayzone | Retest same specimen in duplicate. |
| ≥ 1.21 | REACTIVE | Reactive | Retest same specimen in duplicate. |

| Final ARCHITECT CORE Interpretation |  |  |
| --- | --- | --- |
| Initial Interpretation | Results with Retest | Final Interpretation |
| Nonreactive | No retest required. | Nonreactive |
| Grayzone | If two of the three results are < 1.00 S/CO | Nonreactive |
|  | If two of the three results are ≥ 1.00 S/CO | Reactive |
| Reactive | If both retest results are < 1.00 S/CO | Nonreactive |
|  | If two of the three results are ≥ 1.00 S/CO | Reactive |

• A nonreactive final interpretation indicates that anti-HBc antibodies were not detected in the sample; it is possible that the individual is not infected with HBV.

• A reactive final interpretation indicates presumptive evidence of HBV; anti-HBc antibodies were detected in the sample which suggests either on-going or previous HBV infection.

Flags

Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

## Measuring Range:

N/A

Dilutions

Specimens cannot be diluted for the ARCHITECT CORE assay.

Precautions and Warnings

• For In Vitro Diagnostic Use

Safety Precautions

CAUTION: This product requires the handling of human specimens. It is recommended that all human-sourced materials be considered potentially infectious and handled in accordance with the OSHA Standard on Bloodborne Pathogens. Biosafety Level 2 or other appropriate biosafety practices should be used for materials that contain or are suspected of containing infectious agents.17-20

The following warnings and precautions apply to: Microparticles

WARNING 	Contains methylisothiazolone.

H317 		May cause an allergic skin reaction.

Prevention

P261 		Avoid breathing mist / vapors / spray.

P272 		Contaminated work clothing should not be allowed out of the workplace.

P280 		Wear protective gloves / protective clothing /eye protection.

Response

P302+P352	 IF ON SKIN: Wash with plenty of water.

P333+P313	 If skin irritation or rash occurs: Get medical advice / attention.

P362+P364 	Take off contaminated clothing and wash it before reuse.

Disposal

P501 		Dispose of contents / container in accordance with local regulations.

The following warnings and precautions apply to: Assay Diluent

DANGER: 	Contains polyethylene glycol octylphenyl ether (Triton X-405) and methylisothiazolone.

H318 		Causes serious eye damage.

H317 		May cause an allergic skin reaction.

H412 		Harmful to aquatic life with long lasting effects.

Prevention

P261 		Avoid breathing mist / vapors / spray.

P272 		Contaminated work clothing should not be allowed out of the workplace.

P280 		Wear protective gloves / protective clothing /eye protection.

P273 		Avoid release to the environment.

Response

P302+P352 	IF ON SKIN: Wash with plenty of water.

P305+P351+P338 IF IN EYES: Rinse cautiously with water for several minutes. Remove contact lenses, if

present and easy to do. Continue rinsing.

P310 		Immediately call a POISON CENTER or doctor / physician.

P333+P313	If skin irritation or rash occurs: Get medical advice / attention.

P362+P364 	Take off contaminated clothing and wash it before reuse.

Disposal

P501 		Dispose of contents / container in accordance with local regulations.

Microparticles, Conjugate and Assay Diluent contain sodium azide. Contact with acids liberates very toxic gas. This material and its container must be disposed of in accordance with local regulations.

Safety Data Sheets are available at www.abbottdiagnostics.com or contact your local representative.

For information on the safe disposal of sodium azide and a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8.

## Limitations — interference

LIMITATIONS OF THE PROCEDURE

• Current methods for the detection of anti-HBc antibodies may not detect all infected individuals. A nonreactive test result does not exclude the possibility of exposure to or infection with HBV.

• Specimens from patients who have received preparations of mouse monoclonal antibodies for diagnosis or therapy may contain human anti-mouse antibodies (HAMA). Such specimens may show either falsely elevated or depressed values when tested with assay kits such as ARCHITECT CORE that employ mouse monoclonal antibodies. Additional information may be required for diagnosis.23, 24

• Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays. Patients routinely exposed to animals or to animal serum products can be prone to this interference, and anomalous values may be observed. Additional information may be required for diagnosis.25

• Refer to the SPECIMEN COLLECTION AND PREPARATION FOR ANALYSIS section of this package insert for specimen limitations.

INTERFERNCE

At the concentrations listed below, bilirubin (conjugated and unconjugated), hemoglobin, total protein, and triglycerides showed less than 10% interference in the ARCHITECT CORE assay for high negative samples (S/CO range: 0.60 to 0.99) and low positive samples (S/CO range: 1.00 to 1.40):

• Bilirubin ≤ 20 mg/dL

• Hemoglobin ≤ 500 mg/dL

• Total Protein ≤ 12 g/dL

• Triglycerides ≤ 3000 mg/dL

Contacts:

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Alternative method

Send to sister facility

References

ABBOTT CORE package insert

Abbott Laboratories

Diagnostics Division

Abbott Park, Illinois 60064

G4-7984/R05

Revision 11/14

Effective date

Effective date for this procedure:

Author

Compiled by Abbott Diagnostics

Written by:  	 Cheryl Strauch, MLS(ASCP)