# PRINCIPLE:

The OSOM® hCG Combo Test is a rapid immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine or serum as an aid in the early determination of pregnancy. This test is for professional use only in physicians’ offices and clinical laboratories.

Human chorionic gonadotropin (hCG) is a glycoprotein hormone secreted by the developing placenta. After fertilization, the concentration of hCG rapidly rises in both the urine and serum of pregnant women. The detection of hCG in these fluids is an excellent marker for confirming pregnancy. The OSOM® hCG Combo Test is a rapid test which can detect the presence of hCG in urine or serum. The test utilizes monoclonal and polyclonal antibodies to hCG.

The OSOM® hCG Combo Test is a solid phase, sandwich-format  immunochromatographic assay for the qualitative detection of the hCG. Urine or serum is added to the sample well of the Test Device using the pipette provided. The sample migrates through reaction pads where hCG, if present in the sample, binds to a monoclonal anti-hCG dye conjugate. The sample then migrates across a membrane towards the results window, where the labeled hCG complex is captured at a test line region containing immobilized rabbit anti-hCG. Excess conjugate will flow past the test line region and be captured at a control line region containing an immobilized antibody directed against the anti-hCG dye conjugate (with or without hCG complexed to it.

SPECIMEN INFORMATION:

- SPECIMEN COLLECTION AND PREPARATION

No filtration or centrifugation of urine or serum specimens is required for testing with the OSOM ® hCG Combo Test.

Urine

- Urine specimens may be collected in any clean, dry, plastic, or glass containers.

- For early determination of pregnancy, the first morning specimen of urine is recommended since it is usually contains the highest concentration of hCG.

Serum

- Serum specimens should be obtained aseptically in tubes without anticoagulants.

- Plasma specimens are not suitable and should not be used for the OSOM® hCG Combo Test.

II. SPECIMEN STORAGE

Urine

- Urine specimens may be stored at room temperature 15° to 30° C (59° to 86° F) for up to 8 hours, or refrigerated at 2° to 8° C(35° to 46 ° F ) for up to 72 hours.

Serum

- Serum specimens may be stored at 2° to 8° C (35° to 46 ° F) for up to 48 hours before testing.

- If testing is delayed beyond 48 hours, the serum specimens (separated from the clot) should be frozen at -20°C (-4°F) or colder. Frozen specimens may be stored for up to 1 year. The frozen specimens should be thawed, mixed, and brought to room temperature 15° to 30° C (59° to 86° F) before testing.

REAGENTS AND SUPPLIES:

# I. REAGENTS

# Materials Provided

- Membrane coated with rabbit polyclonal anti-alpha hCG.

- Conjugated pad containing mouse monoclonal anti-beta hCG.

# Materials Required But Not Provided

# Clock or timer

- Sample collection cup or tubes.

- Positive and Negative Controls. OSOM® hCG Urine Control (catalog number 134) and OSOM® hCG Serum Control (catalog number 138).

II. PRECAUTIONS

- For in vitro diagnostic use.

- Do not use beyond the expiration date printed on the kit or foil pouch.

- The lot numbers may be different on the foil pouch and the kit.

- Use appropriate precautions for the collection, handling, and storage of specimens. All human blood products should be treated as potentially infectious and handled with good laboratory practices using appropriate precautions recommended in the Centers for Disease Control/national Institutes of Health manual,  “ Biosafety in Microbiological and Biomedical laboratories”. 1993.

- Dispose of all used Test Devices, pipettes and specimens in suitable biohazard waste containers.

- Test Devices are stable in the unopened foil pouches until the expiration date. Do not remove the Test Device from the pouch until needed.

# III. STORAGE AND STABILITY

- Store OSOM hCG Combo Tests at room temperature, 15° to 30° C (59° to 86°F), out of direct sunlight.

- Test Devices are stable until the expiration date printed on the kit or foil pouch.

- Do not freeze.

- If the control band does not appear when running the test, the Test Cassette or kit may have been stored or handled improperly or the foil pouch may not be intact.

# QUALITY CONTROL:

- Patient test results are only reported when controls perform as expected.

- Kit lot numbers, lot number of control, date opened, and expiration dates are logged in the BhCG Quality Control log book when the initial Quality Control is performed on the kit..

- External Quality Control is performed every 30 days for the same lot number.

- Quality control is run with each new kit lot number or shipment, with quality control reagents from kit in use (see parallel testing procedure in Quality control section).

External

External quality control will be performed with each new lot or shipment.

- Serum:  OSOM positive and negative serum control. (OSOM® hCG Serum Control (catalog number 138)).

- Urine: OSOM positive and negative urine control will be performed only if a urine BHCG is ordered within use of kit. (OSOM® hCG Urine Control (catalog number 134)).

Internal

Several procedural controls are incorporated into each OSOM hCG Combo Test for routine quality checks. These procedural controls are documented for each sample as part of daily quality control.

The same labeled conjugate antibody results in the appearance of both the test and the control bands. The appearance of the control band in the results window is an internal positive procedural control which validates the following:

- The appearance of the control band assures that the detection component of both the test line and control line is intact, that adequate sample volume was added and that adequate capillary migration of the sample has occurred. It also verifies proper assembly of the Test Device.

- The appearance of the control band indicates that an adequate volume of fluid was added to the sample well for capillary migration to occur. If the control band does not appear at the read time, the test is invalid. Appearance of the control band should be documented as part of the daily quality control band should be documented as part of the daily quality control.

- The clearing of the background in the results area may be documented as a negative procedural control. It also serves as an additional capillary flow control. After the read time, the background should appear white to light gray and not interfere with the reading of the test.  The test is invalid if the background fails to clear and obscures the observation of a distinct band.

- If the control band fails to appear with a repeat assay, do not report patient results. Contact Enzyme Diagnostics for technical service: telephone 617-252-7760 or 1-800-332-1042.

PROCEDURE:

# Patient specimens and control material must be brought to room temperature 15° to 30° C (59° to 86°F).

- Remove the Test Device and the pipette from the pouch. Place the Device on a flat surface.

- Squeeze the bulb of the pipette and insert the barrel into the patient sample. Release the bulb and draw up enough sample to fill the barrel to the line indicated on the pipette. Do not overfill.

- Expel the entire contents of the barrel (135 µL) into the sample well of the Test Device. No drop counting required,

- Discard the pipette in a suitable biohazardous waste container.

- Read results:

  - 3 minutes for urine.

  - 5 minutes for serum

  - Strong positive results may be observed sooner.

    - Results are invalid after the stated read time.

Procedural Notes:

- If specimen has been stored refrigerated, allow it to warm to room temperature before use.

- Several tests can be run at the same time. Use a new pipette with each test to avoid contamination errors.

# RESULTS:

# I. INTERPRETATION OF RESULTS:

# Negative Results

If no band appears at “T” and a black or gray band is visible at the “C” Control” position the test can be considered negative, indicating that a detectable level of the hCG is not present.

# Positive Results

Two separate black or gray bands-one at “T: Test” and the other at “ C: Control” – are visible in the results window, indicating that the specimen contains detectable levels of hCG. While the intensity of the test band may vary with different specimens, the appearance of two distinct bands should be interpreted as a positive result.

# Invalid Results

If no band appears at the “C: Control” position, the test is invalid. The test is also invalid if incomplete or beaded bands appear at either the “T: Test” or “C: Control.” The test should be repeated using another Test Device.

Note: the test is valid if the control line appears by the stated read time regardless of whether the sample has migrated all the way to the end of the sample window.

# II. EXPECTED VALUES

hCG is not normally detected in the urine and serum specimens of healthy men and non-pregnant women. In normal pregnancy, 20 mIU/mL hCG is reported to be present in both urine and serum 2 to 3 days before the first missed menstrual period. The level of hCG continue to increase up to 200,000 mIU/mL at the end of the first trimester.

III. REPORTING OF RESULTS:

All results are reported in Meditech (LIS system). All procedural (internal) controls are reported in for that run.

LIMITATIONS:

- This assay is capable of detecting only whole molecule ((intact) hCG. It cannot detect the presence of free hCG subunits. Therefore, this test should only be used for the qualitative detection of human chorionic gonadotropin in urine or serum for the early determination of pregnancy.

- For diagnostic purposes, hCG test results should always be used in conjunction with other methods and in the context of the patient’s clinical information (e. g., medical history, symptoms, results of other tests, clinical impressions, etc.). Ectopic pregnancy cannot be distinguished from normal pregnancy by hCG measurements alone.

- If the hCG is inconsistent with, or unsupported by, clinical evidence, results should also be confirmed by an alternate hCG method. If a serum specimen is initially tested qualitatively, alternate methods may include the quantitative testing of serum or the qualitative testing of urine. The absence of urinary hCG may suggest a falsely elevated serum result. Additionally, results may be confirmed by performing serial dilutions of the sample as usually, but not always, samples that contain interfering substances exhibit nonlinear results when diluted. Test results should be confirmed using quantitative hCG assay prior to the performance of any critical medical procedure.

# PERFORMANCE CHARACTERISTICS:

# I. Sensitivity & Specificity

- The OSOM® hCG Combo Test will detect hCG in urine specimens with concentrations of 20mIU/mL or more and serum specimens with 10mIU/mL or more.

- The expected sensitivity of urine samples at a read time of 3 minutes is

20 mIU/mL.

- The expected sensitivity of serum samples at a read time of 5 minutes is 10 mIU/mL

II. Interfering Substances

- Interfering substances may falsely depress or falsely elevate results. These interfering substances may cause false results over the entire range of the assay, not just at low levels, and may indicate the presence of hCG when there is none. As with any immunochemical reaction, unknown interferences from medications or endogenous substances may affect results.

- Infrequently, hCG levels may appear consistently elevated and could be due to, but not limited to, the presence of the following.

  - Heterophilic antibodies: Patients routinely exposed to animals or to animal serum products, can be prone to this interference and anomalous values may be observed.

  - Trophoblastic or nontrophoblastic neoplasms: Abnormal physiological states that may falsely elevate hCG levels. This test should not be used in the diagnosis of these conditions.

  - Nonspecific protein binding.

  - hCG substances.

    - Specimens from patients who have received preparations of Mouse Monoclonal

Antibodies for diagnosis or therapy may contain Human Anti-Mouse Antibodies

(HAMA). Such specimens may demonstrate either falsely elevated or falsely

depressed results when tested with assay kits which employ Mouse Monoclonal

Antibodies.

- The following potentially interfering substances were added to urine and serum specimens containing 0 or 20 mIU/mL(urine) or 10 mIU/mL (serum) hCG. The substances listed below where not found to affect the performance of the test.

| Urine/Serum |  |
| --- | --- |
| Acetominophen 20mg/L |  |
| Acetylsalicylic acid 20mg/dL |  |
| Amphetamines 10µl/mL |  |
| Ascorbic acid 20 mg/dL |  |
| Caffeine 20mg/dL |  |
| Cannabinol 10mg/dL |  |
| Codeine 10mg/dL |  |
| Diazapam 2mg/dL |  |
| Ephedrine 20mg/dL |  |
| Estradiol 25ng/dL |  |
| Ethanol 200mg/dL |  |
| Gentisic acid 20mg/dL |  |
| Glucose 2000mg/dL |  |
| Hemoglobin(urine) 250 mg/dL/(serum) 500mg/dL |  |
| Human Albumin 2000mg/dL |  |
| Ibuprofen 40mg/dL |  |
| Lithium 3.5 mg/dL |  |
| Methadone 10mg/dL |  |
| Morphine 6µl/mL |  |
| Phenobarbital 15mg/dL |  |
| Progesterone 40ng/mL |  |
| Proteins 2000mg/dL |  |
| Salicylic acid 20mg/dL |  |
| Tetracycline 20mg/dL |  |
| Triglycerides (serum) 2000g/dL |  |
| RF factor 40 IU/mL (serum) |  |

REFERENCES:

OSOM® hCG Combo Test Package Insert  12/07

OSOM® hCG Serum Control Set Package Insert,  11/2004

OSOM® hCG Urine Control Set Package Insert, 07/2001

Compiled by: Elsa Escobar (MT ASCP) 10/06,01/16

| SOP Number: | AS070 |  | 01/05/16 |
| --- | --- | --- | --- |
| Department: | SEROLOGY | Effective Date: | 01/05/16 |
| Author: | ELSA ESCOBAR MT(ASCP) | Version: | 1 |

| Applicable Standards |  |  |  |  | Version History |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Standard |  | Organization |  |  | Version | Effective Date |  | Deactivation Date |
|  |  |  |  |  | 1 | 01/05/16 |  |  |
|  |  | CAP |  |  |  |  |  |  |
| Related Documents |  |  |  |  |  |  |  |  |
| NEW PROCEDURE |  |  |  |  |  |  |  |  |
| CHANGE IN METHOD |  |  |  |  |  |  |  |  |
| Review History (Up to the Last 15 Occurrences) |  |  |  |  |  |  |  |  |
| Date | Version |  | Revision Type |  |  |  | Review By |  |
| 01/05/16 | 1 |  | New Policy/Procedure |  |  |  | DR.RANDALL SIMONSEN |  |
|  |  |  | Director Review |  |  |  |  |  |

Distribution

SEROLOGY SOP CSH-ALICE