##### Principle

The instrument is designed to accurately and precisely measure the sedimentation rate of erythrocytes in ESR-Vacuum Tubes. The results are recorded as mm per hour (Westergren method) by converting to the QuickMode Method, a scientifically developed method for measuring ESR in only 30 minutes. The instrument compensates for temperatures above 26C (78.8F).

Red cell sedimentation is accelerated by an increase in the plasma concentration of “acute phase proteins,” which are increased in acute tissue damage, chronic inflammation, chronic infection, and pregnancy. The ESR reflects both the increase in certain accelerating proteins, such as fibrinogen and gamma globulins, and the decrease in retarding proteins, such as albumin.

##### Specimen Collection

- The ESR-Vacuum Tubes, 1.2ml or 2.0ml, contain 3.2% Sodium Citrate as an anticoagulant.

- ESR-Vacuum Tubes preserve the integrity of the patient sample for up to 72 hours from the time of blood collection when transported or stored at 2 to 10C, or up to 4 hours at 18-30° C.

- Patient samples may be drawn directly into an ESR-Vacuum Tube or transferred from an EDTA tube into an ESR-Vacuum Tube. When transferring from an EDTA tube, be sure the sample is well-mixed and fill the ESR-Vacuum Tube to the indicated mark on the tube. Do not remove the sodium citrate in the ESR-Vacuum Tube when testing control or patient sample. Refer to CLSI guidelines for EDTA tube stability.

##### Calibration

It is recommended to run a system check using the test rack every day the ESR-Auto Plus is in use.

To run a test rack:

1.  From Standby (TIME/NEXT SAMPLE on the display), press and hold X until you hear a beep to access the user menus.

2.  Enter 5, RUN TEST RACK. You will be prompted:

SCANNER TEST: INSERT RACK [OK]

3.  Insert the test rack and press  to start the verification procedure.

4.  After the test rack reading is complete, the results will print on the internal printer.

5.  If the printout indicates SCANNER TEST OK, continue to the QC procedure.

6.  If the printout indicates FAILED TEST, the instrument will ask if you want to retry.

7.  Press  to retry. Press X to exit. If the test rack fails a second time, re-register the test rack. Refer to the ESR-Auto Plus Operator’s Manual for instructions, or contact Streck Technical Services at 800-843-0912 (ext. 7510).

Quality Control

- ESR-Chex is an assayed bi-level control for evaluating the accuracy and precision of automated and manual Erythrocyte Sedimentation Rate (ESR) methods.

- ESR-Chex is designed to be handled in the same manner as a patient specimen. There are several factors concerning the environment and techniques that may affect the outcome of erythrocyte sedimentation rate

- methods. When these are properly controlled, the values recovered with ESR-Chex will be within the Expected Ranges printed on the assay. Accurate analysis of patient specimens is assured.

REAGENTS

ESR-Chex contains stabilized human red blood cells in a preservative medium.

STORAGE AND STABILITY

- ESR-Chex is stable through the expiration date when stored at 2 °C-10 °C. After opening, ESR-Chex is stable throughout the open-vial dating, as indicated on the assay sheet, when stored at 2 °C-30 °C.

- Discoloration of the product may be caused by overheating or freezing during shipping or storage. Gross hemolysis (darkly colored supernatant) may be indicative of product deterioration. However, moderately colored supernatant is normal and should not be confused with deterioration of the product.

- It is good laboratory practice to prepare a fresh sample of ESR-Chex daily to control the ESR procedure.

HANDLING INSTRUCTIONS

1. Remove vials from refrigerator and allow them to equilibrate to room temperature (20-30 minutes).

2. Vortex unopened vials for up to 60 seconds prior to first time use. Unopened vials stored at 2 °C-10 °C may

require extra mixing. Mix previously opened vials through inversion and by vigorously rolling upright between

palms until red cells are completely suspended. Continue to mix for 90 seconds.

3. Allow bubbles to disperse and re-mix by inversion prior to sampling.

4. Draw the sample immediately after thorough mixing is completed.

If mixed vials sit for more than 1 minute before drawing the sample, the vial must be remixed by

repeating steps 2 and 3. Incomplete mixing can invalidate both the sample drawn and the remaining

product in the vial.

5. Follow the manufacturer’s directions for filling the sedimentation rate tube and setup for the method in use.

6. Wipe threads of vial and cap with clean tissue before closing. Recap the vial tightly.

7. Store open vials at room temperature (18 °C-30 °C). Opened ESR-Chex vials are stable throughout the open-vial dating, as indicated on the assay sheet.

EXPECTED RESULTS

- Upon receipt of a new control lot the laboratory will establish its own mean and limits. Per CLSI H26-A21 it is recommended that each level of control be run twice a day for 3-5 days to establish individual lab means for each measurand.  However, the control means established by the laboratory should fall within the Expected Range specified for the control.

- In order to achieve values within the published assay ranges, laboratories should prepare a fresh sample of ESR-Chex daily. Failure to adhere to this guideline may result in values outside the published expected ranges.

- It is advised that laboratories establish their own means and ranges when utilizing control material outside of these handling instructions.

- If established means do not fall within the Expected Range, please evaluate all mechanical and physical

- factors that could affect the outcome, such as handling, temperature, vibration, tube position and product expiration. Inadequate mixing of the product will affect the outcome of the test, see Handling Instructions.

- Sample must be taken within 60 seconds of thoroughly mixing the vial. Assay an unopened vial of ESR-Chex.

- If the values are still outside the Expected Range, contact Streck Technical Services at 800-843-0912 or

- technicalservices@streck.com.

To enter a new control lot number in the QC file:

NOTE – Registering a new control lot will delete all data for the previous lot of that level.

1.  From Standby (TIME/NEXT SAMPLE on the display), press and hold X until you hear a beep to access the user menus.

2.  Enter 4, CONTROL SAMPLES.

3.  Enter 3, REGISTER NEW CONTROL.

4.  Enter Level 1 or Level 2.  Press .  (This action will delete old statistics for this control level.)

5.  Press  to clear old statistics and register a new control lot. Press X to exit.

6.  Enter the lot number of the control. Press and hold a numeric key to enter the corresponding letters A-Z.   Press  when finished.

7.  Enter the expiration date of the control. Press .

8.  Enter the minimum value of the control. Press .

9.  Enter the maximum value of the control. Press .

10. A report of the assay values will be generated on the printer.  Press  to accept the data. Press X to restart the control registration.

To run controls:

1.  From Standby (TIME/NEXT SAMPLE on the display), press and hold X, until you hear a beep to access the user menus.

2.  Enter 4, CONTROL SAMPLES.

3.  Enter 1, RUN CONTROL.

4.  Enter Level 1 or Level 2. Press . The program will automatically name this sample by its lot number followed by a sequence number. Example: The lot number is 1234A. The automatic ID number of the first sample will be 1234A/01. The ID number of the second sample will be 1234A/02, then 1234A/03, etc.

5.  Transfer the well-mixed control material into an ESR-Vacuum Tube.

6.  When prompted, insert the well-mixed sample into any free position in the ESR-Auto Plus. The position light will turn red and the ESR-Auto Plus will initiate testing.

7.  Control results will automatically print when the test is complete. If the result is outside the expected range, an alarm message will be reported on the printout.

NOTE – Controls must be run from the Control menu (not as patient samples) to be stored in the QC log files.

NOTE – Each QC log file stores up to 100 samples per level. Statistical reports can be generated from these log files.  When the QC log reaches capacity, the oldest value will be deleted and the newest value will be added to the QC log file.

To show sample data:

1.  From Standby (TIME/NEXT SAMPLE on the display), press and hold X until you hear a beep to access the user menus.

2.  Enter 4, CONTROL SAMPLES.

3.  Enter 2, SHOW CONTROLS.

a.  Enter 1, SHORT REPORT to generate a summary report.

b.  Enter 2, LONG REPORT to generate a summary report and QC log file printout.

c.  Enter 3, SHOW ON LCD to view individual data points on the display.

d.  Enter 4, CLEAR QC LOG FILE to delete all QC data points.

Running Patient Samples

An ID code is required to run a sample. An ID code may contain up to 13 digits.

Scan ID with the barcode scanner:

1.  From Standby (TIME/NEXT SAMPLE on the display), scan the patient barcode with the barcode scanner.

2.  When prompted, insert the well-mixed patient sample into any free (green) position in the ESR-Auto Plus. The position light will turn red and the ESR-Auto Plus will initiate testing.

Manually enter ID on keypad:

1.  From Standby (TIME/NEXT SAMPLE on the display), enter the patient ID number on the keypad. Use the X key to delete any digits entered incorrectly. Press .

2.  When prompted, insert the well-mixed* patient sample into any free (green) position in the ESR-Auto Plus. The position light will turn red and the ESR-Auto Plus will initiate testing.

##### Result Reporting

- It is good laboratory practice to visually correlate the level of sedimentation in the sample tube to the printed result. The ESR-Auto Plus automatically prints the result on the internal printer and data output port. After measurement is complete, any error codes will be included on the result.

- NOTE – See the ESR-Auto Plus Operator’s Manual for additional software options such as searching for a stored patient result or accessing the prediction mode.

##### Reportable Range

Reportable range:  0 - 120 mm/hr. Sample with a result greater than 120 mm/hr should be reported as “>120 mm/hr.”

##### Limitations of the Procedure

The ESR-Auto Plus uses the QuickMode method which reads the RBC sedimentation level after 30 minutes and

mathematically converts it to an equivalent Modified Westergren result in mm/hr.

Measuring Abnormal Samples:

- The ESR-Auto Plus reads samples by scanning the tube with infrared light (IR). One side of the measuring arm consists of IR-diodes and the other has IR sensors.

- Important! The clinical significance of an ESR result obtained from an abnormal sample, including but not limited to icteric, lipemic, cold agglutinins and other anemias, should be determined by the clinician ordering the test.

Hemolyzed Samples

The IR light is powerful and can transmit through plasma, lipid samples and most hemolyzed samples.

- However, results from hemolyzed samples can cause confusion.  In manual testing, it can be very difficult to see the RBC /plasma interface in a sample that is only partly hemolyzed.

- The ESR-Auto Plus can detect hemolysis to a level of ≤ 3.0 g/dL free hemoglobin due to the strength of the IR light. Try to reproduce the test result manually by holding the hemolyzed sample against a strong light. Re-run the sample on the ESR-Auto Plus and verify the automated result.

Icteric Samples

The ESR-Auto Plus indicates no appreciable shift in ESR recovery at bilirubin concentrations up to 20.2 mg/dL.

Infrared detection is not impeded.

Low Hemoglobin Concentration

The ESR-Auto Plus can read samples with hemoglobin concentrations ≥3.0 g/dL.

Lipemic Samples

The ESR-Auto Plus is unable to read severely lipemic samples. The readability limit of lipemic samples has not

been quantified. The ESR-Auto Plus detects the cellular/plasma interface when reporting the ESR result. That

interface is disguised in severely lipemic samples.

##### Preventive Maintenance

1.  Regularly verify that the print head on the internal printer is free from dust.

2.  Regularly clean the instrument using a soft, damp cloth or paper. The covers may be carefully cleaned with ethanol or isopropanol. Always remove the power cord from the ESR-Auto Plus before cleaning with any liquid. Do not open or remove the ESR-Auto Plus cover during cleaning.

##### References

- ESR-Auto Plus Operator’s Manual, Streck.

- CLSI. Procedures for the Erythrocyte Sedimentation Rate Test; Approved Standard. CLSI document H02.

Compiled by:  Barbara Herro

Revised by:  Elsa Escobar

Well-Mixed:  Follow control manufacturer‘s mixing procedure for QC material or a minimum of 3 minutes on the ESR-657 Mixer for patient samples.

Well-Mixed:  Follow control manufacturer‘s mixing procedure or a minimum of 3 minutes on the ESR-657 Mixer.

Well-Mixed:  Follow laboratory mixing procedure or a minimum of 3 minutes on the ESR-657 Mixer.