Intended use

The ARCHITECT Total β-hCG assay is a Chemiluminescent Microparticle Immunoassay (CMIA) for the quantitative and qualitative determination of beta human chorionic gonadotropin (β-hCG) in human serum and plasma for the early detection of pregnancy.

Summary

Human chorionic gonadotropin (hCG) is a sialoglycoprotein with a molecular weight of approximately 46,000 daltons.1 HCG is initially secreted by the trophoblastic cells of the placenta shortly after implantation of the fertilized ovum into the uterine wall.2,3 The rapid rise in hCG serum levels after conception makes it an excellent marker for early confirmation and monitoring of pregnancy.

Physiologically, hCG appears to maintain the corpus luteum, thereby allowing synthesis of progesterone and estrogens that support the endometrium. As uncomplicated pregnancies progress, the placenta assumes the production of these hormones. The serum hCG levels increase to a peak concentration, then decrease and plateau. HCG circulates as the intact molecule in the serum of normal women who have an uncomplicated pregnancy. The subunits are cleaved rapidly and cleared by the kidney.4

The placental hormone, hCG, is similar to luteinizing hormone (LH), follicle stimulating hormone (FSH), and thyroid stimulating hormone (TSH). All are glycoproteins consisting of two noncovalently bound dissimilar subunits, designated alpha and beta, with attached carbohydrate sidechains. The alpha subunits of these glycoproteins are very similar. In contrast, the beta subunit portions determine the biological and immunochemical specificities.5,6 The beta subunits of hCG and LH exhibit considerable homology in amino acid content. Amino acid residues specific for the beta subunit of hCG confer the immunochemical specificity. 7

With the availability of sensitive quantitative assays for the measurement of serum β-hCG, it has been shown that hCG levels can be useful in prediction of spontaneous abortions,8,9 aiding in the detection of ectopic pregnancy 8,10,11 and multiple gestation.8

## Method

The ARCHITECT Total β-hCG assay is a two-step immunoassay to determine the presence of β-hCG in human serum and plasma using Chemiluminescent Microparticle Immunoassay (CMIA) technology with flexible assay protocols, referred to as Chemiflex.

## Principle

In the first step, sample, and anti-β-hCG coated paramagnetic microparticles are combined. β-hCG present in the sample binds to the anti-β-hCG coated microparticles. After washing, anti-β-hCG acridinium-labeled conjugate is added in the second step. Pre-Trigger and Trigger Solutions are then added to the reaction mixture; the resulting chemiluminescent reaction is measured as relative light units (RLUs). A direct relationship exists between the amount of β-hCG in the sample and the RLUs detected by the ARCHITECT i optical system.

## Specimen collection and handling

The specimen collection tubes listed below were verified to be used with the ARCHITECT Total β-hCG Assay. Other specimen collection tubes have not been tested with this assay.

• Human serum (including serum collected in serum separator tubes)

• Human plasma collected in:

• Lithium heparin (powder or gel)

• Sodium heparin

• Potassium EDTA

• Sodium citrate plasma tubes cannot be used with the ARCHITECT Total β-hCG Assay..

• Liquid anticoagulants may also have a dilution effect resulting in lower

concentrations for individual patient specimens.

• The ARCHITECT i System does not provide the capability to verifyspecimen type. It is the responsibility of the operator to verify that thecorrect specimen types are used in the ARCHITECT Total β-hCG Assay.

Specimen Conditions

• Do not use specimens with the following conditions:

• heat-inactivated

• pooled

• grossly hemolyzed (> 500 mg/dL hemoglobin)

• obvious microbial contamination

• cadaver specimens or any other body fluids

• For accurate results, serum and plasma specimens should be free of fibrin, red blood cells, and other particulate matter. Serum specimens from patients receiving anticoagulant or thrombolytic therapy may

• Use caution when handling patient specimens to prevent cross contamination. Use of disposable pipettes or pipette tips is recommended.

• For optimal results, inspect all specimens for bubbles. Remove bubbles with an applicator stick before analysis. Use a new applicator stick for each specimen to prevent cross contamination.

Preparation for Analysis

• Follow the tube manufacturer’s processing instructions for serum and plasma collection tubes. Gravity separation is not sufficient forspecimen preparation.

• Mix thawed specimens thoroughly by low speed vortexing or by inverting 10 times. Visually inspect the specimens. If layering or stratification is observed, continue mixing until specimens are visibly homogeneous.

• To ensure consistency in results, specimens must be transferred to a centrifuge tube and centrifuged at ≥ 10000 RCF (Relative Centrifugal Force) for 10 minutes before testing if they contain fibrin, red blood cells, or other particulate matter,• they require repeat testing, or they were frozen and thawed.Transfer clarified specimen to a sample cup or secondary tube for  testing.

• Centrifuged specimens with a lipid layer on the top must be transferred to a sample cup or secondary tube. Care must be taken to transfer only the clarified specimen without the lipemic material.

Storage

• Specimens may be stored on or off the clot or red blood cells for:

• up to 24 hours at 15-30°C

• up to 7 days at 2-8°C

• If testing will be delayed more than 24 hours, remove serum or plasma from the clot or red blood cells and store refrigerated at 2-8°C.

• If testing will be delayed more than 7 days, remove serum or plasma from the clot or red blood cells and store frozen at -10°C or colder for up to 12 months.

.

Shipping

• Before shipping specimens, it is recommended that specimens be removed from the clot, red blood cells, or separator gel.

• When shipping specimens, package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.

• Specimens may be shipped: ambient, on wet ice, on dry ice. Do not exceed the storage time limitations listed above.

Materials and Equipment Required

Provided

- 7K78 ARCHITECT Total β-hCG Reagent Kit

Materials Required but not Provided

- ARCHITECT i System

ARCHITECT i

- 7K78-01 ARCHITECT Total β-hCG Calibrators

- 7K78-10 ARCHITECT Total β-hCG Controls

- 7D82-50 ARCHITECT I Multi-Assay Manual Diluent

ARCHITECT

- ARCHITECT

- ARCHITECT

- ARCHITECT

- ARCHITECT

- ARCHITECT

- ARCHITECT

For information on materials required for maintenance procedures,refer to the ARCHITECT System Operations Manual, Section 9.

• Pipettes or pipette tips (optional) to deliver the volumes specified on

## Reagents – working solutions

Reagent Kit, 100 Tests/500 Tests

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT i Systems. Please contact your local distributor.

ARCHITECT  Total β-hCG Reagent Kit (7K78)

- Microparticles:  1 Bottle of 4 bottles (6.6 mL per 100-test bottle/27.0 mL per 500-test bottle) Anti-β-hCG (mouse monoclonal) coated microparticles in TRIS buffer with protein(bovine) stabilizers. Preservatives: antimicrobial agents.

- Conjugate 1 Bottle of 4 bottles (2.1 mL per 100-test bottle/7.4 mL per 500-test bottle) Anti-β-hCG (mouse monoclonal) acridinium-labeled Conjugate in MES buffer with protein (bovine) stabilizers. Minimun concenrations: 2.9 µg/mL. Preservatives: antimicrobial agents.

- Assay Diluent:  1 Bottle (100mL) ARCHITECT I Multi-Assay Manual Diluent containing phosphate buffered saline solution. Peservatives: antimicrobial agents.

Other Reagents

ARCHITECT i Pre-Trigger Solution

Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT i Trigger Solution

Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT i Wash Buffer

Wash Buffer containing phosphate buffered saline solution. Preservative: antimicrobial agents.

## Storage and Stability

•The ARCHITECT Total β-hCG Reagent Kit must be stored at 2-8°C in an upright position and may be used immediately after removal from 2-8°C storage.

• When stored and handled as directed, reagents are stable until the expiration date.

• The ARCHITECT Total β-hCG Reagent Kit may be stored on board the ARCHITECT i System for a maximum of 30 days.

• After these maximum storage times, the reagent kit must be discarded. For information on tracking onboard time, refer to the ARCHITECT System Operations Manual, Section 5.

• Reagents may be stored on or off the ARCHITECT i System. If reagents are removed from the system, store them at 2-8°C (with septums and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded. For information on unloading reagents, refer to the ARCHITECT System Operations Manual, Section 5.

Indications of Reagent Deterioration

When a control value is out of the specified range, it may indicate deterioration of the reagents or errors in technique. Associated test results are invalid and samples must be retested. Assay recalibration may be necessary. For troubleshooting information, refer to the ARCHITECT System Operations Manual, Section 10.

## Calibration

• The ARCHITECT Total β-hCG assay requires calibration:

- Installation

- New lot of reagent

- Controls are out of range

• To perform an ARCHITECT Total β-hCG calibration, test calibrators A, B, C, D, E, and F in replicates of two. A single sample of each Total β-hCG control level must be tested to evaluate the assay calibration. Ensure that assay control values are within the concentration ranges specified in the control package insert. Calibrators should be priority loaded.

• Calibration Range: 0.0 to 15,000.00 MIU/mL.

• For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Quality control

The recommended control requirement for the ARCHITECT Total β-hCG assay is that a single sample of each control level be tested once every 24 hours each day of use.

The Laboratory will run 3 levels of  Bio-Rad Immunoassay Plus controls once every 24 hours.

The ARCHITECT STAT Total β-hCG Assay control values must be within the acceptable ranges specified in the control package insert. If a control is out of the specified range, the associated test results are invalid and samples must be retested. Recalibration may be indicated.

For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Preparation of Working Solutions

Handling Precautions

•   Do not use reagent kits beyond the expiration date.

•   Do not pool reagents within a kit or between reagent lots.

•  Prior to loading the ARCHITECT Total β-hCG Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this package insert.

•   Septums MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

•   To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

•   Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

- Over time, residual liquids may dry on the septum surface. These are typically dried salts, which have no effect on assay efficacy.

- For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay

INSTRUMENT PROCEDURE

1.  The ARCHITECT Total β-hCG assay file must be installed on the ARCHITECT i System with the ARCHITECT i Assay CD-ROM prior to performing the assay.  For detailed information on assay file installation and on viewing and editing assay parameters, refer to the ARCHITECT System Operations Manual, Section 2.

2.  For information of printing assay parameters, refer to the ARCHITECT System Operations Manual, Section 5.

3.  For a detailed description of system procedures, refer to the ARCHITECT System Operations Manual.

- The default result unit for the ARCHITECT Total β-hCG is mIU/mL.  An alternate result unit, IU/L, may be selected for reporting results by editing assay parameter "Result concentration .   conversion factor used by the system is 1.0.”

- Conversion Formula: (Concentration in mIU/mL)x 1.0= IU/L

ASSAY PROCEDURE

1.  Before loading the ARCHITECT Total β-hCG Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment:

a.  Invert the microparticle bottle 30 times.

b.  Visually inspect the bottle to ensure microparticles are resuspended.  If microparticles are still adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

c.  If the microparticles do not resuspend, DO NOT USE.  Contact your local Abbott Laboratories representative.

d.  Once the microparticles have been resuspended, remove and discard the cap.  Wearing clean gloves, remove a septum from the bag.  Squeeze the septum in half to confirm that the slits are open.  Carefully snap the septum onto the top of the bottle.

2.  Order calibration, if necessary

a.  For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

3.  Order tests.

a.  For information on ordering patient specimens and controls, refer to the ARCHITECT System Operations Manual, Section 5.

- Load the ARCHITECT Total β-hCG Reagent Kit on the ARCHITECT i System

  - Verify that all necessary assay reagents are present.  Ensure that septa are present on all reagent bottles.

- The minimum sample volume is calculated by the system and is printed on the Orderlist report.  No more than 10 replicates may be sampled from the same sample cup.  To minimize the effects of evaporation, verify adequate sample cup volume is present prior to running the test.

a.  Priority:  75 uL for the first ARCHITECT Total β-hCG test plus 25 uL for each additional ARCHITECT Total β-hCG test from the same sample cup.

b.  </= 3 hours on board:  150 uL for the first ARCHITECT Total β-hCG test plus 25 uL for each additional ARCHITECT 25-OH Vitamin D test from the same sample cup.

c.  If using primary or aliquot tubes, use the sample gauge to ensure sufficient patient specimen is present.

6.  Prepare calibrators and controls.

a.  ARCHITECT Total β-hCG Calibrators and Controls should be mixed according to instructions in their respective package inserts.

b.  To obtain the recommended volume requirements for the ARCHITECT Total β-hCG Calibrators, hold the bottles vertically and dispense 4 drops of each calibrator or  3 drops of control into each respective sample cup.

7.  Load samples.

- For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

8.  Press RUN.  The system performs the following functions:

a.  Moves the sample to the aspiration point.

b.  Loads a reaction vessel (RV) into the process path.

c.  Aspirates and transfers sample into the RV.

d.  Advances the RV one position and transfers assay diluent and microparticles into the         RV.

e.  Mixes, incubates, and washes the reaction mixture.

f.  Adds conjugate to the RV.

g.  Mixes, incubates, and washes the reaction mixture.

h.  Adds pre-trigger and trigger solutions.

i.  Measures chemiluminescent emission to determine the quantity of troponin-I in the sample.

j.  Aspirates contents of RV to liquid waste and unloads RV to solid waste.

k.  Calculates the result.

9.  For optimal performance, it is important to follow the routine maintenance procedures defined in the ARCHITECT System Operations Manual, Section 9.  If your laboratory requires more frequent maintenance, follow those procedures.

#### Interpretation: reporting results

Calculation

- The ARCHITECT Total β-hCG assay utilizes a 4 Parameter Logistic Curve Fit (4PLC,Y weighted) data reduction method to generate a calibration curve.

Flags

- Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

For qualitative interpretation of the ARCHITECT Total β-hCG test results, specimens with β-hCG levels less than or equal to 5.00 mIU/mL will be reported in the INTERPRETATION field on the test results screen or printout as “NEGATIVE”. Specimens with β-hCG levels greater than or equal to 25.00 mIU/mL will be reported as “POSITIVE”. Specimens with β-hCG levels greater than 5.00 mIU/mL and less than 25.00 mIU/mL will be reported with concentrations only. No interpretation will be reported

for these results.

Expected Values:

Negative: ≤5.00 mIU/mL

Inconclusive: 5.01-24.99 mIU/L

Positive:  ≥25.00 mIU/mL

.

CHRISTUS Spohn Hospital has investigated the transferability of the expected values to its own patient population and determined its own reference range.  For diagnostic purposes, the test frindings should always be assessed in conjunction with the patient’s medical history, clinical examinations, and other findings.

Critical Values: None

## Measuring Range

- The measurement range for the ARCHITECT Total β-hCG assay is 0.0 mIU/mL to 15,000.00 mIU/mL.

## Dilutions:

Specimens can be diluted for the ARCHITECT Total β-hCG assay. Specimens with a Total β-hCG concentration of > 15,000.0 mIU/mL  will be flagged as “> 15,000.0 ng/mL” and should be reported as such. A 1:15 dilution is recommended not to exceed 1:75 dilution.

A result of <7000.0 mIU/mL (1:15 automatic dilution) indicates the need to retest the sample at a lower dilution. The result and interpretation should not be repoted until a sufficient dilution has been reached.

Precautions and Warnings

1.  For In Vitro Diagnostic Use.

This product requires the handling of human specimens.  It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens.  Biosafety Level 2 or other appropriate biosafety practices should be used for materials that contain or are suspected of containing infectious agents.

The following warnings and precautions apply to these components:

• Microparticles

• Assay Diluent

WARNING: Contains methylisothiazolones.

H317	 May cause an allergic skin reaction.

Prevention

P261	 Avoid breathing mist/vapours/spray.

P272 	Contaminated work clothing should not be allowed out of the workplace.

P280 	Wear protective gloves/protective clothing/eyeprotection.

Response

P302+P352 IF ON SKIN: 	Wash with plenty of water.

P333+P313 	If skin irritation or rash occurs: Get medical advice/attention.

P363 		Wash contaminated clothing before reuse

• For information on a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8.

HANDLING PRECAUTIONS

• Do not use reagent kits beyond the expiration date.

• Do not pool reagents within a kit or between reagent kits

.

• The conjugate is temperature sensitive, minimize room temperature

exposure of conjugate bottle.

• Before loading the ARCHITECT Total β-hCG Reagent Kit on

the system for the first time, the microparticle bottle requires mixing

to resuspend microparticles that have settled during shipment. For

microparticle mixing instructions, refer to the PROCEDURE, Assay

Procedure section of this package insert.

• Septums MUST be used to prevent reagent evaporation and

contamination and to ensure reagent integrity. Reliability of assay

results cannot be guaranteed if septums are not used according to

the instructions in this package insert.

• To avoid contamination, wear clean gloves when placing a septum on

an uncapped reagent bottle.

• Once a septum has been placed on the reagent bottle, do not

invert the bottle as this will result in reagent leakage and may

compromise assay results.

• Over time, residual liquids may dry on the septum surface. These

are typically dried salts, which have no effect on assay efficacy.

• For a detailed discussion of handling precautions during system

operation, refer to the ARCHITECT System Operations Manual,

Section 7.

## Limitations — interference

• This assay is capable of detecting whole molecule (intact) hCG as well as free β-hCG subunits.

• For diagnostic purposes, hCG results should always be used in conjunction with other data; e.g., patient’s medical history, symptoms, results of other tests, clinical impressions, etc. Ectopic pregnancy cannot be distinguished from normal pregnancy by hCG measurements alone.10,11 The results from this or any other diagnostic kit should be used and interpreted only in the context of the overall clinical picture.

• If the hCG level is inconsistent with, or unsupported by, clinical evidence, results should be confirmed by an alternate hCG method.

This may include the qualitative testing of urine.29 The absence of urinary hCG may suggest a falsely elevated serum result. Results

may also be confirmed by performing serial dilutions of the sample. Usually, but not always, samples that contain interfering substances exhibit nonlinear results when diluted.

•The ARCHITECT Total  β-hCG assay is cleared for use in the earlydetection of pregnancy only. It is not approved for any other uses such as tumor marker screening, tumor marker monitoring, etc. and it should not be performed for any other uses.

• Infrequently, hCG levels may appear consistently elevated and could be due to, but not limited to, the presence of the following:20,21,28,29,31,32

- heterophilic antibodies

- nonspecific protein binding

- hCG-like substances

- trophoblastic or nontrophoblastic neoplasms

• As with any immunochemical reaction, unknown interferences from medications or endogenous substances may affect results.

• Elevated hCG levels have been associated with some abnormal physiological states (e.g., trophoblastic and nontrophoblastic

neoplasms)16,17 and the results of this test should not be used in the diagnosis of these abnormal states. There have been reports

of people receiving unnecessary medical treatment and surgery, including chemotherapy and hysterectomy, when hCG results were

used in the diagnosis of abnormal states.

• Interfering substances (such as heterophilic antibodies, non-specific proteins, or hCG-like substances) may falsely depress or falsely elevate results. These interfering substances may cause false results over the entire range of the assay, not just at low levels, and may indicate the presence of hCG when there is none.

• Detection of very low levels of hCG does not rule out pregnancy.21 Low levels of hCG can occur in apparently healthy, nonpregnant subjects.26,27 Because hCG values double approximately every 48 hours in a normal pregnancy,21 patients with very low levels of hCG should be resampled and retested after 48 hours.

• Post menopausal specimens may elicit weak positive results due to low hCG levels unrelated to pregnancy. With a weak positive result, it is good laboratory practice to resample and retest after 48 hours, or to test with an alternate hCG method.

• Because of the high degree of sensitivity of the assay, specimens tested as positive during initial days after conception may later

be negative due to natural termination of the pregnancy. Natural termination occurs in 22% of clinically unrecognized pregnancies and 31% of pregnancies overall.30 It is good laboratory practice to resample and retest weak positive results after an additional 48 hours.

• Specimens from patients who have received preparations of Mouse Monoclonal Antibodies for diagnosis or therapy may contain Human Anti-Mouse Antibodies (HAMA). Such specimens may demonstrate either falsely elevated or falsely depressed results when tested with assay kits which employ Mouse Monoclonal Antibodies.18,19 These specimens should not be tested with the Abbott ARCHITECT Total β-hCG assay.

• Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.20,21,32

Patients routinely exposed to animals or to animal serum products, can be prone to this interference and anomalous values may be

observed. Additional information may be required for diagnosis.

____________________________________________________________________________________

Alternative method

Send to Sister Facility

References

ABBOTT Total β-hCG package insert-- Package insert 11/2011

Abbott Laboratories

Diagnostics Division

Abbott Park, Illinois 60064

PN 840174/R3

Revision 07/04

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Effective date

Effective date for this procedure: 01/15/18

Author

Compiled byAbbott Diagnostics

Revised by:  Jessica Mercer - Lead Tech BS, MT (ACSP)

Revised by:Elsa Escobar MT (ASCP)