_____________________________________________________________________________________

Intended use

The ARCHITECT BNP assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of human B-type natriuretic peptide (BNP) in human EDTA plasma on the ARCHITECT i System. BNP values are used as an aid in the diagnosis and assessment of severity of heart failure.

Summary

Heart failure is a syndrome caused by a variety of conditions such as coronary artery disease, hypertension, valve disease, myocarditis and others. Common symptoms of heart failure include shortness of breath, coughing under exertion, swelling of appendages, and dizziness. Heart failure is better defined as the progressive inability of the heart ventricles to pump blood out to the lungs and/or the extremities. Heart failure is either systolic or diastolic or a combination of both. Severity is usually classified into four classes defined by the New York Heart Association (NYHA class I-IV).

BNP is one member of the family of natriuretic peptides that were initially discovered by de Bold, et al.1 Although BNP was first isolated from porcine brain tissue (originally named brain natriuretic peptide),2 the heart has been determined to be the major source.3 BNP is synthesized and released into the blood in response to volume overload or conditions that cause ventricular stretch, to control fluid and electrolyte homeostasis by interaction with the renin-angiotensin-aldosterone system (RAAS).3,4 PreproBNP (134 amino acids) is synthesized in the cardiac myocytes and is processed to a proBNP (108 amino acids) precursor molecule. The proBNP is subsequently cleaved into physiologically active BNP (32 amino acids), and a degradation fragment NT-proBNP (76 amino acids).5,6 BNP, NT-proBNP, and a higher molecular weight form have been detected in peripheral blood.7,8 BNP is cleared from the circulation, with a half-life (t1/2) of approximately 23 minutes, by specific cellular receptors and neutral endopeptidases.9 Numerous studies have indicated that BNP can be used for patient diagnosis, prognosis and therapy monitoring. Levels of BNP have been shown to be elevated in patients with cardiac dysfunction.10,11 Plasma BNP levels provide clinically useful information concerning the diagnosis and management of left ventricular dysfunction and heart failure, which complements other diagnostic testing procedures (e.g., electrocardiograms, chest x-rays, and echocardiograms).12,13 BNP levels can be used to assess the severity of heart failure, as demonstrated by the correlation with New York Heart Association classifications.14 Plasma BNP levels also increase with decreasing physiological functional capacities, as measured by left ventricular ejection fraction (LVEF) or exercise-based evaluations.15,16 The European Society of Cardiology has included the use of natriuretic peptides (e.g., BNP) testing in their guidelines for the diagnosis or rule out of heart failure.17

Others have suggested that BNP has utility in the stratification of patients with heart failure and acute coronary syndrome (ACS). Elevated levels of BNP in heart failure patients predict disease progression and increased morbidity and mortality.18-20 Studies also suggest ACS patients with increased BNP levels have a higher rate of cardiac complications and higher mortality post myocardial infarction.21,22 Preliminary studies have reported the use of BNP measurements to optimize patient treatment / management for heart failure.23-25 Nesiritide (Natrecor), recombinant BNP has been used for treatment in patients with acute, decompensated heart failure.26 The efficacy of BNP monitoring, pre- and post-treatment with Natrecor, has been studied.27 Measurements of BNP two hours or more post-treatment detect only the endogenous levels of BNP.

## Method

The ARCHITECT BNP assay is a two-step immunoassay for the quantitative determination of human B-type natriuretic peptide (BNP) in human EDTA plasma using CMIA technology with flexible assay protocols, referred to as Chemiflex.

## Principle

In the first step, sample and anti-BNP coated paramagnetic microparticles are combined. BNP present in the sample binds to the anti-BNP coated microparticles. After washing, anti-BNP acridinium-labeled conjugate is added to create a reaction mixture in the second step. Following another wash cycle, pre-trigger and trigger solutions are added to the reaction mixture. The resulting chemiluminescent reaction is measured as relative light units (RLUs). A direct relationship exists between the amount of BNP in the sample and the RLUs detected by the ARCHITECT i System optics.

For additional information on system and assay technology, refer to the ARCHITECT System Operations Manual, Section 3.

## Specimen collection and handling

Specimen Types

- EDTA plasma must be used for the ARCHITECT BNP assay. Other sample types, including serum, citrate plasma and heparin plasma are not recommended and have not been evaluated.

- Samples should be collected in plastic collection tubes, because the BNP molecule has been shown to be unstable in glass containers.32,33 Follow the manufacturer’s processing instructions for plasma collection tubes.

- The ARCHITECT i System does not provide the capability to verify sample type. It is the responsibility of the operator to verify that the correct sample type is used in the ARCHITECT BNP assay.

Specimen Conditions

Specimens containing red blood cells or particulate matter may give inconsistent results and must be clarified by centrifugation prior to testing. Gravity separation is not sufficient.

3

- Centrifuged specimens with a lipid layer on top of the liquid must be transferred to a sample cup or secondary tube. Care must be taken to transfer only the clarified specimen and not the lipemic material.

- Inspect all samples for bubbles. Remove bubbles with an applicator stick prior to analysis. Use a new applicator stick for each sample to prevent cross contamination.

- Do not use severely hemolyzed specimens. If a specimen appears to be severely hemolyzed, another specimen should be obtained and tested.

Preparation for Analysis

- Multiple freeze/thaw cycles should be avoided. Samples may undergo up to 3 freeze/thaw cycles. Specimens must be mixed thoroughly after thawing, by LOW speed vortexing or by gently inverting, and centrifuged prior to use, to remove particulate matter and ensure consistency in the results.34

Storage

- Whole blood samples, stored at 2-8°C must be tested within 24 hours of collection.

- Whole blood samples, stored at room temperature must be tested within 4 hours of collection.

- Plasma samples, stored at 2-8°C must be tested within 24 hours of collection.

- Plasma samples, stored at room temperature must be tested within 4 hours of collection.

- If samples cannot be tested within the given times for room temperature or 2-8°C storage, they may be separated by centrifugation and frozen for up to 3 months at -20°C or below in plastic tubes.

Shipping

- Before shipping specimens, it is recommended that specimens be removed from the red blood cells or separator tube.

- When shipping specimens, package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.

- Specimens may be shipped ambient, on wet ice, or on dry ice. Do not exceed the storage limitations listed above.

Materials and Equipment Required

Provided

- 8K28 ARCHITECT BNP Reagent Kit

Materials Required but not Provided

- ARCHITECT i System

3K50 ARCHITECT i	-US- Addition A

3K52 ARCHITECT i	-WW (excluding US)- Addition A

- 8K28-02 ARCHITECT BNP Calibrators

- 8K28-11 ARCHITECT BNP Controls

ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- Pipettes or pipette tips (optional)

For information on materials required for maintenance procedures, refer to the ARCHITECT System Operations Manual, Section 9.

## Reagents – working solutions

Reagent Kit, 100 Tests/500 Tests

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT i Systems. Please contact your local distributor.

ARCHITECT BNP Reagent Kit (8K28)

Microparticles - 1 Bottle (6.6 mL for the 100 test bottle/ 26.8 mL for the 500 test bottle) Anti-BNP (mouse, monoclonal) coated microparticles in TRIS buffer with protein (bovine, mouse) stabilizers. Preservatives: sodium azide and ProClin 950.

Conjugate - 1 Bottle (5.9 mL for the 100 test bottle/26.0 mL for the 500 test bottle) Anti-BNP (mouse, monoclonal) acridinium-labeled conjugate in MES buffer with protein (bovine) stabilizer. Minimum concentration: 0.1 μg/mL. Preservatives: sodium azide and ProClin 300.

Specimen diluents - 1 Bottle (8.9 mL for the 100 test bottle/ 45.4 mL for the 500 test bottle) Specimen Diluent containing TRIS buffer with protein (bovine) stabilizer. Preservatives: sodium azide and ProClin 950.

Other Reagents

ARCHITECT i Pre-Trigger Solution

Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT i Trigger Solution

Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT i Wash Buffer

Wash Buffer containing phosphate buffered saline solution. Preservative: antimicrobial agents.

## Storage and stability

| - The ARCHITECT BNP Reagent Kit must be stored at 2-8°C in an upright position and may be used immediately after removal from 2-8°C storage. - When stored and handled as directed, reagents are stable until the expiration date. - The ARCHITECT BNP Reagent Kit may be stored on board the ARCHITECT i System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking onboard time, refer to the ARCHITECT System Operations Manual, Section 5. - Reagents may be stored on or off the ARCHITECT i System. If reagents are removed from the system, store them at 2-8°C (with septums and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If any reagent bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded. For information on unloading reagents, refer to the ARCHITECT System Operations Manual, Section 5. |  |
| --- | --- |

## Calibration

- To perform an ARCHITECT BNP calibration, test calibrators A, B, C, D, E and F in duplicate. A single sample of each BNP control must be tested to evaluate the assay calibration. Ensure that assay control values are within the concentration ranges specified in the control package insert. Calibrators should be priority loaded.

- Calibrator Range: 0 – 5000 pg/mL.

- Once an ARCHITECT BNP calibration is accepted and stored, all subsequent samples may be tested without further calibration unless:

- A reagent kit with a new lot number is used.

- Controls are out of range.

- For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Quality control

The recommended control requirement for the ARCHITECT BNP assay is that a single sample of each control level be tested once every 24 hours each day of use. The STAT Lab will run 3 levels of  (Architect Controls Low, Med & High) controls once every 24 hours.

The ARCHITECT BNP control values must be within the acceptable ranges specified in the control package insert. If a control is out of the specified range, the associated test results are invalid and samples must be retested. Recalibration may be indicated.

## Preparation of Working Solutions

Handling Precautions

•   Do not use reagent kits beyond the expiration date.

•   Do not pool reagents within a kit or between reagent lots.

•  Prior to loading the ARCHITECT BNP Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this package insert.

•   Septums MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

•   To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

•   Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

- Over time, residual liquids may dry on the septum surface. These are typically dried salts, which have no effect on assay efficacy.

- For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay Procedure

- Before loading the ARCHITECT BNP Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. After the first time the microparticles have been loaded, no further mixing is required.

- Invert the microparticle bottle 30 times.

- Visually inspect the bottle to ensure microparticles are resuspended. If microparticles remain adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

- If the microparticles do not resuspend, DO NOT USE. Contact your local Abbott Representative.

- Once the microparticles have been resuspended, place a septum on the bottle. For instructions on placing septums on bottles, refer to the Handling Precautions section of this package insert.

- Load the ARCHITECT BNP Reagent Kit on the ARCHITECT i System or the ARCHITECT i System with STAT protocol capability. Verify that all necessary assay reagents are present. Ensure that septums are present on all reagent bottles.

- Order calibration, if necessary.

- For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

- Order tests.

- For information on ordering patient specimens and controls and for general operating procedures, refer to the ARCHITECT System Operations Manual, Section 5.

- The minimum sample volume is calculated by the system and is printed on the Orderlist report. No more than 10 replicates may be sampled from the same sample cup. To minimize the effects of evaporation, verify that adequate sample cup volume is present prior to running the test.

- Priority: 150 μL for the first ARCHITECT BNP test plus 100 μL for each additional test from the same sample cup.

- ≤ 3 hours on board: 150 μL for the first ARCHITECT BNP test plus 100 μL for each additional ARCHITECT BNP test from the same sample cup.

- To minimize the effects of evaporation all samples (patient specimens, calibrators and controls) must be tested within 3 hours of being placed on board the ARCHITECT i System.

- If using primary or aliquot tubes, use the sample gauge to ensure that sufficient patient specimen is present.

- Prepare calibrators and controls.

- Mix ARCHITECT BNP Calibrators and Controls by gentle inversion before use.

- To obtain the recommended volume requirements for the ARCHITECT BNP Calibrators and Controls, hold the bottles vertically and dispense a minimum of 5 drops of each calibrator or 5 drops of each control into each respective sample cup.

- Load samples.

- For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

- Press RUN.

- For additional information on principles of operation, refer to the ARCHITECT Operations Manual, Section 3.

- For optimal performance it is important to follow the routine maintenance procedures defined in the ARCHITECT System Operations Manual, Section 9. If your laboratory procedures require more frequent maintenance, follow those procedures.

#### Interpretation: reporting results

Calculation

- The ARCHITECT BNP assay utilizes a point to point data reduction method to generate a calibration curve.

Flags

- Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

- Normal  Range (0-100 pg/mL)

Non-Heart Failure Population

Plasma samples from 890 individuals (465 females, 425 males) who had not been diagnosed with heart failure were tested with the AxSYM BNP assay. This population included non-hospitalized patients with renal disease (not on dialysis), diabetes, hypertension and chronic obstructive pulmonary disease. BNP levels for the patients with renal disease, diabetes, hypertension and chronic obstructive pulmonary disease were not statistically different from the population of apparently healthy individuals. The data from this study are summarized in the following table.*

Non-Heart Failure Population - All (Age Group)

< 45 45-54 55-64 65-74 75+

All Years Years Years Years Years

| Sample Size (N=) | 890 | 205 | 146 | 171 | 248 | 120 |
| --- | --- | --- | --- | --- | --- | --- |
| Median (pg/mL) | 21 | 17 | 9 | 24 | 23 | 31 |
| Mean (pg/mL) | 39 | 28 | 21 | 37 | 47 | 63 |
| SD (pg/mL) | 66 | 36 | 30 | 48 | 80 | 109 |
| 95th Percentile | 135 | 85 | 87 | 119 | 160 | 254 |
| Percentage < 100 pg/mL | 91.5% | 96.6% | 95.2% | 94.2% | 87.1% | 83.3% |
| Minimum (pg/mL) | 0 | 0 | 0 | 0 | 0 | 0 |
| Maximum (pg/mL) | 907 | 263 | 142 | 380 | 907 | 837 |

Non-Heart Failure Population - Males (Age Group)

< 45 45-54 55-64 65-74 75+

All Years Years Years Years Years

| Sample Size (N=) | 425 | 107 | 71 | 94 | 115 | 38 |
| --- | --- | --- | --- | --- | --- | --- |
| Median (pg/mL) | 14 | 12 | 1 | 17 | 21 | 37 |
| Mean (pg/mL) | 30 | 23 | 9 | 26 | 47 | 49 |
| SD (pg/mL) | 61 | 34 | 14 | 45 | 96 | 51 |
| 95th Percentile | 104 | 73 | 40 | 80 | 150 | 121 |
| Percentage < 100 pg/mL | 94.8% | 97.2% | 100.0% | 97.9% | 88.7% | 89.5% |
| Minimum (pg/mL) | 0 | 0 | 0 | 0 | 0 | 0 |
| Maximum (pg/mL) | 907 | 200 | 57 | 380 | 907 | 254 |

Non-Heart Failure Population - Females (Age Group)

< 45 45-54 55-64 65-74 75+

All Years Years Years Years Years

| Sample Size (N=) | 465 | 98 | 75 | 77 | 133 | 82 |
| --- | --- | --- | --- | --- | --- | --- |
| Median (pg/mL) | 26 | 23 | 23 | 37 | 23 | 25 |
| Mean (pg/mL) | 46 | 34 | 34 | 51 | 46 | 69 |
| SD (pg/mL) | 70 | 37 | 36 | 48 | 63 | 126 |
| 95th Percentile | 150 | 89 | 111 | 155 | 159 | 266 |
| Percentage < 100 pg/mL | 88.4% | 95.9% | 90.7% | 89.6% | 85.7% | 80.5% |
| Minimum (pg/mL) | 0 | 0 | 0 | 0 | 0 | 0 |
| Maximum (pg/mL) | 837 | 263 | 142 | 230 | 374 | 837 |

* Representative data; results in individual laboratories may vary from these data.

Due to demographic population differences, the reference range should be established at each laboratory.

Heart Failure Population

Plasma samples from 693 patients with diagnosed heart failure (231 females, 462 males) were tested with the AxSYM BNP assay. All patients in this population were categorized according to the functional classification system published by the New York Heart Association (NYHA).39 This system divides heart failure patients into one of four categories of increasing disease progression (classes I to IV) based upon a subjective assessment of the patient’s clinical signs and symptoms. The data from this study are summarized in the following table.*

Heart Failure Population - All

|  | All | NYHA Functional Class I II III IV |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample Size (N=) | 693 | 124 | 319 | 190 | 60 |
| Median (pg/mL) | 298 | 133 | 266 | 335 | 1531 |
| Mean (pg/mL) | 578 | 320 | 432 | 656 | 1635 |
| SD (pg/mL) | 771 | 388 | 574 | 841 | 1097 |
| 5th Percentile | 14 | 9 | 15 | 12 | 188 |
| 95th Percentile | 2154 | 1257 | 1534 | 2516 | >4000 |
| Percentage _> 100 pg/mL | 74.2% | 58.1% | 73.0% | 79.0% | 98.3% |
| Minimum (pg/mL) | 0 | 3 | 0 | 0 | 14 |
| Maximum (pg/mL) | >4000 | 1651 | >4000 | >4000 | >4000 |

Heart Failure Population – Males

|  | All | NYHA Functional Class I II III IV |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample Size (N=) | 462 | 94 | 215 | 121 | 32 |
| Median (pg/mL) | 268 | 122 | 258 | 293 | 1645 |
| Mean (pg/mL) | 524 | 314 | 409 | 597 | 1646 |
| SD (pg/mL) | 719 | 390 | 539 | 821 | 1032 |
| 5th Percentile | 12 | 9 | 14 | 22 | 265 |
| 95th Percentile | 1976 | 1281 | 1356 | 2288 | 3654 |
| Percentage _> 100 pg/mL | 71.0% | 56.4% | 70.7% | 76.0% | 96.9% |
| Minimum (pg/mL) | 0 | 3 | 0 | 0 | 14 |
| Maximum (pg/mL) | >4000 | 1408 | 3782 | >4000 | >4000 |

Heart Failure Population - Females

|  |  | All |  | NYHA Functional Class I II III IV |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Size (N=) |  | 231 |  | 30 |  | 104 |  | 69 |  | 28 |  |
| Median (pg/mL) |  | 385 |  | 174 |  | 298 |  | 466 |  | 1408 |  |
| Mean (pg/mL) |  | 685 |  | 341 |  | 481 |  | 760 |  | 1623 |  |
| SD (pg/mL) |  | 858 |  | 388 |  | 641 |  | 870 |  | 1186 |  |
| 5th Percentile |  | 16 |  | 14 |  | 21 |  | 12 |  | 244 |  |
| 95th Percentile |  | 2593 |  | 1022 |  | 2031 |  | 2718 |  | >4000 |  |
| Percentage _> 100 pg/mL |  | 80.5% |  | 63.3% |  | 77.9% |  | 84.1% |  | 100.0% |  |
| Minimum (pg/mL) | 0 |  | 10 |  | 0 |  | 0 |  | 173 |  |  |
| Maximum (pg/mL) | >4000 |  | 1651 |  | >4000 |  | >4000 |  | >4000 |  |  |

* Representative data; results in individual laboratories may vary from these data.

A box and whiskers plot of the clinical study population, broken down by NYHA classification, is presented in the following graph. The dashed line represents 100 pg/mL, the suggested decision threshold for the AxSYM BNP assay. In support of previous literature reports,14 these data show a progressive increase in BNP concentrations with increases in NYHA classifications. This analysis indicates that BNP measurements provide objective information for use in the assessment of the severity of heart failure.

Data from the above clinical study were used to generate the Receiver Operating Characteristic (ROC) curve of BNP decision thresholds versus clinical sensitivity and clinical specificity as shown in the following graph. At a decision threshold of 100 pg/mL, the BNP assay demonstrated a clinical sensitivity and specificity of 74.2% and 91.5%, respectively, in this study. The area under the curve (AUC) is 0.90 (0.86 to 0.92, 95% CI).

The ARCHITECT BNP Calibrators are traceable to an internal reference standard that has been prepared gravimetrically with synthetic BNP. The internal reference standard correlates to the AxSYM BNP assay with a decision threshold of 100 pg/mL. There is no internationally recognized BNP standard available at this time.

An age-matched analysis of the heart failure and non-heart failure populations was performed based on the data published by the American Heart Association in the 2000 Heart and Stroke Statistical Update40 and according to the age structure of the United States population.41 The age distributions in the intended use population are approximately as follows: individuals less than 45 years old comprise 9%, individuals 45-54 years old comprise 11%, individuals 55-64 years old comprise 22%, individuals 65-74 years old comprise 26%, and individuals 75 years and older comprise 32%. The resulting combined AUC is 0.87 (0.85 to 0.90, 95% CI).

The clinical sensitivity and specificity using a decision threshold of 100 pg/mL is presented in the following table.*

## Measuring Range:

- The measurement range for the ARCHITECT BNP assay is 10 pg/mL to 5000 pg/mL.

Dilutions

Specimens with a BNP assay value exceeding 5000.0 pg/mL are flagged with the code “>5000.0” and may be diluted using the Automated Dilution Protocol. Retest of samples must be performed within assay specimen handling limits to ensure optimal BNP recovery (refer to the SPECIMEN COLLECTION AND PREPARATION FOR ANALYSIS section of this package insert).

The Manual Dilution Feature of the Assay Parameter is set to ON primarily for use when performing dilution linearity studies.

NOTE: The Automated Dilution Protocol is preferred when diluting specimens due to the known instability of the BNP analyte. If performing manual dilutions, the specimen handling guidelines must be followed (refer to the SPECIMEN COLLECTION AND PREPARATION FOR ANALYSIS section of this package insert).

If using the Automated Dilution Protocol, the system performs a 1:5 dilution of the specimen and automatically calculates the concentration of the sample before dilution and reports the result.

4

- For detailed information on ordering dilutions, refer to the ARCHITECT System Operations Manual, Section 5.

Precautions and Warnings

CAUTION: This product requires the handling of human specimens. It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens28. Biosafety Level 229 or other appropriate biosafety practices30,31 should be used for materials that contain or are suspected of containing infectious agents.

The Microparticles, Conjugate and Specimen Diluent contain methylisothiazolones, which are components of ProClin. These components are classified per applicable European Community (EC) Directives as: Irritant (Xi). The following are the appropriate Risk (R) and Safety (S) phrases.

R43   May cause sensitization by skin contact.

S24   Avoid contact with skin.

S35   This material and its container must be disposed of in a safe way.

S37   Wear suitable gloves.

S46   If swallowed, seek medical advice immediately and show this container or label.

For product not classified as dangerous per European Directive 1999/45/EC as amended - Safety data sheet available for professional user on request.

For a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8

## Limitations — interference

- EDTA plasma, collected in plastic tubes, should be used for this assay. The use of glass collection tubes, or other sample types, such as serum or plasma with other anticoagulants, is not recommended.

- For diagnostic purposes, the ARCHITECT BNP results should be used in conjunction with other clinical data; e.g., symptoms, medical history, etc. If BNP results are not consistent with other clinical observations, additional information may be required for diagnosis.

- Specimens from patients who have received preparations of mouse monoclonal antibodies for diagnosis or therapy may contain human anti-mouse antibodies (HAMA).35 Such specimens may show either falsely elevated or depressed values when tested with assay kits which employ mouse monoclonal antibodies.36 ARCHITECT BNP reagents contain a component that reduces the effect of HAMA reactive specimens. Additional clinical or diagnosis information may be required to determine patient status.

- Heterophilic antibodies in human plasma can react with reagent immunoglobulins, interfering with in vitro immunoassays. The presence of heterophilic antibodies in a patient may cause anomalous values to be observed.37,38

- ARCHITECT BNP results should not be used interchangeably with other
manufacturers’ methods for BNP or NT-proBNP determinations.

- Measurements of BNP should occur prior to Nesiritide (Natrecor), recombinant BNP treatment and 2 hours post-treatment.27

Refer to the SPECIMEN COLLECTION AND PREPARATION FOR ANALYSIS section in this package insert for additional information

Potential interference in the ARCHITECT BNP assay from the following compounds is designed to be ≤ 10%.

Potential interference was evaluated by a study based upon guidance from NCCLS Protocol EP7-A.43 Specimens were supplemented with various drugs and potentially interfering substances (triglycerides, hemoglobin, bilirubin, and total protein) at the levels indicated in the following tables. The average recovery observed during the study ranged from 92 to 110%.*

| Drug Drug Concentration |  | Drug Drug Concentration |  |
| --- | --- | --- | --- |
| Acetaminophen | 30 μg/mL | Indomethacin 36 μg/mL |  |
| Acetylsalicylic Acid | 600 μg/mL | Isosorbide Dinitrate 150 ng/mL |  |
| Amiodarone | 6 μg/mL | Lisinopril | 4 μg/mL |
| Amlodipine besylate | 100 ng/mL | Lovastatin | 20 μg/mL |
| Ampicillin | 53 μg/mL | Methyldopa | 15 μg/mL |
| Ascorbic Acid | 40 μg/mL | Nicotine | 1 μg/mL |
| Atenolol | 10 μg/mL | Nifedipine | 400 ng/mL |
| Caffeine | 60 μg/mL | Nitrofurantoin | 4 μg/mL |
| Captopril | 5 μg/mL | Nitroglycerine | 500 ng/mL |
| Chloramphenicol | 50 μg/mL | Oxazepam | 5 μg/mL |
| Clopidogrel Bisulphate | 2.5 μg/mL | Oxytetracycline | 15 μg/mL |
| Cyclosporine | 2.5 μg/mL | Phenobarbital | 100 μg/mL |
| Diclofenac | 50 μg/mL | Phenytoin | 50 μg/mL |
| Digoxin | 2 ng/mL | Probenecid | 600 μg/mL |
| Diltiazem | 40 μg/mL | Procainamide | 24 μg/mL |
| Dipyridamole | 80 μg/mL | Propranolol | 2 μg/mL |
| Dobutamine | 100 μg/mL | Quinidine | 12 μg/mL |
| Dopamine | 900 ng/mL | Simvastatin | 16 μg/mL |
| Enalapril Maleate | 300 ng/mL | Spironolactone | 600 ng/mL |
| Erythromycin | 60 μg/mL | Sulfamethoxazole | 400 μg/mL |
| Fenofibrate | 45 μg/mL | Trandolapril | 40 μg/mL |
| Furosemide | 60 μg/mL | Trimethoprim | 40 μg/mL |
| Heparin | 8 U/mL | Verapamil | 2 μg/mL |
| Hydralazine | 6.4 μg/mL | Warfarin | 20 μg/mL |
| Hydrochlorothiazide | 6 μg/mL |  |  |

Potentially Interfering

Substance	Concentration

Triglycerides	3000 mg/dL

Hemoglobin	500 mg/dL

Bilirubin	20 mg/dL

Total Protein	3 g/dL

Total Protein	12 g/dL

* Representative data; results in individual laboratories may vary from these data.

## Performance characteristics

Precision

6

The ARCHITECT BNP assay is designed to have an upper 95% Confidence Interval (CI) imprecision of ≤ 12% total CV.

A study was performed with the ARCHITECT BNP assay based on guidance from the National Committee for Clinical Laboratory Standards (NCCLS) Protocol EP5-A2.42 A three-member panel and the ARCHITECT BNP Controls were assayed in replicates of two at two separate times of the day for 20 days. Testing was performed on two ARCHITECT Systems using a single calibration on each instrument. Data from this study are summarized in the following table.*

| Sample | Reagent InstruLot ment |  | n | Mean Conc. (pg/mL) | Within Run SD %CV |  | Total SD %CV |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Panel 1 | 1 | 1 | 80 | 183.6 | 6.9 | 3.8 | 9.8 | 5.3 |
|  | 2 | 2 | 80 | 180.1 | 10.0 | 5.6 | 12.1 | 6.7 |
| Panel 2 | 1 | 1 | 80 | 1004.4 | 26.4 | 2.6 | 41.1 | 4.1 |
|  | 2 | 2 | 80 | 961.6 | 44.8 | 4.7 | 53.7 | 5.6 |
| Panel 3 | 1 | 1 | 80 | 3812.1 | 91.3 | 2.4 | 153.5 | 4.0 |
|  | 2 | 2 | 80 | 3645.1 | 138.4 | 3.8 | 188.9 | 5.2 |
| Low | 1 | 1 | 80 | 92.2 | 3.2 | 3.5 | 4.0 | 4.4 |
| Control | 2 | 2 | 80 | 87.8 | 3.6 | 4.0 | 4.1 | 4.7 |
| Medium | 1 | 1 | 80 | 504.3 | 9.3 | 1.8 | 13.5 | 2.7 |
| Control | 2 | 2 | 80 | 494.1 | 13.1 | 2.6 | 17.2 | 3.5 |
| High | 1 | 1 | 80 | 3572.1 | 32.6 | 0.9 | 62.0 | 1.7 |
| Control | 2 | 2 | 80 | 3423.0 | 80.1 | 2.3 | 112.3 | 3.3 |

* Representative data; results in individual laboratories may vary from these data.

Dilution Linearity/Recovery

The ARCHITECT BNP assay is designed to have a mean recovery of 100 ± 15% of the expected result for diluted specimens.

A dilution linearity study was performed by diluting high concentration BNP specimens with the ARCHITECT BNP Specimen Diluent. The concentration of BNP was determined for each dilution of sample and the percent (%) recovery was calculated. Data from this study are summarized in the following table.*

| Sample | Final Dilution Factor | Value Obtained (pg/mL) | % Recoverya |
| --- | --- | --- | --- |
| 1 | Undiluted | 4772 | - |
|  | 1.33 | 3489 | 98 |
|  | 2 | 2426 | 102 |
|  | 5 | 921 | 97 |
| 2 | Undiluted | 3791 | - |
|  | 1.33 | 2679 | 94 |
|  | 2 | 1918 | 101 |
|  | 5 | 675 | 89 |
| 3 | Undiluted | 4071 | - |
|  | 1.33 | 3107 | 102 |
|  | 2 | 1556 | 77 |
|  | 5 | 739 | 91 |
| 4 | Undiluted | 4985 | - |
|  | 1.33 | 3548 | 95 |
|  | 2 | 2237 | 90 |
|  | 5 | 911 | 91 |
| 5 | Undiluted | 4011 | - |
|  | 1.33 | 3061 | 102 |
|  | 2 | 2032 | 101 |
|  | 5 | 716 | 89 |

Mean recovery across the five diluted samples shown above = 95%

BNP Value Obtained (pg/mL) x100

a % Recovery=
   (Undiluted Value / Dilution Factor)

* Representative data; results in individual laboratories may vary from these data.

Analytical Sensitivity

7

The ARCHITECT BNP assay is designed to have an analytical sensitivity of ≤ 10 pg/mL.

Analytical sensitivity is defined as the concentration calculated as the mean plus two standard deviations of replicates of the ARCHITECT BNP Calibrator A (0 pg/mL). The analytical sensitivity (low-linearity) is defined in the ARCHITECT BNP assay parameters as 10 pg/mL.

Analytical Specificity

The specificity of the ARCHITECT BNP assay is designed to be ≤ 10 pg/mL of measured BNP when tested with human ANP, Angiotensin I, II and III, CNP, and NT-proBNP at the following concentrations. Each potential cross-reactant was added to protease-inhibitor treated plasma and then assayed. Data from this study are summarized in the following table.*

| Potential Cross-reactant | Concentration (pg/mL) | BNP Concentrationa (pg/mL) |
| --- | --- | --- |
| ANP | 1000 | < 10 |
| Angiotensin I | 600 | < 10 |
| Angiotensin II | 600 | < 10 |
| Angiotensin III | 1000 | < 10 |
| CNP | 1000 | < 10 |
| NT-proBNP (1-76) | 1000 | < 10 |

a Cross-reactivity = BNP Value Obtained (pg/mL) - Endogenous BNP Level (pg/mL)

* Representative data; results in individual laboratories may vary from these data.

Carryover

Carryover from a sample with a high BNP value (approximately 25000 pg/mL) to an adjacent sample of ARCHITECT BNP Calibrator A (0 pg/mL) was determined to be below the analytical sensitivity of the assay.*

* Representative data; results in individual laboratories may vary from these data.

Method Comparison

The ARCHITECT BNP assay is designed to have a slope of 1.0 ± 0.2 and a correlation coefficient (r) of _> 0.90 when compared to the AxSYM BNP assay.

A study was performed with guidance from NCCLS Protocol EP9-A244 to compare the ARCHITECT BNP assay to the AxSYM BNP assay. EDTA plasma samples from 171 individuals (128 heart failure patients, 43 non-heart failure individuals) were tested with both assays. These samples were collected from populations of individuals with and without heart failure. The results from the Passing-Bablok45 linear regression analysis are summarized in the following table.*

ARCHITECT BNP vs. AxSYM BNP

Regression Specimen	Correlation Intercept	Slope

Method	Type	n Coefficient (95% CI)	(95% CI)

Passing	EDTA	171	0.96	-38.32	1.03

Bablok†	Plasma	(-48.26 to	(0.98 to

-28.50)	1.09)

Sample Range (ARCHITECT): 0 – 3702 pg/mL

Sample Range (AxSYM): 50 – 3103 pg/mL

† A linear regression method with no special assumptions regarding the distribution of the samples and measurement errors.

* Representative data; variables such as differences in sampling size and population may impact the correlation of the assay, therefore, results in individual laboratories may vary from these data.

Contacts:

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Alternative method

Send to sister facility

References

- de Bold AJ, Borenstein HB, Veress AT, et al. A rapid and potent natriuretic response to intravenous injection of atrial myocardial extract in rats. Life Sci 1981;28(1):89-94.

- Sudoh T, Kangawa K, Minamino N, et al. A new natriuretic peptide in porcine brain. Nature 1988;332:78-81.

- Yandle TG. Biochemistry of natriuretic peptides. J Intern Med 1994;235: 561-76.

- Davidson NC, Struthers AD. Brain natriuretic peptide. J Hypertens 1994;12(4):329-36.

- Kambayashi Y, Nakao K, Mukoyama M, et al. Isolation and sequence determination of human brain natriuretic peptide in human atrium. FEBS Lett 1990;259(2):341-5.

- Sudoh T, Maekawa K, Kojima M, et al. Cloning and sequence analysis of cDNA encoding a precursor for human brain natriuretic peptide. Biochem Biophys Res Commun 1989;159(3):1427-34.

- Shimizu H, Masuta K, Aono K, et al. Molecular forms of human brain natriuretic peptide in plasma. Clin Chim Acta 2002;316:129-35.

- Hunt PJ, Espiner EA, Nicholls MG, et al. The role of the circulation in processing pro-brain natriuretic peptide (proBNP) to amino-terminal BNP and BNP-32. Peptides 1997;18(10):1475-81.

- Holmes SJ, Espiner EA, Richards AM, et al. Renal, endocrine, and hemodynamic effects of human brain natriuretic peptide in normal man. J Clin Endocrinol Metab 1993;76(1):91-6.

- Sagnella GA. Measurement and importance of plasma brain natriuretic peptide and related peptides. Ann Clin Biochem 2001;38:83-93.

- Sagnella GA. Measurement and significance of circulating natriuretic peptides in cardiovascular disease. Clin Sci 1998;95:519-29.

- Maisel AS, Krishnaswamy P, Nowak RM, et al. Rapid measurement of B-type natriuretic peptide in the emergency diagnosis of heart failure. N Engl J Med 2002;347(3):161-7.

- Latini R, Masson S, Anand I, et al. Effects of valsartan on circulating brain natriuretic peptide and norepinephrine in symptomatic chronic heart failure: the Valsartan Heart Failure Trial (Val-HeFT) Circulation 2002;106:2454-8.

- Wieczorek SJ, Wu AHB, Christenson R, et al. A rapid B-type natriuretic peptide assay accurately diagnoses left ventricular dysfunction and heart failure: a multicenter evaluation. Am Heart J 2002, 144(5):834-9.

- Fischer Y, Filzmaier K, Stiegler H, et al. Evaluation of a new, rapid bedside test for quantitative determination of B-type natriuretic peptide. Clin Chem 2001;47(3):591-4.

- Wieczorek SJ, Hager D, Barry M-B, et al. Correlation of B-type natriuretic peptide level to 6-min walk test performance in patients with left ventricular systolic dysfunction. Clin Chem Acta 2003;328:87-90.

- Remme WJ, Swedberg K. Task Force Report: Guidelines for the diagnosis and treatment of chronic heart failure. Eur Heart J 2001;22(17):1527-60.

- McDonagh TA, Cunningham AD, Morrison CE, et al. Left ventricular dysfunction, natriuretic peptides, and mortality in an urban population. Heart 2001;86:21-6.

- Anand IS, Fisher LD, Chiang Y-T, et al. Changes in brain natriuretic peptide and norepinephrine over time and mortality and morbidity in the Valsartan Heart Failure Trial (Val-HeFT). Circulation 2003;107: 1278-83.

- Silver MA, Maisel A, Yancy CW, et al. BNP concensus panel 2004: a clinical approach for the diagnostic, prognostic, screening, treatment monitoring, and therapeutic roles of natriuretic peptides in cardiovascular diseases [published correction appears in Congest Heart Fail 2005;11(2):102]. Congest Heart Fail 2004;10(5 suppl 3): 1–30.

- Omland T, Aakvaag A, Bonarjee VVS, et al. Plasma brain natriuretic peptide as an indicator of left ventricular systolic function and long-term survival after acute myocardial infarction. Comparison with plasma atrial natriuretic peptide and N-terminal proatrial natriuretic peptide. Circulation 1996;93(11):1963-9.

- De Lemos JA, Morrow DA, Bentley JH et al. The prognostic value of B-type natriuretic peptide in patients with acute coronary syndromes. N Engl J Med 2001;345(14):1014-21.

- Lee S-C, Stevens TL, Sandberg SM, et al. The potential of brain natriuretic peptide as a biomarker for New York Heart Association class during the outpatient treatment of heart failure. J Card Fail 2002;8(3):149-54.

- Richards AM, Lainchbury JG, Nicholls MG, et al. BNP in hormone-guided treatment of heart failure. Trends Endocrinol Metab. 2002;13(4):151-5.

- Troughton RW, Frampton CM, Yandle TG, et al. Treatment of heart failure guided by plasma aminoterminal brain natriuretic peptide (N-BNP) concentrations. Lancet 2000;355:1126-30.

- Burger AJ, Horton DP, LeJemtel T, et al. Effect of nesiritide (B-type natriuretic peptide) and dobutamine on ventricular arrhythmias in the treatment of patients with acutely decompensated congestive heart failure: the PRECEDENT study. Am Heart J 2002;144(6):1102-8.

- Maisel AS, Cremo R, Gardetto N, et al. The effects of nesiritide on serum levels of B-type natriuretic peptide (BNP) in patients admitted for decompensated congestive heart failure [Abstr]. Circulation (Suppl II) 2002;106(19):II-565.

- US Department of Labor, Occupational Safety and Health Administration, 29 CFR Part 1910.1030, Bloodborne Pathogens.

- US Department of Health and Human Services. Biosafety in Microbiological and Biomedical Laboratories, Fifth Edition. Washington, DC: US Government Printing Office, January 2007.

- World Health Organization. Laboratory Biosafety Manual. Geneva: World Health Organization, 2004.

- Clinical and Laboratory Standards Institute. Protection of Laboratory Workers from Occupationally Acquired Infections: Approved Guideline— Third Edition. CLSI Document M29-A3. Wayne, PA: Clinical and Laboratory Standards Institute, 2005.

- Shimizu H, Aono K, Masuta K, et al. Stability of brain natriuretic peptide (BNP) in human blood samples. Clin Chim Acta 1999;285:169-72.

- Shimizu H, Aono K, Masuta K, et al. Degradation of human brain natriuretic peptide (BNP) by contact activation of blood coagulation system. Clin Chim Acta 2001;305:181-6.

- Constantine NT, Abouseif NE, and Fox E. To mix or not to mix: the effects of not mixing sera on HIV serologic results. Lab Medicine 1990;21(11):749-51.

- Schroff RW, Foon KA, Beatty SM, et al. Human anti-murine immunoglobulin responses in patients receiving monoclonal antibody therapy. Cancer Res 1985;45:879-85.

Primus FJ, Kelley EA, Hansen HJ, et al. “Sandwich”-type immunoassay of carcinoembryonic antigen in patients receiving murine monoclonal antibodies for diagnosis and therapy. Clin Chem 1988;34(2):261-4.

8

- Nahm MH, Hoffmann JW. Heteroantibody: phantom of the immunoassay. Clin Chem 1990;36(6):829.

- Boscato LM, Stuart MC. Heterophilic antibodies: a problem for all immunoassays. Clin Chem 1988;34(1):27-33.

- The Criteria Committee of the New York Heart Association. Nomenclature and criteria for diagnosis of diseases of the heart and great vessels. 9th ed. Boston, Mass: Little, Brown & Co; 1994:253-6.

- American Heart Association. 2000 Heart and Stroke Statistical Update. Dallas, TX: American Heart Association;1999:18-19.

- MacKay AP, Fingerhut LA, Duran CR. Adolescent Health Chartbook. Health, United States, 2000. Hyattsville, MD: National Center for Health Statistics; 2000:123.

- National Committee for Clinical Laboratory Standards. Evaluation of precision performance of quantitative measurement methods; Approved Guideline-Second Edition. NCCLS Document EP5-A2, Wayne, PA: NCCLS, 2004.

- National Committee for Clinical Laboratory Standards. Interference testing in clinical chemistry; Approved Guideline. NCCLS document EP7-A, Wayne, PA: NCCLS, 2002.

- National Committee for Clinical Laboratory Standards. Method comparison and bias estimation using patient samples; Approved Guideline-Second Edition. NCCLS Document EP9-A2, Wayne, PA: NCCLS, 2002.

- Passing H, Bablok W. A new biometrical procedure for testing the equality of measurements from two different analytical methods. J Clin Chem Clin Biochem 1983; 21(11):709-20.

Effective date

Effective date for this procedure: 01/10/18

Author

Compiled by Abbott Diagnostics

Revised by:  	 Jessica Mercer, MLS (ASCP)

Brooke Ross, MLS (ASCP)

Elsa Escobar MT (ASCP)