Intended use

The ARCHITECT TROPONIN I assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cardiac troponin-I  in human EDTA plasma on the ARCHITECT i System. Troponin-I values are used to assist in the diagnosis of myocardial infarction (MI).

Summary

Troponin-I (TnI) is a regulatory subunit of the troponin complex associated with the actin thin filament within muscle cells.  TnI, in conjunction with troponin-C and troponin-T, plays an integral role in the regulation of muscle contraction.  Three distinct tissue specific isoforms of TnI have been identified from skeletal and cardiac muscles.  The cardiac isoform exhibits only 60% similarity with the skeletal muscle isoform and contains additional amino acids at the N-terminus; cardiac troponin-I (cTnI) has a molecular weight of approximately 24,000 daltons.

Clinical studies have demonstrated the release of cTnI into the blood stream within hours following acute myocardial infarctions (AMI) or ischemic damage.  Elevated levels of cTnI (above the values established for non-AMI specimens) are detectable in serum within 4 to 6 hours after the onset of chest pain, reach peak concentrations in approximately 8 to 28 hours, and remain elevated for 3 to 10 days following AMI.  The temporal pattern of cTnI release following an infarction thus extends across the diagnostic windows of both creatine kinase-MB (CK-MB) and lactate dehydrogenase (LD).  The clinical utility of cTnI measurements for the assessment of myocardial damage has been demonstrated in several clinical studies indicating improved cardiac specificity of cTnI over CK-MB.  The high specificity of cTnI measurements is beneficial in identifying cardiac injury for clinical conditions involving skeletal muscle injury resulting from surgery, trauma, extensive exercise, or muscular disease.

The World Health Organization (WHO) criteria for defining AMI are the presence of two of the following three elements:  unequivocal ECG changes, unequivocal serum cardiac enzyme changes, and prolonged chest pain.  The Joint European Society of Cardiology/American College of Cardiology Committee current guideline redefines MI and supports the use of cTnI as a preferred marker for myocardial injury.  Their definition of MI is a typical rise and gradual fall of cardiac troponin (or more rapid fall of CK-MB) with at least one of the following:  ischemic symptoms, pathological Q waves on electrocardiogram (ECG), ischemic ECG changes, or coronary artery intervention.  Serial sampling is recommended to detect the temporal rise and fall of troponin levels characteristic of MI.  An elevated troponin alone is not sufficient to make the diagnosis of myocardial infarction.  Other markers such as CK-MB and myoglobin can be used in conjunction with troponin-I results in aiding the diagnosis of MI.

Several major studies have shown that cTnI is also useful as a predictor of cardiac risk in patients with unstable angina.  Previous studies showed that during a 30-day follow-up, patients with acute coronary syndromes (including unstable angina) were at greater risk of progressing to MI if cTnI is elevated.  Results from the PRISM trial showed that elevated cTnI levels could help to identify patients with unstable angina who had additional cardiac risk (especially within the first 72 hours after onset of symptoms) and who could benefit from treatment with a glycoprotein IIb/IIIa-receptor antagonist.  Thus, cTnI can play an important role in identifying patients with acute coronary syndromes who are at greater risk for cardiac events.  The American College of Cardiology and the American Heart Association current guidelines also recommend using troponin results when making treatment decisions regarding unstable angina and non-ST segment elevation MI (NSTEMI).

## Method

The ARCHITECT STAT Troponin-I assay is a two-step immunoassay to determine the presence of cTnI in human serum and plasma using CMIA technology with flexible assay protocols, referred to as Chemiflex.

## Principle

In the first step, sample, assay diluent and anti-troponin-I antibody-coated paramagnetic microparticles are combined.  Troponin-I present in the sample binds to the anti-troponin-I coated microparticles.  After incubation and wash, anti-troponin-I acridinium labeled conjugate is added in the second step.  Following another incubation and wash, pre-trigger and trigger solutions are then added to the reaction mixture.  The resulting chemiluminescent reaction is measured as relative light units (RLUs).  A direct relationship exists between the amount of Troponin-I in the sample and the RLUs detected by the ARCHITECT i system optics.  The concentration of troponin-I is read relative to a standard curve established with calibrators of known troponin-I concentrations.

For additional information on system and assay technology, refer to the ARCHITECT i System Operations Manual, Section 3.

## Specimen collection and handling

1.  The following specimen collection tubes may be used for the ARCHITECT STAT Troponin-I assay.

Serum:	Glass: 	Will not be used

Plastic:  Serum separator tubes with clot activators

Plasma

Glass:	 Will not be used

Plastic:  Lithium Heparin

Plasma separator tubes with lithium heparin

Sodium Heparin

Other anticoagulants have not been validated for use with the ARCHITECT STAT Troponin-I assay.

2.  Refer to the specimen collection tube manufacturer's instructions as well as these package insert instructions for specimen collection and preparation for analysis.  Each laboratory should determine the acceptability of its own specimen collection and processing procedures.

3.  Inadequate centrifugation of the specimen may cause an erroneous result.

4.  For serum specimens, ensure that complete clot formation has taken place prior to centrifugation.  Some specimens, especially those from patients receiving anticoagulant or thrombolytic therapy, may exhibit increased clotting times.  If the specimen is centrifuged before complete clot formation, the presence of fibrin may cause erroneous results.

5.  If a lipid layer forms on the specimen surface, avoid the lipid layer when withdrawing the specimen.

6.  If testing will be delayed more than 8 hours, remove the plasma or serum from the cells, clot, or gel.  Specimens removed from the cells, clot, or gel may be stored up to 72 hours at 2-8 degrees C. or stored frozen (-10 degrees C. or colder) prior to being tested.

7.  Specimens can be stored up to 30 days frozen at -10 degrees C. or colder.

8.  Thaw frozen specimens and mix thoroughly by LOW speed vortexing or by gently inverting, then centrifuge at 2,500-3,000 x g for 10 minutes prior to use to remove particulate matter and to ensure consistency in the results.  Thaw specimens only once.

9.  Ensure specimens are free of fibrin, red blood cells, and other particulate matter.

10.  Inspect all samples for bubbles.  Remove bubbles prior to analysis.

11.  When serial specimens are being evaluated, use the same specimen type throughout the evaluation.

12.  The ARCHITECT i System does not provide the capability to verify specimen type.  It is the responsibility of the operator to verify the correct specimen type is used in the ARCHITECT STAT Troponin-I assay.

13.  Use caution when handling patient specimens to prevent cross contamination.  Use of disposable pipettes or pipette tips is recommended.

14.  Test all samples (patient specimens, controls, and calibrators) within 3 hours of being placed on board the ARCHITECT i System.  Refer to the ARCHITECT System Operations Manual, Section 5, for a more detailed discussion of on-board sample storage constraints.

15.  Performance has not been established using cadaver specimens or body fluids other than human plasma or serum.

16.  Do not use samples with obvious microbial contamination.

17.  Do not use heat-inactivated specimens.

18.  When shipping specimens, package and label specimens in compliance with applicable state, federal, and international regulations covering the transport of clinical specimens and infectious substances.  Prior to shipping, remove the plasma or serum specimen from the cells, clot, or gel.  Ship specimens frozen on dry ice.

Materials and Equipment Required

Provided

- 2K41 ARCHITECT STAT Troponin-I Reagent Kit

Materials Required but not Provided

- ARCHITECT i System

3K51 ARCHITECT i	-US- Addition B

3K53 ARCHITECT i	-WW (excluding US)- Addition B

- 2K41-01 ARCHITECT STAT Troponin-I Calibrators

- 2K41-10 ARCHITECT STAT Troponin-I Controls

ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- ARCHITECT i

- Pipettes or pipette tips (optional)

For information on materials required for maintenance procedures, refer to the ARCHITECT System Operations Manual, Section 9.

## Reagents – working solutions

Reagent Kit, 100 Tests/500 Tests

NOTE: Some kit sizes are not available in all countries or for use on all ARCHITECT i Systems. Please contact your local distributor.

ARCHITECT STAT TROPONIN-I Reagent Kit (2K41)

1.  Microparticles:  1 or 4 Bottles (6.6 mL/27.0 mL) Anti-troponin-I (mouse, monoclonal) coated microparticles in TRIS buffer with protein (bovine and goat) stabilizers.  Preservative:  antimicrobial agents.

2.  Conjugate:  1 or 4 Bottles (5.9 mL/26.3 mL) Anti-troponin-I(mouse, monoclonal) acridinium labeled conjugate in MES buffer with protein (bovine) stabilizer.  Preservative:  ProClin 300.

3.  Assay Diluent:  1 or 4 Bottles (10.0 mL/50.9 mL) Troponin-I Assay Diluent, containing protein (bovine and goat) stabilizers in phosphate buffer.  Preservative:  ProClin 300.

Other Reagents

ARCHITECT i Pre-Trigger Solution

Pre-Trigger Solution containing 1.32% (w/v) hydrogen peroxide.

ARCHITECT i Trigger Solution

Trigger Solution containing 0.35 N sodium hydroxide.

ARCHITECT i Wash Buffer

Wash Buffer containing phosphate buffered saline solution. Preservative: antimicrobial agents.

## Storage and stability

| 1. The ARCHITECT STAT Troponin-I Reagent Kit must be stored at 2-8 degrees C. in an upright position and may be used immediately after removal from 2-8 degrees C. storage. 2. When stored and handled as directed, reagents are stable until the expiration date. 3. The ARCHITECT STAT Troponin-I Reagent Kit may be stored on-board the ARCHITECT i System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking on-board time, refer to the ARCHITECT System Operations Manual, Section 5. 4. Reagents may be stored on or off the ARCHITECT i System. If reagents are removed from the system, store them at 2-8 degrees C. (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright. If the microparticle bottle does not remain upright (with a septum installed) while in 2-8 degrees C. storage off the system, the reagent kit must be discarded. After reagents are removed from the system, one must initiate a scan to update the on-board stability timer. INDICATIONS OF REAGENT DETERIORATION When a control value is out of the specified range, it may indicate deterioration of the reagents or errors in technique. Associated test results may be invalid and may require retesting. Assay recalibration may be necessary. For troubleshooting information, refer to the ARCHiTECT System Operations Manual, Section 10. |  |
| --- | --- |

## Calibration

1.  To perform an ARCHITECT STAT Troponin-I calibration, test the Calibrators A, B, C, D, E, and F in duplicate.  A single sample of each ARCHITECT STAT Troponin-I Control level must be tested to evaluate the assay calibration.  Ensure that assay control values are within the concentration ranges specified in the control package insert.  Calibrators should be priority loaded.

2.  Calibration range:  0.00 - 50.00 ng/mL.

3.  Once an ARCHITECT STAT Troponin-I calibration is accepted and stored, all subsequent samples may be tested without further calibration unless:

a.  A reagent kit with a new lot number is used

b.  Controls are out of range

4.  For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Quality control

The recommended control requirement for the ARCHITECT STAT Troponin-I assay is that a single sample of each control level be tested once every 24 hours each day of use. The STAT Lab will run 3 levels of  (Architect Controls Low, Med & High) controls once every 24 hours.

The ARCHITECT STAT Troponon-I control values must be within the acceptable ranges specified in the control package insert. If a control is out of the specified range, the associated test results are invalid and samples must be retested. Recalibration may be indicated.

For detailed information on how to perform an assay calibration, refer to the ARCHITECT System Operations Manual, Section 6.

## Preparation of Working Solutions

Handling Precautions

•   Do not use reagent kits beyond the expiration date.

•   Do not pool reagents within a kit or between reagent lots.

•  Prior to loading the ARCHITECT STAT Tropnin-I Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment. For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this package insert.

•   Septums MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septums are not used according to the instructions in this package insert.

•   To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

•   Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

- Over time, residual liquids may dry on the septum surface. These are typically dried salts, which have no effect on assay efficacy.

- For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 7.

## Assay Procedure

INSTRUMENT PROCEDURE

1.  The ARCHITECT STAT Troponin-I assay file must be installed on the ARCHITECT i System with STAT protocol capability from the ARCHITECT i Assay CD-ROM Addition B prior to performing the assay.  For detailed information on assay file installation and on viewing and editing assay parameters, refer to the ARCHITECT System Operations Manual, Section 2.

2.  For information of printing assay parameters, refer to the ARCHITECT System Operations Manual, Section 5.

3.  For a detailed description of system procedures, refer to the ARCHITECT System Operations Manual.

4.  The default result unit for the ARCHITECT STAT Troponin-I assay is ng/mL.  An alternate result unit, ug/L, may be selected for reporting results by editing assay parameter "Result concentration units", to ug/L.  The conversion factor used by the system is 1.0.

ASSAY PROCEDURE

1.  Before loading the ARCHITECT STAT Troponin-I Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment:

a.  Invert the microparticle bottle 30 times.

b.  Visually inspect the bottle to ensure microparticles are resuspended.  If microparticles are still adhered to the bottle, continue to invert the bottle until the microparticles have been completely resuspended.

c.  If the microparticles do not resuspend, DO NOT USE.  Contact your local Abbott Laboratories representative.

d.  Once the microparticles have been resuspended, remove and discard the cap.  Wearing clean gloves, remove a septum from the bag.  Squeeze the septum in half to confirm that the slits are open.  Carefully snap the septum onto the top of the bottle.

2.  Order calibration, if necessary

a.  For information on ordering calibrations, refer to the ARCHITECT System Operations Manual, Section 6.

3.  Order tests.

a.  For information on ordering patient specimens and controls, refer to the ARCHITECT System Operations Manual, Section 5.

4.  Load the ARCHITECT STAT Troponin-I Reagent Kit on the ARCHITECT i System with STAT protocol capability.

a.  Verify that all necessary assay reagents are present.  Ensure that septa are present on all reagent bottles.

5.  The minimum sample volume is calculated by the system and is printed on the Orderlist report.  No more than 10 replicates may be sampled from the same sample cup.  To minimize the effects of evaporation, verify adequate sample cup volume is present prior to running the test.

a.  Priority:  125 uL for the first ARCHITECT STAT Troponin-I test plus 75 uL for each additional ARCHITECT STAT Troponin-I test from the same sample cup.

b.  </= 3 hours on board:  150 uL for the first ARCHITECT STAT Troponin-I test plus 75 uL for each additional ARCHITECT STAT Troponin-I test from the same sample cup.

c.  If using primary or aliquot tubes, use the sample gauge to ensure sufficient patient specimen is present.

6.  Prepare calibrators and controls.

a.  ARCHITECT STAT Troponin-I Calibrators and Controls should be mixed according to instructions in their respective package inserts.

b.  To obtain the recommended volume requirements for the ARCHITECT STAT Troponin-I Calibrators, hold the bottles vertically and dispense 9 drops of each calibrator into each respective sample cup.  Dispense 150 uL of each control into each respective sample cup.

7.  Load samples.

a.  For information on loading samples, refer to the ARCHITECT System Operations Manual, Section 5.

8.  Press RUN.  The system performs the following functions:

a.  Moves the sample to the aspiration point.

b.  Loads a reaction vessel (RV) into the process path.

c.  Aspirates and transfers sample into the RV.

d.  Advances the RV one position and transfers assay diluent and microparticles into the RV.

e.  Mixes, incubates, and washes the reaction mixture.

f.  Adds conjugate to the RV.

g.  Mixes, incubates, and washes the reaction mixture.

h.  Adds pre-trigger and trigger solutions.

i.  Measures chemiluminescent emission to determine the quantity of troponin-I in the sample.

j.  Aspirates contents of RV to liquid waste and unloads RV to solid waste.

k.  Calculates the result.

9.  For optimal performance, it is important to follow the routine maintenance procedures defined in the ARCHITECT System Operations Manual, Section 9.  If your laboratory requires more frequent maintenance, follow those procedures.

#### Interpretation: reporting results

Calculation

- The ARCHITECT STAT Troponin-I assay utilizes a 4 Parameter Logistic Curve Fit (4PLC,Y weighted) data reduction method to generate a calibration curve.

Flags

- Some results may contain information in the Flags field. For a description of the flags that may appear in this field, refer to the ARCHITECT System Operations Manual, Section 5.

- Normal Range  Female 0.000-0.013 ng/mL

- Normal Range Male  0.000-0.033 ng/mL

Critical Values:

- Troponin I -- ≥ 0.290 ng/mL

- Call all critical values to the appropriate nursing unit/floor using the 1READS or1CALL comment according yo the Critical Values Policy LAB-106-P

Any condition resulting in myocardial cell damage can potentially increase cardiac troponin-I levels.  Published studies have documented that these conditions include, but are not limited to, angina, unstable angina, congestive heart failure, myocarditis, cardiac surgery, or invasive testing and non-cardiac related causes such as pulmonary embolism, renal failure, and sepsis.

Serial sampling is recommended to detect the temporal rise and fall of troponin levels characteristic of MI.

For diagnostic cutoff and additional information, refer to the SPECIFIC PERFORMANCE CHARACTERISTICS, Clinical Performance section in this package insert.

A reference range study was conducted based on guidance from National Committee for Clinical Laboratory Standards (NCCLS) Protocol C28-A2.  Apparently healthy individuals were evaluated in replicates of one using the ARCHITECT STAT Troponin-I assay.  The observed values are summarized in the following table.*

APPARENTLY HEALTHY POPULATION

Population:  Female

n:  225

Age Range:  18-62

99th Percentile (ng/mL):  0.0130

Population:  Male

n:  224

Age Range:  18-63

99th Percentile (ng/mL):  0.033

Population:  TOTAL

n:  449

Age Range:  18-63

99th Percentile:  0.028

*Representative data; results in individual laboratories may vary from these data.

ALTERNATE RESULT UNITS

1.  The default result unit for the ARCHITECT STAT Troponin-I assay is ng/mL.  When the alternate result unit, ug/L, is selected, the conversion factor used by the system is 1.0.

2.  Conversion Formula:  (Concentration in ng/mL) x (1.0) = ug/L.

## Measuring Range:

- The measurement range for the ARCHITECT STAT Troponin-I assay is 0 ng/mL to 50 ng/mL.

Dilutions

Specimens with an troponin-I value exceeding 50.00 ng/mL are flagged with the code "> 50.00" and may be diluted with the Automated Dilution Protocol or the Manual Dilution Procedure.

AUTOMATED DILUTION PROTOCOL

1.  If using the Automated Dilution Protocol, the system performs a 1:9 dilution of the specimen and automatically calculates the concentration of the specimen before dilution and reports the result.

2.  Specimens with a troponin-I value exceeding 468.50 ng/mL are flagged with the code "> 468.50" when run using the Automated Dilution Protocol.  These specimens may be diluted by following the Manual Dilution Procedure.

MANUAL DILUTION PROCEDURE

1.  Manual dilutions should be performed as follows:

a.  The suggested dilution for a troponin-I is 1:20.

b.  Prior to diluting the specimen, dispense several drops of ARCHITECT STAT Troponin-I Calibrator A into a clean test tube for use in the next step.

c.  Add 10 uL of the patient specimen to 190 uL of ARCHITECT STAT Troponin-I Calibrator A.

d.  The operator must enter the dilution factor in the Patient or Control order screen.  The system will use this dilution factor to automatically calculate the concentration of the sample before dilution.  This will be the reported result.  The concentration of the specimen diluted (before dilution factor is applied) should be 2.5 ng/mL or greater.

e.  For detailed information on ordering dilutions, refer to the ARCHITECT System Operations Manual, Section 5.

Precautions and Warnings

1.  For In Vitro Diagnostic Use.

This product requires the handling of human specimens.  It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with the OSHA Standard on Bloodborne Pathogens.  Biosafety Level 2 or other appropriate biosafety practices should be used for materials that contain or are suspected of containing infectious agents.

1.  The ARCHITECT STAT Troponin-I Assay Diluent and Conjugate contain a mixture of:  5-chloro-2-methyl-4-isothiazolin-3-one and 2-methyl-4-isothiazolin-3-one (3:1), which is a component of ProClin, and are classified per applicable European Community (EC) Directives as:  Irritant (Xi).  The following are the appropriate Risk (R) and Safety (S) phrases.

R43  May cause sensitization by skin contact.

S24  Avoid skin contact.

S35  This material and its container must be disposed of in a safe way.

S37  Wear suitable gloves.

S46  If swallowed, seek medical advice immediately and show this container or label.

2.  ARCHITECT i Trigger Solution contains sodium hydroxide (NaOH) and is classified per applicable European Community (EC) Directives as:  Irritant (Xi).  The following are the appropriate Risk (R) and Safety (S) phrases.

R41  Risk of serious damage to eyes.

S25  Avoid contact with eyes.

S26  In case of contact with eyes, rinse immediately with plenty of water and seek medical advice.

S35  This material and its container must be disposed of in a safe way.

S36/39  Wear suitable protective clothing and eye/face protection.

S46  If swallowed, seek medical advice immediately and show this container or label.

3.  Information for European customers:  For product not classified as dangerous per European Directive 1999/45/EC - Safety data sheet available for professional user on request.

4.  For a detailed discussion of safety precautions during system operation, refer to the ARCHITECT System Operations Manual, Section 8.

HANDLING PRECAUTIONS

1.  Do not use reagent kits beyond the expiration date.

2.  Do not pool reagents within a kit or between reagent kits.

3.  Prior to loading the ARCHITECT STAT Troponin-I Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.  For microparticle mixing instructions, refer to the PROCEDURE, Assay Procedure section of this procedure.

4.  Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity.  Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in your package insert.

5.  To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

6.  Prior to placing the septum on an uncapped reagent bottle, squeeze the septum in half to confirm that the slits are open.  If the slits appear sealed, continue to gently squeeze the septum to open the slits.

7.  Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

8.  Over time, residual liquids may dry on the septum surface.  These are typically dried salts which have no effect on assay efficacy.

9.  For a detailed discussion of handling precautions during system operation, refer to the ARCHITECT i System Operations Manual, Section 7.

## Limitations — interference

.  Cardiac troponin-I levels can be increased in any condition resulting in cardiac cell damage.  For MI diagnostic purposes, the ARCHITECT STAT Troponin-I results should be used in conjunction with other information such as cardiac marker results (e.g., CK-MB and/or myoglobin), ECG, clinical observations and symptoms, etc.

2.  Specimens from patients who have received preparations of mouse monoclonal antibodies for diagnosis or therapy may contain human anti-mouse antibodies (HAMA).  Such specimens may show either falsely elevated or depressed values when tested with assay kits that employ mouse monoclonal antibodies.  Additional clinical or diagnostic information may be required to determine patient status.

3.  Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.  The presence of heterophilic antibodies in a patient specimen may cause anomalous values to be observed.  Additional information may be required for diagnosis.

4.  Although the ARCHITECT STAT Troponin-I assay is specifically designed to minimize the effects of HAMA and heterophilic antibodies, assay results that are not consistent with other clinical observations may require additional information for diagnosis.

Contacts:

Abbott Diagnostics Technical Support:  1-877-4ABBOTT (1-877-422-2688)

Alternative method

Send to sister facility

References

ABBOTT STAT Troponin-I package insert

Abbott Laboratories

Diagnostics Division

Abbott Park, Illinois 60064

PN 840174/R3

Revision 07/04

Effective date

Effective date for this procedure: 01/10/18

Author

Compiled by Abbott Diagnostics

Revised by: Jessica Mercer MLS (ASCP)

Revised by:	 Elsa Escobar