- Instructions

Hold the pipette vertically when aspirating the liquid and place the tip only a few

millimeters into the liquid. When dispensing, pipette should be held at a 30° to 45° angle,

Insuring that the tip does not touch the micro tube. Bubbles do not affect the reaction.

- Equipment

- Biohit Electronic Pipettor

- MTS Pipette Tips

- Precautions

- Avoid contamination of the cone tip

- Avoid leaving the pipettor in its side with liquid in the tip which might seep back into the mechanism.

- Never strike the cone tip against the tray when mounting tips as this can damage several internal components.

- Do not use excessive force when mounting the tips as this can hamper the tip ejection.

- Do not drop the pipettor as this may cause excessive shock.

- Avoid rough handling, moderate pressure is all that is required when using the keyboard.

- Programs

Pipetting 50µl of RBC’s

50µl x 1 Biohit Program 1

50µ x 6 Biohit  Program 2

Pipetting 10µl of RBC’s

10µl x 6 Biohit Program 9

Pipetting 25µl of plasma

25µl x 3   Biohit Program 3

25µl x 12 Biohit  Program 5

References:

- Quick Guide for Biohit Electronic Pipettor

Revised By: Jessica Mercer,   SAH Clinical Coordinator

Revised by: Elsa Escobar