- Fit Testing

  - Rescheduled for this afternoon at 13:00 to 16:00

  - This will take place in the Staff Education and Development office

  - It will take 10-15 minutes per person; please go up two people at a time

  - Please remember nothing to eat drink or smoke for 30 minutes prior to evaluation

- Glucometre Ward Audit

  - We have once again failed our ward audit on the glucometre.

  - This is a bit embarrassing as we are considered to be the standard set for the rest of the facility.

  - The QC reagent WAS dated, but the test strips were not dated.

- Microscopes

  - Microscope service has been performed

  - A camera has been installed on the microscope in Microbiology; this camera can be moved around by removing the eye piece.

  - We are waiting to have the software installed on desktop workstations

  - Joel has used this type of camera before; Mike will consult with Joel for assistance in usage

- Fluorescent Microscope

  - Marcus, the microscope technologist, has pointed out that the cable coming from the light box to the micr5oscope needs to be kept as straight as possible.

  - As we have not maintained the cable properly it will need to be replaced.

  - Mike will follow up with Sherri once she returns

- Immersion Oil

  - Two of our 50x lenses have been seized or destroyed because of people not maintaining the lenses properly by removing the immersion oil from the objective

  - These cost $1200 apiece to replace

- LIS Down Time

  - Scheduled Oracle update 01:00 to 04:00 (overnight)

  - Mohammed is currently the scheduled on-call technologist

- Pizza Day Friday

  - Friday is Jessie’s last day and we will be having a pizza lunch.

  - Please see Carolyn to contribute

- Roundtable:

  - Jennifer:

    - The Vitros 5600 is still set to go LIVE, on August 12, 2015.

    - The ABL837 interface is still on schedule to go LIVE next week – this was originally scheduled for Monday but, that is a STAT holiday. The interface will go LIVE Tuesday or Wednesday: Until then please continue to use the manual reporting procedures outlined previously.

    - The Clinitek Advantus is scheduled to go LIVE September 16, 2015 to allow for sufficient time for procedure development and training: There are significant changes to the work process and will include the assignment of POCT devices such as the iSTAT and glucometre.

  - Amy: Wing night at 19:00 for Jessie’s going away event. We have a table reserved a Sam’s Monkey Tree

  - Mike: New parameters for smear review and send out have been written. Mike will have Amy review the document and will send it down to Dr. S. Brown to review for appropriateness. There will be

  - Erin: The virtual servers have for TEST, LIVE and UPG which were scheduled for Aug.12 but of course were performed the day April left for her three week holiday. Erin has been working on evaluating the upgrades and has not encountered any issues at this time.

  - Joel: Microbiology has asked for an anti-vibration pad for one of their microscopes. Carolyn followed up with Marcus about a suitable option; however, they are quite costly. As such, funding approval must be sought.