| Test | Sample Type | Ext. | AMR | Maximum Dilution | QC | Available STAT | Run Frequency |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ACET | Lithium Heparin Plasma (Minty Green) | -08 | 66 – 1323 umol/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| ALB | Lithium Heparin Plasma (Minty Green) | -08 | 10 – 60 g/L | Cannot be auto-diluted See IFU | CHEM1 & CHEM2 |  | Daily |
| ALP | Lithium Heparin Plasma (Minty Green) | -08 | 20 – 1500 IU/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| ALT | Lithium Heparin Plasma (Minty Green) | -08 | 6 – 1000 IU/L | 1:20 | CHEM1 & CHEM2 |  | Daily |
| AST | Lithium Heparin Plasma (Minty Green) | -08 | 3 – 750 IU/L | 1:20 | CHEM1 & CHEM2 |  | Daily |
| B-hCG | Lithium Heparin Plasma (Minty Green) | -08 | 2.39 -15000 IU/L | 1:400 | IMMUN1, IMMUN2 & IMMUN3 |  | Daily |
| BU/BC | Lithium Heparin Plasma (Minty Green) | -08 | 0 – 462 umol/L | 1:2 | CHEM1, CHEM2 & NBIL2 |  | Daily |
| CA | Lithium Heparin Plasma (Minty Green) | -08 | 0.25 – 3.49 mmol/L | 1:2 | CHEM1 & CHEM2 |  | Daily |
| CSF Glucose | Tube 1 Supernatant | -1C | 1.11 – 36.08 mmol/L | Do Not Dilute | CSF1 & CSF2 |  | Ad Hoc |
| CSF Protein | Tube 1 Supernatant | -1C | 0.27* – 3.00 mmol/L | 1:5 | CSF1 & CSF2 |  | Ad Hoc |

* AMR adjusted higher than vendor specifications due to validation recommendations

| Test | Sample Type | Ext. | AMR | Maximum Dilution | QC | Available STAT | Run Frequency |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CK | Lithium Heparin Plasma (Minty Green) | -08 | 20 – 1600 IU/L | 1:20 | CHEM1 & CHEM2 |  | Daily |
| CL | Lithium Heparin Plasma (Minty Green) | -08 | 50 – 175 mmol/L | Do Not Dilute | CHEM1 & CHEM2 |  | Daily |
| CARB | Serum (Red Top) | -13 | 12.7 – 84.6 µmol/L | 1:5 | IMMUN1, IMMUN2 & IMMUN3 |  |  |
| CREA | Lithium Heparin Plasma (Minty Green) | -08 | 4 – 1238 umol/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| CRP | Lithium Heparin Plasma (Minty Green) | -08 | 5-90 mg/L | Do Not Dilute | CRP1 CHEM1 |  | Daily |
| DIG | Serum (Red Top) | -13 | 0.51 – 5.12nmol/L | 1:5 | IMMUN1, IMMUN2 & IMMUN3 |  |  |
| ECO2 | Lithium Heparin Plasma (Minty Green) | -08 | 5 – 40 mmol/L | 1:2 | CHEM1 & CHEM2 |  | Daily |
| FERR | Serum (Gold Top) | -04 | 0.29 – 1000 ng/mL | 1:20 | IMMUN1, IMMUN2 & IMMUN3 |  | Monday to Friday |
| Free T3 | Serum (Gold Top) | -04 | 0.77 – 35.0 pmol/L | Do Not Dilute | IMMUN1& IMMUN2 |  | Monday to Friday |
| Free T4 | Serum (Gold Top) | -04 | 0.88 – 90.0 pmol/L | Do Not Dilute | IMMUN1& IMMUN2 |  | Monday to Friday |
| GGT | Lithium Heparin Plasma (Minty Green) | -08 | 10 – 1400 IU/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| GLU | Lithium Heparin Plasma (Minty Green) | -08 | 1.1 – 34.7 mmol/L | 1:5 | CHEM1 & CHEM2 |  | Daily |

| Test | Sample Type | Ext. | AMR | Maximum Dilution | QC | Available STAT | Run Frequency |
| --- | --- | --- | --- | --- | --- | --- | --- |
| HbA1c | EDTA (Lavender Top) | -02 | 4.00 – 14.00 % | Do Not Dilute | HBA1-1, HBA1-2 & HBA1-3 |  | Monday to Friday |
| HBSAB | Serum (Gold Top) | -04 | 4.23 – 1000 mIU/mL | Do Not Dilute | VIROCL & VIRO2 |  | Monday to Friday |
| HBSAG | Serum (Gold Top) | -04 | Negative <0.9 Referred >/= 0.9 | Do Not Dilute | VIROCL & VIRO1 |  | Monday to Friday |
| HCV | Serum (Gold Top) | -04 | Negative <0.9 Referred >/= 0.9 | Do Not Dilute | VIROCL & VIRO1 |  | Monday to Friday |
| HIV | Serum (Gold Top) | -04 | Negative <0.9 Referred >/= 0.9 | Do Not Dilute | VIROCL & VIRO1 |  | Monday to Friday |
| K | Lithium Heparin Plasma (Minty Green) | -08 | 1.0 – 14.0 mmol/L | Do Not Dilute | CHEM1 & CHEM2 |  | Daily |
| LDH | Lithium Heparin Plasma (Minty Green) | -08 | 100 – 2150 IU/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| LIPA | Lithium Heparin Plasma (Minty Green) | -08 | 10 – 2000 IU/L | 1:20 | CHEM1 & CHEM2 |  | Daily |
| LITH | Serum (Red Top) | -13 | 0.20 – 4.00 mmol/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| MG | Lithium Heparin Plasma (Minty Green) | -08 | 0.08 – 4.11 mmol/L | 1:2 | CHEM1 & CHEM2 |  | Daily |

| Test | Sample Type | Ext. | AMR | Maximum Dilution | QC | Available STAT | Run Frequency |
| --- | --- | --- | --- | --- | --- | --- | --- |
| NA | Lithium Heparin Plasma (Minty Green) | -08 | 75 – 250 mmol/L | Do Not Dilute | CHEM1 & CHEM2 |  | Daily |
| NTBNP | Lithium Heparin Plasma (Minty Green) | -08 | 11.1 – 35000 pg/mL | 1:10 | TROP1, TROP2 & TROP3 |  | Daily |
| PHOS | Lithium Heparin Plasma (Minty Green) | -08 | 0.16 – 4.20 mmol/L | 1:2 | CHEM1 & CHEM2 |  | Daily |
| PHYT | Serum (Red Top) | -13 | 11.9 – 158.4 umol/L | 1:5 | IMMUN1, IMMUN2 & IMMUN3 |  |  |
| SALI | Serum (Red Top) | -13 | 0.07 – 2.90 | 1:5 | CHEM1 & CHEM2 |  | Daily |
| SYPH | Serum (Gold Top) | -04 | Negative <0.8 Referred >/= 0.8 | Do Not Dilute | VIROCL & VIROSY |  | Monday to Friday |
| TBIL | Lithium Heparin Plasma (Minty Green) | -08 | 2 – 462 umol/L | 1:100 | CHEM1 & CHEM2 |  | Daily |
| TP | Lithium Heparin Plasma (Minty Green) | -08 | 20 – 110 g/L | 1:5 | CHEM1 & CHEM2 |  | Daily |
| TPSA | Serum (Gold Top) | -04 | 0.064 – 100 ug/L | 1:20 | IMMUN1, IMMUN2 & IMMUN3 |  | Monday to Friday |
| TRIG | Serum (Gold Top) | -04 | 0.11 - 5.93 mmol/L | 1:5 | Chem 1 & LIPID2 |  | Daily |
| TROPI | Lithium Heparin Plasma (Minty Green) | -08 | 0.012 – 80.0 ng/mL | 1:10 | TROP1, TROP2 & TROP3 |  | Daily |
| TSH | Serum (Gold Top) | -04 | 0.015 – 100.0 mIU/L | Do Not Dilute | IMMUN1, IMMUN2 & IMMUN3 |  | Monday to Friday |

| Test | Sample Type | Ext. | AMR | Maximum Dilution | QC | Available STAT | Run Frequency |
| --- | --- | --- | --- | --- | --- | --- | --- |
| UREA | Lithium Heparin Plasma (Minty Green) | -08 | 0.7 – 42.8 mmol/L | 1:2 | CHEM1 & CHEM2 |  | Daily |
| URIC | Lithium Heparin Plasma (Minty Green) | -08 | 30 – 1011 umol/L | 1:3 | CHEM1 & CHEM2 |  | Daily |
| Urine Albumin | Random Urine | -80 | 6.0-190.0 mg/L | 1:6 | URIN1 & URIN2 |  | Monday to Friday |
|  | 24 Hour Urine | -81 |  |  |  |  |  |
| Urine CREA | Random Urine | -80 | 0.11 – 30.63 mmol/L | Do Not Dilute | URIN1 & URIN2 |  | Monday to Friday |
|  | 24 Hour Urine | -81 |  |  |  |  |  |
| Vancomycin | Serum (Red Top) | -13 | 5 – 50 mg/L | 1:4 | CHEM1 & CHEM2 |  | Daily |