- Prepared by: D. Napert

- Adopted: 12/4/13   David Morris PhD____________________________

Annual Review

| Reviewed by Medical Director | Date | Reviewed by Medical Director | Date |
| --- | --- | --- | --- |

Revisions:

| New 12/4/13 |  |  |  |
| --- | --- | --- | --- |
| Change tube EDTA to LiHep 8/2019 DN |  |  |  |

Lifespan AMC-Department of Pathology

The Miriam Hospital ___		___ Rhode Island Hospital

164 Summit Avenue		593 Eddy Street

- Providence, Rhode Island 02906		                        Providence, Rhode Island 02903

This procedure is valid for the following chemistry analyzers:

| - AU480 | - AU680 |
| --- | --- |
| - AU 5800 |  |

# PRINCIPLE:

Ammonia is derived from the catabolism of amino acids and from the action of intestinal bacteria on dietary protein and is converted to urea in the liver hepatocytes and so rendered non-toxic.  Under normal circumstances the concentration of ammonia in the circulation remains low, typically less than 50 µmol/L (85 µg/dL).  Studies have shown that excess ammonia can have a toxic effect on the central nervous system and clinical manifestations are typically neurological disturbances.1,2,3

Elevated levels of ammonia may be either due to inborn errors of metabolism or as secondary to other conditions.

Inborn errors of metabolism are the major cause of elevated ammonia in infants and usually the result of urea cycle enzyme deficiencies.  Inherited disorders affecting the metabolism of the dibasic amino acids (lysine and ornithine) and those involving the metabolism of organic acids may also produce elevated levels of circulating ammonia.  Elevated ammonia may also be observed in severe liver failure as may occur in Reye’s Syndrome, viral hepatitis, or cirrhosis.1,2,3

# INTENDED USE:

System reagent for the quantitative determination of in ammonia (NH3) human plasma on Beckman Coulter AU Clinical Chemistry analyzers.

# METHODOLOGY:

A number of methods have been developed for the estimation of plasma ammonia and these can be broadly classified into either indirect or direct methods.

In the indirect procedures, ammonia is first of all isolated, for example by the addition of alkali or the use of a cation exchange resin.  After that it is measured colorimetrically by nessierization or Berthelot reaction.  These procedures are not easily automated or require dedicated equipment.

Direct procedures, such as enzymatic methods, are more widely used in routine laboratories; as they do not require the separation of ammonia from the specimen prior to the analytical step.  Direct procedures are therefore more easily automated.  The Beckman Coulter AU System Ammonia reagent is a direct enzymatic procedure based on the following reaction sequence.1

NH4    + α-ketoglutarate + NADH     (GLDH)    Glutamate + NAD + H2O

The reagent contains LDH in excess, to rapidly reduce endogenous pyruvate so that it does not interfere with the assay system.  The Beckman Coulter AU Ammonia reagent also incorporates a patented stabilization process which renders the reagent stable in the liquid phase.

# SPECIMEN:

## Patient Preparation:

None required.  Plasma collected by standard venipuncture techniques.

## Type:

Sample Volume 17.0 uL

Plasma, collected with lithium heparin (not ammonium heparin) is the acceptable specimen.  EDTA plasma is not acceptable unless the tube is completely filled and there is a consult with Medical Director or charge person.  Result discrepancies have been identified when using EDTA tubes that are not full volume.

Ideally, the collection tube should be completely filled with blood and immediately placed on ice.  Centrifuge (cold) the sample as soon as possible and separate the plasma.  Keep samples at 2-8°C until analysis.1

## Handling Conditions:

Ammonia samples are stable for 4 hours (spun) at 2-8°C or for 48hours when frozen < -20°C.1

Tubes should be filled completely, mixed gently by inversion, placed on ice, centrifuged immediately for 10 minutes. If the spun sample is not able to be processed within 4 hours, the plasma may be frozen. The tubes should be tightly sealed at all times.

Criteria for Unacceptable Specimens

The sample must be properly labeled with the minimum of the patient’s name and Date of Birth.  For more information on the acceptability of samples, see section “Specimen” at beginning of this policy.

Samples must be received on ice and centrifuged within 90 minutes.  These samples, once spun are stable for 4 hours.

If the specimen arrives at room temperature, or left unspun greater than 90 minutes, the sample is unacceptable and cannot be analyzed.  The unit should be notified and a footnote entered into the system.

Serum samples are unacceptable and should be treated as above.

Due to interferences with the analysis, testing  on samples with  a HEMOL index ≥2 will not be performed.  The result should be entered as “Hemolyzed”, the floor notified and a footnote including the name of the person taking the information, the date and time of the call, and the tech’s ID.

Moderate levels of lipemia may also interfere with the assay.  Samples should not be ultra-centrifuged due to the stability of the analyte.

# EQUIPMENT AND MATERIALS:

## Equipment:

Beckman Coulter AU400/AU400e, AU480, AU600, AU640/AU640e, AU680, AU2700, and AU5400 analyzers.

## Materials:

Beckman Coulter AU System Ammonia Reagent

Final concentration of reactive ingredients:

| α-Ketoglutarate | 7.5 mmol/L |
| --- | --- |
| NADH | > 0.2 mmol/L |
| GLDH (Micro-organism) | > 4000 U/L |
| LDH (Micro-organism) | >30,000 U/L |
| Tris buffer | 100 mmol/L |

Also contains preservatives.

Reagent storage location in this laboratory: Reagents are stored in the Chemistry refrigerator

Test tubes 12 -16 mm in diameter or sample cups 	(Cat No. AU1063).

Storage location of test tubes or sample cups in this laboratory: Sample tubes and cups are stored in the main chemistry lab area

Beckman Coulter AU System Ammonia Standard (included in the kit)

Active Ingredient: Ammonium sulphate 		59 µmol/L (100 µg/dL)

Storage location of the calibrator in this laboratory: Calibrators are stored in the Chemistry refrigerator

### Precautions:

- For in vitro diagnostic use.

- Do not ingest. Harmful if swallowed.

- Avoid contact with skin and eyes.  If spilled, thoroughly wash affected areas with water.

- Contains sodium azide as a preservative that may react with lead joints in copper plumbing to form explosive compounds. Even though the reagent contains minute quantities of sodium azide, drains should be well flushed with water when discarding the reagent.

## Preparation:

The Beckman Coulter AU System Ammonia Reagent is liquid, ready for use. No preparation is needed.

### Storage Requirements:

1.	The unopened reagents are stable until the expiration date printed on the label when stored at 2 - 8°C.

- Opened reagent is stable for 14 days when stored in the refrigerated compartment of the analyzer.

- Opened Ammonia Standard is stable until the expiration date printed on the label as long as the bottles are capped when not in use and stored 2 - 8C.

### Indications of Deterioration:

Discoloration of the reagent, visible signs of microbial growth, turbidity or precipitation in reagent may indicate degradation and warrant discontinuance of use.

# PERFORMANCE PARAMETERS:

The following data was obtained using this Ammonia Reagent on Beckman Coulter AU analyzers according to established procedures. Results obtained at individual facilities may differ.

## Precision:8

Estimates of precision, based on CLSI recommendations7, are consistent with typical performance. The within run precision is less than 5%CV and total precision is less than 5%CV. Assays of control sera were carried out and data reduced following CLSI guidelines.

| N=20 | Within run |  | Total |  |
| --- | --- | --- | --- | --- |
| Mean, μmol/L | SD | CV% | SD | CV% |
| 116 | 5.0 | 4.3 | 5.1 | 4.4 |
| 293 | 4.3 | 1.5 | 8.4 | 2.8 |

## Method Comparison:8,9

44 patient samples were used to compare this Beckman Coulter AU Ammonia Reagent (Method 1) versus Arkray ammonia method (Method 2).

| Y Method | Beckman Coulter |
| --- | --- |
| X Method | Arkay |
| Regression Equation | Method 1 = 1.06x + 2.7 |
| Correlation Coeff. (r) | 0.9373 |
| No. of Samples (n) | 44 |
| Range | 17-122 μmol/L |

### Sensitivity:

When run as recommended the lowest detection limit is 10 μmol/L or 17 μg/dL

The lowest detection limit (LDL) for this method was determined by assaying replicates of analyte free water.  The mean and standard deviation were determined and LDL was calculated using the formula:

LDL = X + (2 x s)

Where	X = mean value of replicates

S = standard deviation of replicate (n-1)

# CALIBRATION:

## Standard Preparation:

Perform a single-point calibration (AB) using a water blank (blue rack) and the Ammonia kit standard in a yellow calibration rack.  The frequency of calibration is 7 days. Calibration of this procedure is accomplished by use of the Beckman Coulter AU System Ammonia Standard provided in the kit.  The standard has been verified against an in-house Ammonium Standard Solution which is traceable to Standard Reference Material from the US National Institute of Standards and Technology (NIST).

## Calibration Procedure:

Recalibration of this test is required when any of these conditions exist:

- A reagent lot number has changed or there is an observed shift in control values.

- Major preventative maintenance was performed on the analyzer.

3.	A critical part was replaced.

# QUALITY CONTROL:

During operation of the Beckman Coulter AU analyzer at least two levels of an appropriate quality control material should be tested a minimum of once a day. In addition, controls should be performed after calibration, with each new lot of reagents, and after specific maintenance or troubleshooting steps described in the appropriate AU User’s Guide. Quality control testing should be performed in accordance with regulatory requirements and each laboratory’s standard procedure.

Location of controls used at this laboratory. Controls are located in Chemistry refrigerator

# ANALYZER PARAMETERS:

A complete list of test parameters and operating procedures can be found in the appropriate User’s Guide and at www.beckmancoulter.com.

# CALCULATIONS:

Parameters are set for reporting in μmol/L.  To work in μg/dL, multiply the results by 1.7.

# REPORTING RESULTS:

## Reference Ranges:

Adult (Lifespan):	2-50 μmol/L

Critical Orange Value:    >100 umol/L

Reportable Range <10 - >600 umol/L

Expected values may vary with age, sex, diet and geographical location. Each laboratory should determine its own expected values as dictated by good laboratory practice.

Hemolysis Index:

If hemolysis index is 1, report results with comment @HEM2

If hemolysis index is >1, cancel as .HEMOLYZED and call.

## Procedures for Abnormal Results:

Abnormal results are flagged by the listed analyzers according to the normal values entered by the user into the instrument parameters.

## Reporting Format:

Results are automatically printed for each sample in μmol/L at 37°C.

All Critical Orange results must be called back to the requesting location and reported within 8 hours.  A footnote must then be entered which includes the person to whom the result was reported, the date and time the result was reported and the tech who reported the value.  All results must be read back to the caller to include the full name of the patient and the patient’s DOB and the footnote noted as such.

# LIMITATIONS:

The Beckman Coulter AU System Ammonia procedure is linear from 10 to 600 μmol/L (17-1020 μg/dL). Samples exceeding the upper limit of linearity should be diluted with ammonia free water and repeated. The sample may be diluted, repeated and multiplied by the dilution factor automatically utilizing the AUTO REPEAT RUN.

## Interfering Substances:

- Hemolyzed samples should not be used as erythrocytes contain levels of ammonia approximately 3 times that of plasma.1

- No interference from pyruvate was observed up to a level of 0.75 mmol/L or 0.01 mg/dL.

- No interference from ALT was observed up to a level of 4000 U/L.

- Bilirubin: No significant interference up to 17.4 mg/dL unconjugated Bilirubin.

- Lipemia: No significant interference up to 50 mg/dL (Intralipid).

- Reliable estimations of ammonia can only be achieved if steps are taken to avoid contamination from ammonia.  Sources of contamination include, but are not restricted to cigarette smoking (patient and collection staff), laboratory atmosphere, and laboratory glassware.

- Young DS4 has published a comprehensive list of drugs and substances, which may interfere with this assay.

The information presented is based on results from Beckman Coulter studies and is current at the date of publication. Beckman Coulter Inc., makes no representation about the completeness or accuracy of results generated by future studies. For further information on interfering substances, refer to Young4 for a compilation of reported interferences with this test.

# REFERENCES:

1.	Clinical Chemistry Infobase: A Scientific & Management Cyclopedia.  Pesce-Kaplan Pulbishers 1996: 2246-2320.

2.	Tietz Textbook of Clinical Chemistry.  Burtis CA and Ashwood ER (eds).      2nd Edition, WB Saunders Company, 1994; 32:1485-88.

3.	The Diagnosis of Urea Cycle Disorders. Lab Medica International, May/June 1993; 13-17.

4.	Young, D.S., Effects of Drugs on Clinical Laboratory Tests, 3rd edition, 1990; 3: 30-2.

5.	Ektachem Multilayer Dry Film Assay for Ammonia Evaluated. J. Clin Chem 1985: Vol 31: 12:2012-2014.

6.	Wachtel M et al, Creation and Verification of Reference Intervals. Laboratory Medicine 1995; 26:593-7.

7.	Clinical and Laboratory Standards Institute.  User evaluation of Precision Performance of Clinical Chemistry Devices.  CLSI/NCCLS, 1984, CLSI/NCCLS Publication EP5-T.

8.	Meller A. “Evaluation and Implementation of an Alternative Method for Measurement of Blood Ammonia”.  Department of Clinical Biochemistry, Newcastle upon Tyne Hospitals NHS Trust, 2006.

9.  Data is on file for specific AU analyzers.