**PROCEDURE: CSF SIMPLEXA HSV 1 & 2 DIRECT PCR**

1. **PRINCIPLE**

The Simplexa HSV 1 & 2 Direct assay system is a real-time PCR that enables the direct amplification, detection and differentiation of Herpes Simplex Virus (HSV) -1 and/or HSV-2 DNA from unprocessed Cerebral Spinal Fluid (CSF) specimens without nucleic acid extraction. The system consists of the Simplexa HSV 1 & 2 Direct, the 3M Integrated Cycler (with 3M Integrated Cycler Studio Software), the Direct Amplification Disc, and associated accessories.

In the Simplexa HSV 1 & 2 Direct assay, bi-functional fluorescent probe-primers are used together with corresponding reverse primers to amplify HSV-1, HSV-2 and internal control targets. Well conserved regions of the HSV-1 and HSV-2 DNA polymerase genes are targeted to identify HSV-1 and HSV-2 DNA respectively in the specimen. An internal control is used to detect PCR failure and/or inhibition.

1. **AVAILABLILITY**

M-F: 7:30AM-2:00PM

1. **TEST CODE**

CSHSV

1. **SPECIMEN**
	1. CSF

**Remnant specimen used only if no sterile specimen exists. Clinical parameters and consultation with lab director used to determine significance of results in this situation**

1. **MATERIALS AND EQUIPMENT**
2. Materials
	1. Simplexa HSV 1 & 2 Positive Control Pack (MOL2160)
	2. 200uL pipette for specimen transfer and 100uL pipette for Reaction Mix transfer
	3. Sterile, nuclease free disposable pipette tips with filters (Art XL P-200 and 100uL)
	4. Direct Amplification Disc Kit (MOL1455) used on the 3M Integrated Cycler
	5. DiaSorin Molecular LLC Simplexa HSV 1 & 2 Direct Reaction Mix kit (MOL2150)
	6. Disposable, powder free gloves
	7. Synthetic CSF (HSP-515) No Target Control (NTC) purchased from SeraCare
3. Equipment
	1. 3M Integrated Cycler with 3M Integrated Cycler Studio Software version 6.0 or higher
	2. Freezer (manual defrost) at -10 to -30 oC
	3. Refrigerator at 2 to 8 oC
	4. Sartorius 200uL pipette for specimens
	5. Eppendorf 100uL pipette for reaction mix
4. **STORAGE AND HANDLING**
5. Specimen should be transported to the lab immediately THEN stored at 2-8 oC until it is processedB. If there is a greater than 7 day delay in running, the specimen should be held at -70 oC
6. DiaSorin Molecular LLC Simplexa HSV 1 & 2 Direct Reaction Mix kit (MOL2150) and Simplexa HSV 1 & 2 Positive Control Pack (MOL2160) should be immediately stored in a -10 to -30 oC manual frost freezer
7. Direct Amplification Disc Kits (MOL1455) can be kept at room temperature (18-25 oC)
8. Do not refreeze or vortex Reaction Mix
9. Do not use Simplexa Kits or any component of this test past its expiration date
10. **QUALITY CONTROL/ QUALITY ASSURANCE**
11. Commercially purchased external *Positive Control* and *No Target Control* are run once daily. The control result is recorded on the CYCLER HSV DAILY QC RECORD found in the binder by the cycler. No patient results will be released unless controls are valid and perform as expected. Notify Senior Medical Technologist or Manager of invalid control results. Repeat entire run (patients and controls) Positive controls, once thawed, must be used within 24 hours
12. New lots and/or new shipments of DiaSorin Molecular LLC Simplexa HSV 1 & 2 Direct Reaction Mix kits (MOL2150) must be QC’d using the commercially purchased controls. Results are recorded on the NLNS Cycler Simplexa HSV sheet found in the binder by the cycler
13. Use Simplexa HSV 1 & 2 Positive Control Pack (MOL2160) and AcroMetrix HHV Negative Control (Thermo Scientific) for the positive and negative control, respectively
14. Environmental wipe testing is performed monthly. Areas are swabbed and run as test patients. Refer To **Appendix ENVIRO** for detailed directions.
15. Positivity rate is monitored on a monthly basis
16. All results must be entered, verified then rechecked against the Simplexa printout before finalizing results. A report must be printed and given to a Senior Medical Technologist along with the tasklist for final review
17. Periodic Maintenance is done annually by a Field Application Scientist from DiaSorin Molecular LLC. A call will be made to the lab to schedule this PM
18. **TEST PROCEDURE**

**USE ONLY CLEAN UNGLOVED HANDS TO SET UP THE INSTRUMENT**

1. Make a tasklist for all pending specimens (Refer to Appendix A for a quick set up checklist).
2. Label one (1) snap cap tube for each specimen being set up
3. 3M Integrated Cycler set up
4. Turn on the instruments in the following order: cycler, computer, then printer
5. Log onto the computer using *computer user* and *integratedcycler* as the user name and password, respectively
6. Double click on the *Integrated Cycler Studio* icon in the center of the screen. Log on using *Virology* and *rihvirolab* as the user name and password, respectively
7. Block the large barcode and scan the small 3D barcode found on the card that corresponds to the in-use lot of Reaction Mix. The card can be found in the FOCUS QC binder on the counter by the cycler
8. Open a new sterile disc. Carefully, without touching the bottom of the disk, turn it inside the pouch until the barcode shows on the edge. Scan the barcode. Put on gloves
	1. Alternatively, a used disc with an adequate number of remaining slots can also be used
9. One by one, scan the specimens to be run in the order of the tasklist. Remove gloves
10. For controls
	1. Type *Positive* for the positive control and change the type to PC-HSV
	2. Type *Negative* for the negative control and change the type to NTC
11. Open the lid of the cycler by pressing the grey button on the front.
12. Bring the disc and a cooling plate to the biosafety hood
	1. Cooling plates are kept refrigerated (top shelf in a properly labeled box)
13. Specimen/hood set up. Put on gloves

**CHANGE GLOVES IMMEDIATELY IF THEY BECOME CONTAMINATED BY SPECIMEN**

* 1. Bring the specimens over to the molecular hood and quick vortex
	2. Using graduated disposable transfer pipettes, transfer spinal fluid from the primary container to the labeled snap cap tube
		1. Bring CSF to the top of the second narrowest portion of the pipette and dispense at the bottom of the snap cap tube. See graphic below

 

* 1. Change gloves
	2. Place the HSV-VESICLE and REACTION MIX pipette along with the appropriate tips into the hood. Use 200uL Art XL tips for specimen
	3. From the -20 oC freezer obtain the positive control, negative control and Reaction Mix. Place tubes in a mini tube rack
		1. Each positive control vial can be used twice. Mark the cap to indicate the control has been used once
		2. Select one vial of Reaction Mix for each sample including controls

1. Disc inoculation
	1. Remove a tip from the tip box and use the narrow end to gently lift the tabs away from the disc. Place tip in sharps container in hood
	2. Work with only one specimen/control opened at a time
	3. Starting at disc space 1, peel the foil back to reveal the two wells to be used being careful not to remove the foil entirely from the disc. Do not touch the sticky foil underside. See figures below
	4. Open specimen #1 and add 50uL of fluid using the 200uL pipette and XL tips to the front well labeled **SAMPLE**
	5. Re-cap the specimen
	6. Add 50uL of reaction mix using the 100uL pipette and 100uL tips to the back well labeled **R.**
		1. Reaction Mix tubes are single use
		2. DO NOT VORTEX Reaction Mix Tubes
	7. Replace the foil seal over the wells being careful not to touch the sticky surface. Press out any bubbles or wrinkles. Remove tab at perforations.
	8. Repeat steps c-g for each specimen and controls





1. Starting the run
	1. Carry the cooling plate with the disc to the cycler
	2. Place the disc in the cycler. Remove gloves
	3. Shut the lid, and use the mouse to press the run button on the screen
	4. Wipe the cooling disc with alcohol and return the disc to the refrigerator
2. Post analysis
	1. Select the print button on the screen; check the “CT” box at the bottom of the screen to print an entire report
	2. Open the lid of the cycler by pressing the grey button on the front of the instrument. With a gloved hand, remove the disc and place in the biohazard bag. Remove glove
	3. Shut down the instruments in this order: the computer, cycler, and then printer.
	4. ***Lightly*** wet gauze or Wipeall with alcohol and wipe the keyboard and inside of the cycler
	5. Close the lid of the cycler and the laptop
	6. Using a ***lightly*** dampened gauze or Wipeall clean the outside surfaces of the laptop, cycler, and printer with bleach, DI water and 70% alcohol in that order
	7. Positive specimens are stored in the Virology -70 oC freezer.
	8. All negative spinal fluids are placed in the refrigerator in the Chemistry Lab
3. **INTERPRETATION**
4. “Detected” result indicates the presence of HSV-1 and/or HSV-2 DNA in the patient sample
5. “Not Detected” result indicates the absence of HSV-1 and/or HSV-2 DNA in the patient sample
6. “Invalid” results indicate the inability to determine presence or absence of HSV-1 and/or HSV-2 DNA in the patient sample. This result may be due to: (1) DNA Internal Control (DNA IC) failure or (2) failure to detect sufficient specimen. An invalid sample needs to be retested. If retesting does not resolve the issue, refer to Senior Medical Technologist
7. “EC500” result indicates a data quality error for the particular vial analyte(s). The software was unable to determine a valid amplification for that analyte(s).
	1. If repeat is invalid, notify Director, Asst. Director, and/or Senior Medical Technologist
	2. Fill out “FOCUS Simplexa Invalid Record” sheet in QC Binder
	3. Repeat specimen
	4. Senior Medical Technologist will notify DiaSorin Molecular, LLC Customer Service. See contact information below
8. Refer to Soft Resulting (Appendix B) for complete instructions about entering results
9. **LIMITATIONS**
10. Contamination of patient specimens or reagents can produce erroneous results. Use good laboratory practices and control workflow
11. Deviations from the package insert or the use of times or temperatures other than those specified may give invalid results
12. Assay setup should be performed at room temperature (approximate range 18 to 25 oC).
13. Use appropriate fixed volume pipettes or equivalent for the addition of sample and reaction mix to the disc
14. Avoid touching the underside of the foil that will be in contact with the wells and disc surface which may cause contamination
15. To prevent potential erroneous results, make sure the sample and Reaction Mix is added to the corresponding well
16. To prevent contamination finish loading and applying adhesive foil cover to one set of Sample and Reaction wells before opening the foil of adjacent set(s) of Sample and Reaction wells
17. Initiate the run within 30 minutes of removing the Reaction Mix vial from the freezer
18. Do not attempt to remove adhesive foil cover wedges that have been used or attempt to re-use Sample and Reaction ports that have been used in previous runs
19. If kit contents or packaging appear to be broken or damaged, do not use and contact DiaSorin Molecular LLC. Refer to last page of procedure for contact information.
20. The spectral matrix must be installed in each 3M Integrated Cycler and should not be changed unless an updated QR code for the instrument is provided by DiaSorin Molecular LLC. The spectral matrix is unique to each 3M Integrated Cycler. The spectral matrix was provided with the 3M Integrated Cycler instrument on the cover of the 3M Integrated Cycler Hardware Manual. If the matrix label will not scan or cannot be found contact DiaSorin Molecular LLC. Refer to last page of procedure for contact information
21. Failure to install or changing the spectral matrix can result in false results
22. The detection of viral nucleic acid is dependent upon proper sample collection, transport, handling, and storage. Failure to observe proper procedures in any one of these steps can lead to incorrect results
23. False-negative results may occur if the viruses are present at a level that is below the analytical sensitivity of the assay or if the virus has genomic mutations, insertions, deletions, or rearrangements or if performed very early in the course of illness
24. When very high levels of HSV-2 are present with very low levels of HSV-1, the signal from the HSV-1 reaction may not be adequate to be detected, due to competitive interference
25. The performance if this test has not been established for screening of blood or blood products for the presence of HSV or for use with samples other than CSF
26. The performance of this test has not been established for monitoring treatment of HSV infection of the CNS
27. **NOTES**
28. Information on the Simplexa HSV 1 & 2 Direct Reaction Mix vial can only be transferred into the 3M Integrated Cycler Studio through a bar-code scanner. If the scanner is not working, or if you are unable to transfer the information for any reason, contact DiaSorin Molecular LLC Technical Services
29. Wear protective equipment, such as (but not limited to) gloves and lab coats when handling kit reagents, controls, and patient specimens. Wash hands thoroughly when finished running the test
30. Treat all specimens and discs as capable of transmitting infectious agents.
31. Discs may be reused until all 8 wedges have been used. Dispose of used discs without detaching foil cover in biohazard waste container
32. After each use, store discs flat with the numbered foil side up
33. Reaction Mix contains > 1% glycerol, which may cause irritation upon inhalation or skin contact. Upon inhalation or skin contact, first aid measures should be taken
34. **TECHNICAL SUPPORT**

For Technical Assistance: Phone 1-800-838-4548

 Fax 1-562-240-6526

1. **REFERENCES**
2. DiaSorin Molecular LLC Simplexa HSV 1 & 2 Direct Package Insert
3. CID 2013:57 Guide to Utilization of the Microbiology Lab *table II-1*
4. **REVISIONS**
	1. 02/10/2021 Do not need to repeat CT values >35

**APPENDIX A – Simplexa HSV 1 & 2 Checklist**

* Run a pending report, find any missing specimens, set up a tasklist
* Number specimens according to tasklist
* Label a snap cap tube for each specimen in a large CSF collection tube
* Turn on the cycler, then the computer, then the printer.
* Log into Windows
* Log into Cycler program
* Scan the small barcode on the Master Mix ID card
* Open a new disk or find a used disk for use with test. Scan.
* Scan specimens
* Enter control information changing the type where appropriate
* Bring specimens to hood, vortex, dispense 50mL into snap cap tube
* Retrieve pipette, XL tips, reaction mix tubes, controls and cooling plate. Place in hood.
* Place disk on cooling plate
* Use a tip to lift the tabs from the disk
* One by one, peel back the foil wedge, place 50uL of specimen and reaction mix to the appropriate wells and replace foil wedge. Remove tab at perforations
* Carry disk and cooling plate to the cycler
* Place the disk on the cycler
* Close cover, hit run
* When test is complete, hit the print button, check the three boxes at the bottom of the screen, and select the print button.
* Remove used disk, place in biohazard bag
* Shut down computer, printer and cycler
* Clean inside of cycler, computer keyboard and cooling plate with alcohol only. Close lids
* Clean the exterior surfaces and pipette with bleach, water and alcohol lightly sprayed on gauze or wipe-all. Do not use excess liquid.
* Clean hood and tip box as normal

**APPENDIX B – Soft Resulting for FOCUS HSV CSF and Vesicles**

**REFER TO CRITICAL RESULTS NOTIFICATION FOR PHYSICIAN CALLING POLICY**

1. Click on the PRINT button on the FOCUS computer
2. Click “show CT”
3. Review CT values and QC
4. Print report off the FOCUS computer and close screen.
5. If results and QC are ok, click EXPORT LIS button.
6. Resulting negative HSV samples
7. All negative results will automatically upload to Soft and will post to patient order number and autoverify.
8. Resulting Positive HSV samples
	1. From SoftLab, go to “interfaces”, and “Instrument Menu”
	2. Select “RFOCS” “Focus cycler” (#67)
	3. Select “Loadlist and todays results”, “Not Posted”, “By Sequence”
	4. Each order will be highlighted individually. Verify the result against the instrument printout.
9. Phone reports:
	1. Highlight the order number on left of screen
	2. At bottom of screen click on *Lab Results*
	3. Open “Comment” box and add comment/phone report using @CALM
	4. Click back to *Instrument* tab and save when asked
	5. Click **Post All** to verify the report
	6. Order number should disappear from list on left
10. Click *Verify All*
11. Click SAVE
12. Go to next specimen. Repeat steps 7-10 until all samples are resulted

**NOTE:** *All Invalids/ERROR specimens will upload to soft. Please do not upload these runs to LIS. Manually result by following these steps*

1. Double click *Resulting Worklist* Icon in *Softlab*
2. Choose Tasklist from the *Select Tests By* drop-down list
3. Type in the Tasklist ID in the appropriate box-Select OK
4. In the left column, highlight the first specimen in the tasklist
5. In the resulting area, select one of three choices for both HSV 1 and HSV 2 (refer to next page for visual reference)
	1. For HSV 1 POSITIVE
		1. Select Detected (2) from the HERP keypad to the right of CHSV1 line
		2. Select Not Detected (1) from the HERP keypad to the right of CHSV2 line
		3. The footnote on line 3 will self-populate
	2. For HSV 2 POSITIVE
		1. Select Not Detected (1) from the HERP keypad to the right of CHSV1 line
		2. Select Detected (2) from the HERP keypad to the right of CHSV2 line
		3. The footnote on line 3 will self-populate
	3. For an invalid test
		1. Select Invalid@HSVI (3) in both line 1 and line 2
		2. The footnote on line 3 will self-populate
6. If a call must be made, enter the information in the comment box
	1. Open the comment box
	2. Type @CALM to populate the ”called to” template
	3. Enter the nurse you spoke to along with the time and date of the call
	4. Press OK
7. Click *Verify All*
8. Click SAVE
9. Go to next specimen. Repeat steps 15-19 until all samples are resulted
10. Once complete, a report must be printed
	1. Select the Print Icon
	2. Choose Worklist from the print menu
	3. Under Layout select RE\_TASKREP, click OK
	4. Give cycler printout, Tasklist and printed Soft report to Senior Medical Technologist for review