PROCEDURE: Microbiology Safety Guidelines (NPH)

- Principle

  - The Microbiology Safety Guidelines provides essential standards to ensure a secure work environment at NPH. By mandating proper PPE, strict disinfection, and standardized waste disposal, these guidelines minimize infection risks and protect patient privacy. Additionally, specific protocols for Biological Safety Cabinets and chemical labeling ensure equipment integrity and personnel safety through consistent monitoring and clear hazard communication.

- Procedure

  - Microbiology General Safety

    - Reference:

      - Laboratory-wide Policy: Safety Preparedness Binder (NPH)

    - Departmental Personal Protection

      - These guidelines apply specifically to the Microbiology laboratory.

      - Gloves and disposable lab coats/gowns must be worn during processing, instrument maintenance, troubleshooting, and cleaning.

      - Protective eyewear is available, and its use is strongly recommended when performing procedures with a risk of splattering.

      - The department-specific eyewash station is in the center of the main lab.

      - Protective clothing i.e., gloves, disposable lab coats/gowns, are not to be worn outside the laboratory and once removed, hands should be washed.

      - Always wash hands before leaving the laboratory. There are separate hand-washing and testing sinks in the laboratory. There are also hand sanitizer dispensers at the entrance of each room. Remember, everything in the laboratory is considered contaminated, including fomites such as telephones, pens, chairs, etc.

      - Lab personnel may not eat, drink, or apply cosmetics while in the laboratory.

      - Never mouth pipette.  Use automatic pipettes or controller bulbs.

      - In the event of accidental exposures to infectious samples, hands or other exposed areas must be washed with soap and wash or antiseptic solution immediately.

      - In case of eye exposure to infectious materials, and or chemical exposure immediately flush eyes at the emergency eye wash station located at the sink area with copious amounts of water for 15 minutes.

      - Within the first hour of the accidental exposure, report the exposure to the laboratory manager and EMPOLYEE HEALTH SERVICES (EHS) for evaluation and follow-up.

    - Equipment & Bench Maintenance

      - Microscope Care

        - Stages and lenses must be kept free of oil

        - Use ONLY lens paper on glass surfaces

        - Optical Lens cleaning solution is available for cleaning of the glass surfaces of the microscope

        - Kim wipes may be used for all other surfaces of the microscope

        - Ethanol may be used to clean non-glass microscope surfaces

    - Bench Disinfection

      - Benches:

        - Clean with approved disinfectant (ie.  Sani-cloth wipes)

      - Hoods:

        - Use 10% bleach, followed by DI water, followed by 70% ethanol.

        - If there is a suspected case of CJD, the processing area must be cleaned with 1N NaOH and allowed to set for 1 hour before washing thoroughly with soap. Instrumentation should follow the same procedure.

    - Hand Off Communication

      - When information needs to be ‘handed off’ to the next shift, the staff will utilize Microsoft Teams or written Communication Log at the receiving desk.

  - Infectious and HIPAA Waste Workflow

    - Biohazard Non-Sharp Waste (Red Bag)

      - Discard all non-sharp biohazard waste into large grey bins double lined with red biohazard bags.

    - Sharps Disposal

      - Sharps containers - large and small containers available from ESD

      - Sharps (needles, scalpels, broken glass) are disposed of in designated sharps containers.

      - Containers must have 10-20% empty space (or be no more than 2/3 full)

      - Containers must always remain capped except when waste is being added.

    - HIPAA/Confidential Waste

      - Any paper or LIS waste containing HIPAA information must be placed in the HIPAA shred barrels located in the front of the main lab.

    - General Waste (Non-HIPAA/Non-Biohazard)

      - Standard paper towels and non-recyclable paper without patient information go into clear/white bags.

    - Service Requests and Safety Contacts

      - Waste Removal: Call ESD at 401-255-3545 once the container is full and labeled for removal.

      - Environmental Safety: For questions or to report spills, contact the Environmental Safety Manager (401-255-2758).

      - Emergency: Dial ext. 5-3333 for all uncontrollable spills and report a Code Orange.

  - Specialized Radiation (UV & Fluorescent)

    - UV Sterilization

      - It is strictly prohibited to work in the biological safety hoods while UV lights are active.

  - Microbiology Biological Safety Cabinet (BSC) Operational Guidelines

    - Reference

      - Emergency Response/Hazardous Waste Contingency Plan for Hazardous Materials Releases, Hazardous Materials Management-HM-2 and Chemical Hygiene Plan-TMH on Policy Stat

    - Technical Specifications & Inventory

      - The laboratory utilizes Class II BSC, Type A or A2.

    - Specimen Handling

      - All specimen processing must occur on or above the work tray (depressed area).

      - The front grill is designated as a "non-safe" area.

    - Airflow Integrity:

      - Never cover hood grates.

      - Avoid rapid or excessive arm movements in and out of the cabinet.

      - Use electric incinerators only; open flames are strictly prohibited.

    - Biohazard Spill Response inside the BSC

      - Do not turn off the cabinet during a spill or during the cleanup process.

      - Decontamination:

        - Pour ready mixed 10% bleach and let sit for at least 15 minutes.

        - Follow up with DI water and 70% alcohol.

    - Specialized Maintenance & Monitoring

      - Airflow must be monitored and recorded every day the unit is in use.

      - Cleaning & Material Care:

        - Surfaces must be disinfected at the start of every shift.

          - Use 10% bleach followed by DI water followed by 70% ethanol.

        - Areas under the grates that are inaccessible during daily cleaning must be disinfected and recorded on a quarterly basis.

      - Units must be professionally certified every 6 months to verify HEPA filter integrity and airflow rates.

  - Reagent Labeling & Signage Guide

    - Reference

      - All aliquoted reagents must be labeled according to the hospital-wide GHS (Globally Harmonized System) standards.

    - Core Labeling Requirements

      - Every aliquoted or reconstituted reagent bottle must display:

        - Content & Quantity: Full chemical name (no formulas) and volume.

        - Concentration/Titer: Specific strength of the reagent.

        - Storage Requirements: Temperature or safety cabinet instructions.

        - Key Dates: Preparation/reconstitution date, open date, and expiration date.

    - Microbiology Standard

      - If no manufacturer expiration is provided, the expiration will be one year from the date of preparation.

    - Signage Reference

      - When preparing secondary labels, use the hospital-approved pictograms to indicate primary hazards.

  - Bioterrorism Readiness

    - Newport Hospital Microbiology is a Satellite Laboratory.  Specimens are generally processed no differently from usual practice, that being plated cultures sent to RIH Microbiology for ID and sensitivity.

    - If a bioterrorism event is suspected notify the following:

      - Newport Hospital Infection Control Department (845-1557).

      - RIH Microbiology at (444-5273).

      - RIDH either by phone (222-5586/5585) or fax (222-4572).

    - Statewide Drills organized by RIDH: Newport Hospital participates as a Satellite Laboratory in the statewide bioterrorism drills organized by RIDH. The testing can include gram staining.  Results are reported to RIDH who then compiles the data from all the Sentinel and Satellite laboratories in the state. Full participation in these exercises is to ensure Newport Hospital’s proficiency in these areas. This type of testing would only be performed at Newport Hospital during an event causing an overwhelming workload at RIH Microbiology.

- References

  - Laboratory Safety Preparedness Binder (NPH)

  - GHS (Globally Harmonized System) standards

- Revisions

  - 02/12/2020- Removal of Appendix AP25 and edit of pathway to Safety Manual on Intranet

  - 02/17/2026 – Removed topics no longer performed in the NPH Microbiology Lab. Added protocol regarding BSC cabinets and Bioterrorism readiness. Reformatted for consistency between affiliated laboratories.