Laboratories at Washington Adventist Hospital, Shady Grove Medical Center and Shady Grove Adventist Emergency Center at Germantown

Date: 4/22/2015

LABORATORY ALERT

Subject: Coag Analyzers added to Instrument Manager

(DI-Data Innovations) and BioRad Unity Real Time

The Laboratory is informing you of the following change:

| Effective date: | April 28, 2015 |
| --- | --- |
| Purpose: | The Stago instruments will be migrated to DI and BioRad |
| Auto Logoff: | Users will automatically be logged off from Instrument Manager after a defined period of inactivity. The timer is set to 60 minutes. |
| Coag QC: | QC will flow to BioRad |
| Method Codes: | Site; Instrument; Old Instrument Method Code; New Instrument Method Code; Auto File Tech Code. GEC; Stago Compact-1; ST1G; GS1; 894. SGMC; Stago Compact-1; ST1S; SS1; 893. Stago Compact-2; ST2S; SS2. WAH; Stago Compact-1; STAW1; WS1; 895. Stago Compact-2; STAW2; WS2. |
| Work List: | Site; Coagulation; Combined; Quality Control. GEC; GEC COAG HOLD; GEC CHEM & COAG HOLD; GEC QC LIST. SGMC; SGAH COAG HOLD; SGAH CHEM & COAG HOLD; SGAH QC LIST. WAH; WAH COAG HOLD; WAH CHEM & COAG HOLD; WAH QC LIST. |
| Downtime: | Short term downtime will be handled via manual data entry. |
| Contacts: | Rob SanLuis |
| LBarrett/4.22.15 |  |

G:\AHC_Lab\Clinical Folders\Communication_General Lab\Lab Alerts

- TRAINING UPDATE

| Lab Location: | GEC, SGMC & WAH | Date Distributed: | 4/22/2015 |
| --- | --- | --- | --- |
| Department: | Core (Coag) | Due Date: | 5/6/2015 |
|  |  | Implementation: | 4/28/2015 |

- DESCRIPTION OF PROCEDURE REVISION

Name of procedure:

Activated Partial Thromboplastin Time (APTT); SGAH.G02.4; WAH.G02.4; GEC.G02.4. Prothrombin Time (PT) and INR; SGAH.G03.3; WAH.G03.3; GEC.G03.3. D-Dimer; SGAH.G04.3; WAH.G04.3; GEC.G04.3. Fibrinogen; SGAH.G05.4; WAH.G05.4; N/A. Thrombin Time; N/A; WAH.G06.3; N/A. STA Compact Operating Instructions; SGAH.G07.3; WAH.G08.3; GEC.G05.3.

Description of change(s):

All of the above analyte SOPs Section; Reason. 6.4, 6.6; Replace LIS with Unity Real Time. Revisions to D-Dimer only also included Section; Reason. 10.5; Correct <Mmin reporting value. A copy of D-DIMER SOP showing the revisions is attached Change to STA Compact Operating SOP Section; Reason. 5.4; replace LIS with Unity Real Time, remove LIS QC codes. 6; add Bio-Rad Unity Real Time SOP. These revised SOPs will be implemented on April 28, 2015

Document your compliance with this training update by taking the quiz in the MTS system.

Approved draft for training all sites (version 3)

Technical SOP

| Title | D-Dimer |  |
| --- | --- | --- |
| Prepared by | Ashkan Chini | Date: 4/7/2011 |
| Owner | Robert SanLuis | Date: 6/3/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	6

7.	Equipment And Supplies	8

8.	Procedure	9

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	12

13.	Procedure Notes	11

14.	Limitations Of Method	12

15.	Safety	13

16.	Related Documents	14

17.	References	14

18.	Revision History	14

19.	Addenda	15

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| D - Dimer Quantitative | Immunoturbidometric / STA® Compact | DDIMER |

Synonyms/Abbreviations

D - Dimer

Department

Coagulation

# ANALYTICAL PRINCIPLE

This assay is based on the change in turbidity of a microparticle suspension that is measured by photometry. A suspension of latex microparticles, coated by covalent bonding with monoclonal antibodies specific for D-dimer, is mixed with the test plasma whose D-dimer level is to be assayed.  An antigen-antibody reaction takes place, leading to an agglutination of the latex microparticles which induces an increase in turbidity of the reaction medium.  This increase in turbidity is reflected by an increase in absorbance, the latter being measured photometrically. The increase in absorbance is a function of the D-dimer level present in the test sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting plasma may be used for samples to be analyzed by this method. Vacutainer tube must be filled to the line to ensure the proper ratio of blood to anticoagulant. |
| Special Collection Procedures | If hematocrit >55%, refer to appendices A and B for collection instructions. |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Whole Blood (sodium citrate) |  |
| -Other Acceptable | None |  |
| Collection Container | Light blue top tube (3.2% sodium citrate) Citrated blood 9:1 (blood to anticoagulant) |  |
| Volume - Optimum - Minimum - Optimum | 2.7 mL (9:1 blood to anticoagulant) in a 2.7 ml tube 2.4 mL (9:1 blood to anticoagulant) in a 2.7 ml tube 1.8 mL (9:1 blood to anticoagulant) in a 1.8 mL tube |  |
| - Minimum | 1.8 mL (9:1 blood to anticoagulant) in a 1.8 mL tube |  |
| Transport Container and Temperature | Light blue vacutainer (as above) or a clean plastic screw capped vial at room temperature. |  |
| Stability & Storage Requirements | Room Temperature: (20 ± 5 C) | 8 hours |
|  | Refrigerated: | Not recommended |
|  | Frozen plasma: | 1 month at -20 C. |
| Specimen preparation | Centrifuge whole blood for specified time /speed documented on each centrifuge for preparing platelet-poor plasma. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Clotted or under-filled tubes are not accepted. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Moderate to gross hemolysis. Reject sample and request a recollection. Credit the test with appropriate LIS English text code HMM (Specimen moderately hemolyzed) or HMT (Specimen markedly hemolyzed) Lipemia: Acceptable Icterus: Acceptable |  |
| Other Considerations | None |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| STA - LIATEST® D – DI: Buffer & Latex | Diagnostic Stago (REF 00515) |
| STA – Owren-Koller Buffer | Diagnostic Stago (REF 00360) |
| Pure Reagent Grade water | NERL Diagnostics (Cat. No. 0015) |

  - Reagent Preparations and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent 1 & 2 | STA - LIATEST® D – DI : Buffer and Latex |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | 2-8C |
| Stability | Unopened reagents are stable until expiration date indicated on the box label. With the STA-mini Reducer and perforated cap in place the stability of Reagents 1 and 2 after opening and in their original vials is 15 days on the Stago. |
| Preparation | Allow the reagents 1 and 2 to stand at room temperature (18-25C) for 15 minutes before use. Mix the reagents by gentle swirling of the vials without creating any bubbles. Then place the perforated cap on each vial. |

| Reagent 3 | STA – Owren-Koller Buffer |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | 2-8C |
| Stability | The buffer solution in intact bottles is stable until the expiration date indicated on the box label. After opening it remains stable for 3 days. |
| Preparation | Allow it to stand at room temperature (18-25C) for 30 minutes before use. |

| Reagent 4 | NERL Reagent Grade water |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | Room temperature |
| Stability | Stable 30 days after opening. |
| Preparation | Ready to use |

# CALIBRATORS/STANDARDS

1.   The kit reagents are pre-calibrated: this calibration is identical for all the reagents of each lot.

2.   Entering the data for the calibration curve:

 	The database of the STA® Compact monitors all reagent lot numbers.  When the operator scans a new lot of Liatest D-Dimer reagent, the STA® Compact will request the operator to scan the bar code printed on the bar code insert across the STA® Compact bar codes reader.

 	The calibration curve will be validated for the lot being used once the two Liatest D-Dimer controls have been run.  If the validation controls are outside the assayed range, The STA® Compact will not run patient samples.

3.   Examine calibration curve on screen:

 	Through the MAIN MENU under CALIB/CONTROL select CALIBRATION.

- Move the cursor to D-Dimer and press Enter .  Curve will appear on STA® Compact screen.

- Print calibration curve:

 	While examining the curve on the STA® Compact screen, press ESC key for options.

  	Select print Option Enter .  Select Execute Enter .

 	The curve will print along with the information on all reagents and control lot numbers.  Also included are control results and ranges.

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalogue Number |
| --- | --- |
| STA® Coag control N + P | Diagnostic Stago (REF 00526) |

6.2 	Control Preparations and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name, (2) lot number, (3) date of preparation, (4) time prepared, (5) expiration date and time, (6) initials of tech, and (7) any special storage instructions; check for visible signs of degradation.

| Control | Coag Controls N + P |
| --- | --- |
| Preparation | Reconstitute each vial of Reagent 1 or 2 with exactly 1 mL of Reagent Grade water. Allow the reconstituted material to stand at room temperature for 30 minutes. Then, swirl the vial gently before use. |
| Storage/Stability | 2-8° C The reagents in intact vials are stable until the expiration date indicated on the box label, when stored at 2-8° C. Once reconstituted, Reagents 1 and 2 remain stable for 8 hours. |

- After the reconstitution period, request the product drawer to open through the MAIN MENU under LOADING and bar code the controls.  Place the controls into the appropriate drawer.

- QC can be run automatically at pre-set intervals (in Test Set-up) or by ordering manually from the Quality Control Menu.

- All control ranges are monitored automatically by the STA® Compact.  If any controls are outside the  2 SD range, the instrument will audibly and visually alarm the operator.  Otherwise, the results can be found in the individual QC files.  Control results are automatically filed in the STA® Compact QC file.  All results for a 24-hour period are converted to a “mean” value at midnight.  This mean is used in the statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- To print all the QC data points for the D-Dimer test, perform the following procedure prior to midnight.  From the MAIN MENU under CAL. /CONTROL select QUALITY CONTROL press Enter .  Cursor to the D-Dimer test and press Enter  to view the Levy-Jennings chart.  Press F1 to view the results in tabular form. Press F6, select Execute then press Enter  to print the individual values under current controls.  Press ESC key to exit (back to graph).  Press F2 or F3 to view other levels and continue with F1 to view the result list.

  - Frequency

Both controls are run at the beginning of each shift, every 4 hours after, and with the change of any reagent used in test performance.

Both controls are run after any maintenance is performed on the analyzer.

  - Tolerance Limits

| Step | Action |
| --- | --- |
| 1 | The established QC ranges are in the QC file of the STA Compact. The quality control results from the instrument are transmitted to the LIS Unity Real Time and can be viewed in that program OEM function. Any out-of-range QC results will be flagged by the LIS. |
| 2 | If all controls are within QC parameters all sample results can be reported. |
| 3 | Rejected runs must be effectively addressed by corrective action. Steps taken in response to QC failures must be documented. Patient samples in failed analytical runs must be reanalyzed. Supervisor may override rejection of partial or complete runs only with detailed documentation that follows criteria that is approved by the Medical Director. |
| 4 | Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, and the date and initials of the person recording the information. See the QC/QA SOP “QC Responsibilities” for more detail. |
| 5 | If the assay is down and results will not be reported in the scheduled turnaround time, clients will be notified of the situation. |

6.5	Review Patient Data

Technologist must review each result print-out for error messages. Refer to the STA® Compact system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

  - Documentation

- QC tolerance limits are programmed into the instrument and the LIS Unity Real Time; it calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

  - Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

STA® Compact – Analyzer

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge calibrated for preparing platelet-poor plasma

7.3	Supplies

- Cuvette Roll – Diagnostic Stago

- STA – brass adaptors

- Plastic micro cups

- STA Mini-Reducer

- Plastic transfer pipettes

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection is required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.  A current Package Insert is included as a Related Document.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1 | At the start of each shift, verify instrument temperatures and availability of cuvettes and cleaner solution by accessing the System Status screen from the main bar. |
| 2 | Record the temperatures on the maintenance sheet. If the reagent arm 2, measuring block, or reagent drawer temperatures are out of range, corrective action must be taken prior to patients being run. |
| 3 | Make sure that there is an adequate supply of reagents in the analyzer, and they are in date. |
| 4 | Load cuvettes and cleaner/wash solution on the analyzer if needed. |

| 8.2 | Analytical Procedure |
| --- | --- |
| 1 | Refer to START-UP procedure for STA® Compact before running patient specimens on the STA® Compact at the start of each shift. |
| 2 | Request quality control. Through MAIN MENU under CALIB. /CONTROL select QUALITY CONTROL and press Enter  . Cursor to the D-Dimer test. Select D-Dimer by pressing F1 and then F10. Type in your Access Code to run the QC. |
| 3 | Load patients’ samples: Access the sample drawer(s) through the MAIN MENU, under LOADING, Select Sample, press Enter  . After the drawer opens, identify the type of specimen, such as micro sample (press F8), or stat (press F12). Identify the sample by bar coding or manually entering on the keyboard the patient identification number and then placing the specimen into the drawer. |
| 4 | In MANUAL MODE, the operator must order the test(s) from the Selection menu or from the Recorded Profile/s Cursor to the test and press Enter  to select. When all tests are ordered, press F10 to save. |
| 5 | In AUTO MODE, the STA®/STA® Compact will automatically order the test(s) selected in the AUTO MODE profile. |
| 6 | If TELELOADING is selected as the AUTO MODE profile, the STA®/STA® Compact will query the host computer and download the test(s) as well as assign the status (i.e. stat). |
| 7 | As soon as the sample drawer closes, the TEST STATUS screen will appear. If there is not enough reagent(s) to run the test(s), the suspect reagent(s) will appear in red with the amount of depletion. This depletion of reagent will BLOCK the SAMPLE PIPETTING. When this occurs, add the necessary reagent(s) to run the samples by responding N (NO) to the warning message ‘NEW TESTS ARE DELAYED - REACTIVATE?’ Reagents can then be loaded in the drawer. By responding Y (YES) to the warning message ‘NEW TESTS ARE DELAYED - REACTIVATE?’, the instrument will continue to perform all tests for which there is sufficient reagent (i.e. while waiting for reagents to stabilize after reconstitution) |
| 8 | If the patients’ results fall outside the assay reportable range 4.0 μg/mL the instrument will automatically do a 1:5 dilution on the samples in question. This auto dilution will let the instrument report results up to 20.0 μg/mL. |
| 9 | All patient results are displayed on the TEST PANEL screen and automatically print out and transmit if selected on the system status menu. |
| 10 | For results in question that need operator intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will display the FILE PROCESSING screen. Follow the options on the left-hand side of the screen (i.e. F3 - rerun test). |

# CALCULATIONS

1.  	The STA® Compact automatically plots the results in delta OD off of a standard curve and converts the results to μg/ml FEU.

2.  	The assay uses the sample undiluted.  If the result is greater than the reportable range, 4.0, a dependent test with a 1:5 dilution will be ordered to take the reportable range to 20.0. The STA® Compact automatically corrects the result for the dilution change.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

N/A

10.2	Rounding

No rounding is necessary.  The instrument reports results up to two decimal points.

10.3	Units of Measure

μg/mL FEU

10.4	Clinically Reportable Range (CRR)

0.27 – 20.0 μg/mL FEU

  - Repeat Criteria and Resulting

The printout from the STA Compact is reviewed for repeat criteria and samples are repeated if needed. Results will be transmitted to the LIS and released using the OEM function.

| IF the result is … | THEN… |
| --- | --- |
| < Mmin | Repeat, check for clots. If result is still <Mmin, report as <0.27 μg/mL FEU-REP |
| > Mmax | Check for clots, repeat using the D-Di 1:5 test |
| If D-Di 1:5 is quantifiable | Report the result with comment REP |
| If D-Di 1:5 is > Mmax | Repeat. Report the result as > 20.0 µg/mL FEU-REP |
| For any of the above situations, be sure the specimen is not under-filled or over-filled, then check the Hematocrit (HCT) result. If the HCT is greater than 55%, refer to appendices A and B for special tube preparation. |  |

| Message | Code |
| --- | --- |
| Verified by repeat analysis | Append –REP to the result. |

Note: 	All patient results are reported with the following comment:

Less than ≤0.50 µg/mL FEU = Negative

Greater than >0.50 µg/mL FEU = Positive

Positive results are non-specific and are seen in a variety of conditions including DVT, pulmonary embolism, recent surgery, cancer, trauma and pregnancy.  Values greater than 0.50 µg/mL FEU may also be seen in otherwise healthy patients >70 years of age.

# EXPECTED VALUES

11.1	Reference Ranges

≤ 0.5 μg/mL FEU

  - Critical Values

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

D-Dimer quantitative assay detects the presence of Disseminated Intravascular Coagulation (DIC). In DIC the fibrinolytic system is activated and therefore the D-Dimer level increases.  D-Dimer assays can help in the diagnosis of DIC in these cases.  It is established that a normal D-Dimer level is an important factor to rule out the diagnosis of deep venous thromboses (DVT) or pulmonary embolisms (PE).  The decrease of D-Dimer levels during heparin therapy for a DVT allows the monitoring of evolution and prognosis of the thrombosis.  This decrease reflects the quality of the endogenic thrombolysis.  The D-Dimer level increases during the activation states of coagulation because they induce the production of thrombin which is followed by the formation of fibrin and leads to fibrinolysis, the latter being most frequently reactive.  The D-Dimer level thus increases following coagulation activation.

Increased levels of D-Dimer have been reported in post-operative period, cancers, cirrhosis, and hemorrhages.

# PROCEDURE NOTES

- FDA Status:  Approved/cleared

- Validated Test Modifications:  None

1.  The detection threshold of the STA® Liatest® D-Dimer on the STA® Compact is 0.27 μg/mL                                     FEU.  The printout limits are pre-set at 0.27 – 4.00 μg/mL FEU.  When a dependent test is set-up to extend the reportable range of the main test, the print out limit should be extended to 20.00 μg/mL FEU.

2.  The STA® Liatest® D-Dimer results are expressed in FEU, Fibrinogen Equivalent Units.  By definition, an FEU is the quantity of fibrinogen initially present that leads to the observed level of D-Dimer.  In general, the actual quantity of D-Dimer is approximately half of an FEU.

3.  A >Mmax  result on the primary assay dilution (1:1 dilution) indicates a result that is greater than 4.00 μg/mL FEU.  In this case the analyzer will automatically do a 1:5 dilution to obtain the result.

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0.27 – 4.00 μg/mL FEU

14.2	Precision

Different plasmas were used for the intra assay and inter assay reproducibility studies on the STA® Compact.

|  | Intra-Assay Reproducibility |  | Inter-Assay Reproducibility |  |
| --- | --- | --- | --- | --- |
| Sample | Sample 1 | Sample 2 | Sample 3 | Sample 4 |
| n mean (seconds) SD (seconds) | 21 0.29 0.04 | 21 2.71 0.08 | 10 0.32 0.05 | 10 2.78 0.14 |

14.3	Interfering Substances

1.  	Cloudy plasmas may lead to an under-estimation of the D-Dimer level.  Ensure that the absorbance value at 540 nm of the plasma diluted 1:6 with STA® - Owren-Koller buffer is < 0.35.

2.	Concentration of Fibrinogen Degradation Products greater than 15 μg/ml may lead to an over-estimation of the D-Dimer level.

3.	The presence of rheumatoid factor at a level greater than 50 IU/ml may lead to an over-estimation of the D-Dimer level.

4.	The STA® Liatest®  D-Dimer is insensitive to fibrinogen and the E fragment.  A low cross-reactivity is observed with the D fragment.

5.	The STA® Liatest® D-Dimer is insensitive to the following substances:

hemoglobin (up to 2 g/l); conjugated bilirubin (up to 290 mg/L); unconjugated bilirubin (up to 200 mg/L); unfractionated heparin (up to 2 IU/mL; LMWH (up to 2 anti-Xa IU/ml)

14.4	Clinical Sensitivity/Specificity/Predictive Values

N/A

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Laboratory Quality Control Program

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Hemolysis, Icteria and Lipemia Interference (Lab policy)

- Repeat Testing Requirements (Lab policy)

- STA Compact Operating Instructions, Coagulation procedure

- Verification of Platelet Poor Plasma, Coagulation procedure

- Current package insert for STA® LIATEST D-DIMER

# REFERENCES

- van der Graaf F, et. al., Exclusion of Deep Venous Thrombosis with D-Dimer Testing, Thromb Haemost. 2000;83:191-198

- Diagnostic Stago STA® LIATEST D-DIMER package insert:  Revised August 2012.

- STA®-Coag Control N + P (REF 00526): citrated control plasmas normal and abnormal levels; Control Plasmas for Assays of Coagulation Parameters on STA, Revised July 2011.

- STA® Compact Operators Manual.  STA® DSI-TSD-SM August 2004, STA® DSI-TSD-US April 2003, and V1.3 revised February 2003.

- Diagnostic Stago STA® Owren-Koller Buffer Solution for Coagulation Tests.  Revised: April 2012.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Supersedes G003.006 |  |  |
| 000 | 6/11/12 | 2.0 | Update to match package insert | J.Buss | J. Buss, RSL |
| 000 | 6/11/12 | 3.2 | Add frozen temperature | J.Buss | J. Buss, RSL |
| 000 | 6/11/12 | 4.1 | Remove Millipore water | J. Buss | J. Buss, RSL |
| 000 | 6/11/12 | 4.2 | D-DI reagent open stability edited | J.Buss | J. Buss, RSL |
| 000 | 6/11/12 | 6.3 | Add QC performed after maintenance | J.Buss | J. Buss, RSL |
| 000 | 6/11/12 | 15 | Update to standard wording | L. Barrett | J. Buss, RSL |
| 001 | 6/3/14 |  | Update owner | L Barrett | R SanLuis |
| 001 | 6/3/14 | 3.1 | Add reference to Appendices | A Chini | R SanLuis |
| 001 | 6/3/14 | 3.2 | Update tube volumes | A Chini | R SanLuis |
| 001 | 6/3/14 | 4.2 | Change storage temp and prep for buffer | A Chini | R SanLuis |
| 001 | 6/3/14 | 6.2 | Add step to print QC | A Chini | R SanLuis |
| 001 | 6/3/14 | 10.4 | Change CRR lower value | A Chini | R SanLuis |
| 001 | 6/3/14 | 10.5 | Add instruction for Hct >55 | A Chini | R SanLuis |
| 001 | 6/3/14 | 13, 14.1 | Change lower value of analytical range | A Chini | R SanLuis |
| 001 | 6/3/14 | 14.3 | Update to match package insert | A Chini | R SanLuis |
| 001 | 6/3/14 | 16 | Update titles | L Barrett | R SanLuis |
| 001 | 6/3/14 | 19 | Add Appendix A and B | A Chini | R SanLuis |
| 001 | 6/3/14 | Footer | Version # leading zero’s dropped due to new EDCS in use as of 10/7/13. | L Barrett | R SanLuis |
| 2 | 4/7/15 | 6.4, 6.6 | Replace LIS with Unity Real Time | L Barrett | R SanLuis |
| 2 | 4/7/15 | 10.5 | Correct <Mmin reporting value | L Barrett | R SanLuis |

# ADDENDA

A.	Instructions for Preparing Collection Tube for Hematocrit >55%

B.	Phlebotomist Instructions for Blood Collection

Appendix A

Instructions for Preparing Collection Tube for Hematocrit >55%

Explanation:

Polycythemia is a disease state in which the proportion of blood volume that is occupied by red blood cells increases - basically when Hematocrit (HCT) is greater than 55%.  It can cause prolonged coagulation results.

When a prolonged coagulation result is obtained, check the specimen for a clot first.

If the specimen is not clotted, be sure the specimen is not under-filled or over-filled, then check the HCT result.

If a HCT result of greater than 55% is obtained, immediately notify the doctor or attending nurse and ask for a redraw using a special tube prepared by the lab.

To prepare a special tube in the lab use the following instructions and formula:

The anticoagulant volume in the collection tube must be adjusted to obtain a 9:1 ratio of blood to Sodium Citrate.  Under or over-filling of the specially prepared collection tube is not acceptable.

The vacuum in the collection tube will be broken to adjust the volume of collection anticoagulant.  Because of this special collection technique, the stability for these whole blood specimens is reduced to four (4) hours after collection.

Formula to calculate the anticoagulant volume is:

Anticoagulant in mL = [(100 – HCT) / (595 – HCT)] x Volume of blood

Example 1:  Specimen with a 70% HCT in a 2.7 mL tube:

Patient with HCT of 70%

Using a 2.7 mL tube

Anticoagulant in mL = [(100 – 70) / (595 – 70)] x 2.7 = 0.15 mL or 150 uL

Pipette a 2.7 mL tube in a way to leave only 150 uL of anticoagulant in there.

A 2.7 mL tube contains 0.3mL anticoagulant; therefore remove 0.15mL

Example 2:  Specimen with a 70% HCT in a 1.8 mL tube:

Patient with HCT of 70%

Using a 1.8 mL tube

Anticoagulant in mL = [(100 – 70) / (595 – 70)] x 1.8 = 0.1 mL or 100 uL

Pipette a 1.8 mL tube in a way to leave only 100 uL of anticoagulant in there.

A 1.8 mL tube contains 0.2mL anticoagulant; therefore remove 0.1mL

Example 3:  Specimen with a 60% HCT in a 2.7 mL tube:

Patient with HCT of 60%

Using a 2.7 mL tube

Anticoagulant in mL = [(100 – 60) / (595 – 60)] x 2.7 = 0.2 mL or 200 uL

Pipette a 2.7 mL tube in a way to leave only 200 uL of anticoagulant in there.

A 2.7 mL tube contains 0.3mL anticoagulant; therefore remove 0.1mL

Example 4:  Specimen with a 60% HCT in a 1.8 mL tube:

Patient with HCT of 60%

Using a 1.8 mL tube

Anticoagulant in mL = [(100 – 60) / (595 – 60)] x 1.8 = 0.13 mL or 130 uL

Pipette a 1.8 mL tube in a way to leave only 130 uL of anticoagulant in there.

A 1.8 mL tube contains 0.2mL anticoagulant; therefore remove 0.07mL

Appendix B

Phlebotomist Instructions for Blood Collection

The technologist will prepare a special tube in which the anticoagulant has been adjusted, therefore the tube is not vacuumed.  The technologist will inform the phlebotomist of the exact amount of blood needed to fill the tube.

Equipment and Supplies

Latex gloves

Latex free tourniquet

Latex free Band Aid or Tape

Alcohol Prep (70% alcohol)

2x2 sterile gauze

Collection tube

Blood Collection Set 21 or 23 gauge winged set

Blood Transfer Device

3mL syringe

Biohazard bag

Biohazard sharps container

LIS collection list and label/Lab requisition

Collection Steps

- Introduce yourself to the patient by stating your first and last name.

- Positively identify the patient according to the SOP ‘Patient Identification’, Phlebotomy procedure manual.

- Wash hands. Apply gloves.

- Explain the procedure to the patient and obtain patient’s consent to draw blood.

- Collect equipment and correct technologist-provided collection tube.

- Assemble equipment and break needle and syringe seals in the presence of the patient.

- Apply tourniquet about midway between the elbow and the shoulder 3-4 inches above the venipuncture site).  Place patient’s arm in a downward position to prevent reflux of ‘backflow’ of blood from the tube into the venous system.  Ask the patient to close hand gently.

- Palpate/feel for vein locating a vein that will flow fast (reducing the possibility of the blood clotting).

- Clean the area for venipuncture with a 70% alcohol pad decontaminating the collection site.

- Allow the area to air-dry completely.

- Assemble the 21 or 23 gauge winged set to the 3mL syringe.  Pull back the plunger to dispel all the air out of the syringe.

- With the bevel up, align the needle with the vein while holding the skin taut.  Insert the needle at a 15-30 degree angle with the skin.  Remove your hand from drawing the skin taut.  Grasp the syringe and draw back bringing the plunger tip to the exact amount of blood requested by the technologist.

- Release the tourniquet. Ask the patient to open hand.

- Place gauze above the puncture site and remove the needle while simultaneously applying pressure on the puncture sit.  Firmly activate needle safety shield, a click must be heard to ensure that the safety shield is secure.

- Remove 21 or 23 gauge winged set from syringe.

- Attach the blood-filled syringe to the Blood Transfer Device.

- Connect the Blood Transfer Device to the un-vacuumed tube, provided by the technologist, and slow and gently fill the collection tube.  DO NOT FORCE blood into tube.  Pressure can lead to tube explosion and blood exposure.

- Place the cap on the tube and invert a few times to make sure the anticoagulant is mixed with blood.

- Dispose of all blood collection equipment into the nearest sharps container.  DO NOT disassemble the syringe from the Blood Transfer Device.

- Dispose of all other used materials in appropriate container and wash hands.

- Label the sample with the LIS collection label and write the time, date, and your tech code.

- Transport specimen to the Lab.