## TRAINING UPDATE

| Lab Location: | SGMC and WAH | Date Implemented: | 11.20.2017 |
| --- | --- | --- | --- |
| Department: | Processing | Due Date: | 12.11.2017 |

- DESCRIPTION OF PROCEDURE REVISION

Name of procedure:

Aliquoting Specimens

Description of change(s):

- Only one specimen may be aliquoted at a time. - Never mix specimens from more than one tube. - Critical steps added to the procedure: - Compare patient name on the aliquot label to the patient name on original label. They must match EXACTLY. - Compare patient MRN on the aliquot label to the patient MRN on the original label. They must match EXACTLY. - Document tech code on the aliquot label and apply to tube with name near the tube opening. - Add an appropriate amount of specimen to the tube.

Non-Technical SOP

| ###### Title | Aliquoting Specimens |  |
| --- | --- | --- |
| ###### Prepared by | Leslie Barrett | Date: 7/10/2009 |
| ###### Owner | Stephanie Codina | Date: 11/13/2017 |

| ###### Laboratory Approval |  |  |
| --- | --- | --- |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Local Issue Date: | Local Effective Date: |  |

| Review: |  |  |
| --- | --- | --- |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	PURPOSE	2

2.	SCOPE	2

3.	RESPONSIBILITY	2

4.	DEFINITIONS	2

5.	PROCEDURE	2

6.	RELATED DOCUMENTS	3

7.	REFERENCES	3

8.	REVISION HISTORY	3

9.	ADDENDA AND APPENDICES	3

# PURPOSE

To provide a process for separating and labeling specimens for analysis.

# SCOPE

This procedure applies to all specimens that are shared between internal laboratory sections, instruments and/or reference laboratories.

# RESPONSIBILITY

All specimen processing and technologists must process specimens according to the procedure.

# DEFINITIONS

None

# PROCEDURE

| Step | Action |
| --- | --- |
| 1 | Never handle or aliquot a specimen without appropriate personal protective equipment to include labcoat, gloves, and safety shield/mask. |
| 2 | Check the specimen requirements using Sunquest function MIQ. - In-house tests may be aliquoted to decrease turn-around-time when a single specimen has ordered tests that will be run on different instruments. - Referral tests are aliquoted into plastic tubes to prevent specimen breakage and leaks during transport. |
| 3 | Aliquot ONLY one specimen at a time. - Physically separate the specimen being aliquoted from other specimens. - Do not mix specimens from different tube even when they were collected from the same patient. - A single specimen may be aliquoted into more than one tube. |

| Step | Action |
| --- | --- |
| 4 | Print the required number of labels for the aliquot(s) per procedure. |
| Start Critical Step |  |
| 5 | Compare the patient name on EACH printed label to the patient name on the original specimen. Names must match exactly. |
| 6 | Compare the patient medical record number on EACH printed label to the medical record number on the specimen. Medical record numbers must match exactly. |
| 7 | Write your tech code on EACH lab label to indicate you verified the name and medical record number on the lab label to the name and medical record number on the tube. |
| 8 | Apply one label to each aliquot tube. The label should be applied directly onto the tube. |
| 9 | Transfer the appropriate amount of specimen into the aliquot tube and cap the tube. - Do not combine specimens from more than one tube into a single aliquot tube even when they were collected from the same patient. - Never pour from an aliquot tube back to the primary tube. |
| 10 | Compare the name and medical record number on each aliquot tube to the name and medical record number on the primary tube to verify correct labeling. |
| End Critical Step |  |
| 11 | Deliver the aliquot to the appropriate testing department or store send out tests at the appropriate temperature based on handling requirements. |

# RELATED DOCUMENTS

None

# REFERENCES

N/A

# REVISION HISTORY

| Version | Date | Reason for Revision | Revised By | Approved By |
| --- | --- | --- | --- | --- |
|  |  | Supersedes SOP S007.001 |  |  |
| 000 | 2/6/12 | Section 5: add no combining or pour back | LBarrett | SKhandagale |
| 001 | 10/13/17 | Header: added WAH Footer: version # leading zero’s dropped due to new EDCS in use as of 10/7/13 Added critical steps and initials to align with secondary labeling SOP; updated format and wording for clarity. | SCodina | NCacciabeve |

# ADDENDA AND APPENDICES

None