| ## POLICY STATEMENT | ImmunoCard STAT! CAMPY is an immunochromatographic test for the qualitative detection of specific Campylobacter in human stool specimens. ImmunoCard STAT! CAMPY detects C. jejuni and C. coli in human stool. Test results are used in conjunction with the patient’s clinical symptoms and other laboratory tests to aid in the diagnosis of diseases caused by Campylobacter. |
| --- | --- |
| ## PURPOSE | This procedure provides technical instruction for the performance of the ImmunoCard STAT! CAMPY test. |
| ## SCOPE | This procedure applies to technical personnel authorized to perform testing. This group includes, but is not limited to Medical Technologists as well as Leads and Supervisory personnel. |
| RESPONSIBILITY | All the above personnel are responsible for following this procedure without exception. In addition, testing personnel are also responsible for evaluating the results and taking proper remedial action. |
| RELATED DOCUMENTS | MICR 6915 F ImmunoCard STAT! CAMP QC chart |

PRINCIPLE

ImmunoCard STAT! CAMPY is a lateral flow-based immunoassay for the direct detection of Campylobacter antigen in stool.  ImmunoCard STAT! CAMPY assay uses monoclonal antibodies specific for an antigen common to C. jejuni and C. coli.  Campylobacter antigen in the sample binds to the monoclonal antibody-colloidal gold conjugate as the sample moves through the device.  The Campylobacter-capture monoclonal antibody bound to the assay membrane at the Test position of the device binds antigen-Campylobacter antibody-colloidal gold complex. The Control Line is a goat anti-mouse antibody bound at the Control position of the device.

SPECIMEN HANDLING

Stool specimens should be collected in Cary-Blair based media and stored at 2-8°C for up to 48 hours.

MATERIALS:

- ImmunoCard STAT! CAMPY Test Device:  Plastic cassette containing a test strip with an immobilized monoclonal antibody specific to C. jejuni and C. coli.  The Test Device is provided in a sealed foil pouch with desiccant

- ImmunoCard STAT! CAMPY Sample Diluent/Negative Control: A buffered protein solution containing 0.094% sodium azide and 0.03% gentamicin as preservatives

- ImmunoCard STAT! CAMPY Positive Control: Inactivated Campylobacter jejuni in a buffered protein solution containing 0.094% sodium azide and 0.03% gentamicin as preservatives

- Transfer pipettes

Storage Requirements: Do not use kit components beyond labeled expiration date.

The expiration date is indicated on the kit label. Store the kit at 2-8°C and return the kit promptly to the refrigerator after each use. Do not mix lot components.

QUALITY CONTROL

Internal controls are contained within the test strip and therefore are evaluated with each test. Daily quality control requires documentation of these procedural controls for each day samples are tested. A PINK-RED band appearing at the Control Line serves as an internal positive control and indicates the test has been performed correctly, that sample flowed properly and that test reagents were active at the time of use. A colorless background around the Control or Test Lines serves as a negative procedural control.  Control or Test Lines that are obscured by heavy background color may invalidate the test and may be an indication of reagent deterioration, use of an inappropriate sample or improper test performance.

Positive and Negative Controls are performed with each new lot or shipment of test kits received and weekly thereafter.  For each new lot or shipment of kits, run the old kit controls with the new kit to perform a parallel crosscheck before the putting the new kit in use.

PROCEDURE:

- Bring all Test Devices, reagents and samples to room temperature (20-25°C) before testing.

- Mix stool as thoroughly as possible prior to pipetting.

- Mix Sample Diluent/Negative Control thoroughly prior to use.

- Controls

- Positive: Add exactly 6 drops of the Positive Control to the sample port of the device marked for the Positive Control.

- Negative: With the transfer pipette provided in the kit add 175 µL of the Sample Diluent/Negative Control (second mark from tip of pipette) to the sample port of a device marked for the Negative Control.

- Specimen Preparation - Cary-Blair stool

- Using the dropper assembly provided with the Sample Diluent/Negative Control, add 350 µL of Sample Diluent to test tube.

- Using a transfer pipette provided with the kit, add 50 µL of thoroughly mixed stool (first mark from pipette tip).

- Vortex for a minimum of 15 seconds.

- Save the transfer pipette in the sample for later use.

- Remove the ImmunoCard STAT! CAMPY test device from its foil pouch and label with the patient identification.

- Using the transfer pipette provided in the kit, add 175 µL of the diluted specimen (second mark from tip of pipette) to the sample port of the device.

- Incubate the test at 20-25 C for 20 minutes.

- Read the results within 1 minute of the end of incubation.

INTERPRETATION OF RESULTS

Positive test: PINK-RED bands at the Control and Test Line positions.  The appearance of a Test Line, even if very weak, indicates the presence of Campylobacter antigen.  The intensity of the Test Line can be less than that of the Control Line.

Negative test: PINK-RED band at the Control Line position.  No other bands are present.

Invalid Test Results:

- No band at the designated position for the Control Line.  The test is invalid as the absence of a control band indicates the test procedure was performed improperly or that deterioration of reagents has occurred.

- PINK-RED band appearing at the Test Line position of the device after the defined incubation limit.  Falsely positive results may occur if tests are incubated too long.

- Band of any color other than PINK-RED.  Bands with colors other than PINK-RED may indicate reagent deterioration.

- If any result is difficult to interpret, the test should be repeated with the same sample to eliminate the potential for error.  Obtain a new sample and retest when the original sample repeatedly produces unreadable results.

REPORTING

Positive test:  Report test result as “POSITIVE”. All positive tests will reflex to a Campylobacter culture. Subculture all positive antigen specimens to a Campy BAP and incubate in the appropriate environment for 72 hours.

Negative test:  Report test results as “NEGATIVE”

LIMITATIONS OF THE PROCEDURE

- The test is qualitative and no quantitative interpretation should be made with respect to the intensity of the positive line when reporting the result.

- Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

WARNINGS

- The ImmunoCard STAT! CAMPY Positive Control Reagent contains inactivated C. jejuni.  Handle as a potentially biohazardous material.

- Patient specimens may contain infectious agents and should be handled and disposed of as potential biohazards.

- Dispose of all used test materials in an appropriate container.  Treat waste as a potential biohazard.

REFERENCES

Meridian Bioscience ImmunoCard STAT® CAMPY package insert 06/09