| ## POLICY STATEMENT | Streptococcus pneumoniae is the leading cause of community-acquired pneumonia. Pneumococcal pneumonia has a mortality rate as high as 30% depending on bacteremia, age, and underlying diseases. When not properly diagnosed and treated, S. pneumoniae infection can lead to bacteremia, meningitis, pericarditis, empyema, purpura fulminans, endocarditis and/or arthritis. The BINAX NOW ® Streptococcus pneumoniae TEST is an Immuno-chromatographic membrane assay used to detect pneumococcal soluble antigen in human urine and CSF. |
| --- | --- |
| ## PURPOSE | This procedure provides technical instruction for the performance of the Streptococcus pneumoniae antigen test. |
| ## SCOPE | This procedure applies to testing personnel authorized to perform testing. This group includes, but is not limited to Laboratory Technologists as well as leads and supervisory personnel. |
| ## RESPONSIBILITY | All the above personnel are responsible for following the Streptococcus pneumoniae antigen procedure without exception. In addition, testing personnel are also responsible for evaluating the results and taking proper remedial action. |
| ## RELATED DOCUMENTS | MICR 6955 F Strep pneumo antigen QC Chart |

PRINCIPLE

The BINAX NOW ® Streptococcus pneumoniae TEST is an Immuno-chromatographic membrane assay used to detect pneumococcal soluble antigen in human urine.  Rabbit anti-S. pneumoniae antibody, the Sample line, is absorbed onto nitrocellulose membrane.  Control antibody is absorbed onto the same membrane as a second stripe.  Both rabbit anti-S. pneumoniae and anti-species antibodies are conjugated to visualizing particles that are dried onto an inert fibrous support.  The resulting conjugate pad and the striped membrane are combined to construct the test strip.

### SPECIMEN HANDLING

Collect urine sample in a standard urine container.  Specimen should be sent to lab as soon as possible. Samples can be stored at room temperature if tested within 24 hours of collection.  Alternatively, specimens may be stored at 2° to 8C for up to 14 days or at -10C to -20C for longer periods of time.  Boric acid may be used as a preservative.

NOTE: The test has not been approved for testing on plasma, serum, and other body fluids.

REAGENTS

Binax NOW ® Streptococcus pneumoniae Test Kit:

- Test Devices: A membrane coated with rabbit antibody specific for S. pneumoniae antigen and with control antibody is combined with rabbit anti-S. pneumoniae antigen and anti-species conjugates in a hinged test device.

- Reagent A: Citrate/Phosphate buffer with sodium lauryl sulfate, Tween 20, and sodium azide.

- Sample swabs: Designed for use in the Binax NOW ® Streptococcus pneumoniae Test.  Do not use other swabs.

- Positive Control Swab: Heat inactivated S. pneumoniae dried onto the swab.

- Negative Control Swab: S. pneumoniae negative swab.

Store the kits at 2°C to 30C.  The test cards and reagents are stable until the expiration dates marked on their outer packaging and containers.  DO NOT use the kit beyond its labeled expiration date. Do not mix lot components.

## QUALITY CONTROL

The Binax NOW ® Streptococcus pneumoniae Test contains built-in procedural controls.  Daily quality control requires documentation of these procedural controls for each sample tested each day. The pink-to-purple line at the “Control” position can be considered an internal positive procedural control. The background color in the window should be light pink to white within 15 minutes and should not interfere with the reading of the test result.

External Positive and Negative Controls are performed with each new lot or shipment of test kits received and weekly thereafter.

### PROCEDURE

- Bring patient sample(s) to room temperature, then swirl gently to mix.

- Remove room temperature device from its pouch just before use and lay flat.

- Dip a Binax swab (Do not use other swabs) into the sample to be tested, completely covering the swab head. If the swab drips, touch the swab to the side of the collection container to remove excess liquid.

- There are two holes on the inner right panel of the device. Insert specimen swab or control swab into the BOTTOM hole (swab well).  Firmly push upwards so that the swab tip is fully visible in the Top hole.  DO NOT REMOVE SWAB.

- Hold Reagent A vial vertically, ½ to one inch above the device.

- Control Swabs:  Slowly add six (6) free falling drops of Reagent A to the BOTTOM hole.

- Specimen Swabs:  Slowly add three (3) free falling drops of Reagent A to the BOTTOM hole.

- Immediately peel off adhesive liner from the right edge of the test device.  Close and securely seal the device.

- Read result in window 15 minutes after closing the device.  Results read beyond 15 minutes may be inaccurate. However, some positive patients may produce a visible sample line in less than 15 minutes.

### REPORTING

- A positive sample will give two pink-to-purple colored lines.  This means that antigen was detected.  Specimens with low levels of antigen may give a faint patient line. Any visible line is positive. Streptococcus pneumoniae vaccine may cause false positive results in the 48 hours following vaccination. Therefore, it is recommended that the Test not be administered within 5 days of receiving the S. pneumoniae vaccine.

Enter results in Meditech as:

POSITIVE: positive for pneumococcal pneumoniae antigen.

- A negative sample will give a single pink-purple colored Control Line in the top half of the window, indicating a presumptive negative result.

Enter results in Meditech as:

NEGATIVE:  presumptive negative for pneumococcal pneumonia antigen, suggesting no current or recent pneumococcal infection.  Infection due to S. pneumoniae cannot be ruled out since the antigen present in the sample may be below the detection limit of the test.

- If no lines are seen, or if just the Sample line is seen, the assay is invalid. Invalid tests should be repeated.

REFERENCES

Binax NOW ® Streptococcus pneumoniae Test package insert 2012