| ## POLICY STATEMENT | Bacteria containing the enzyme catalase break down hydrogen peroxide into water and oxygen which results in the appearance of gas bubbles. |
| --- | --- |
| ## PURPOSE | This procedure provides technical instruction for the performance of the Catalase test. |
| ## SCOPE | This procedure applies to testing personnel authorized to perform testing. This group includes, but is not limited to Laboratory Technologists as well as leads and supervisory personnel. |
| ## RESPONSIBILITY | All the above personnel are responsible for following the Catalase procedure without exception. In addition, testing personnel are also responsible for evaluating the results and taking proper remedial action. |
| ## RELATED DOCUMENTS | MICR 6416 F Catalase QC Chart |

#### REAGENTS

- 3% Hydrogen Peroxide

- 30% Hydrogen Peroxide (extremely caustic to skin – if contact occurs, wash immediately with 70% ethyl alcohol, not water)

Store at 2ºC to 8ºC. Protect from light. DO NOT use the beyond the expiration date.

#### QUALITY CONTROL

Perform with each new lot or shipment.

- Positive Control = Staphylococcus aureus ATCC 25923

- Negative Control = Streptococcus pyogenes ATCC 19615

Do not report patient results if expected results are not obtained from QC organisms.

#### PROCEDURE

- Touch the center of an 18-24 hour, well isolated colony from CHOC or BAP. If colony is from BAP do not pick up blood agar with the colony.

- Using an inoculating needle or applicator stick, pick a well-isolated colony and transfer to clean glass slide. Be sure colony is visible to the naked eye on the slide. If colony is from a blood agar plate, use care not to pick up blood.

- Add 1 drop of the Catalase Reagent to the slide – do not mix

- Aerobes = use 3% hydrogen peroxide

- Neisseria = use 30% hydrogen peroxide

- Anaerobes = Expose colony to air for 30 minutes before testing and use 15% hydrogen peroxide (dilute 30% 1:1with deionized water)

- Examine immediately for the rapid production of gas bubbles.

# REPORTING

- Positive = immediate appearance of bubbles

- Weak reaction = one to two bubbles.

- Negative = no bubbles or a few bubbles after 20 seconds

Enter the result in the appropriate spot in the Meditech workcard.

#### REFERENCES

Garcia, L. Clinical Microbiology Procedure Handbook. 3nd Edition, Volume 1, Catalase test, section 3.17.10, American Society for Microbiology, Washington, D.C