| ## POLICY STATEMENT | The oxidase test is a qualitative procedure for determining the presence or absence of cytochrome c oxidase activity in bacteria. The cytochrome oxidase produced by the organism oxidizes the cytochrome c, which in turn oxidizes the p-phenylenediamine reagent to form a purple colored compound. |
| --- | --- |
| ## PURPOSE | This procedure provides technical instruction for the performance of the DrySlide Oxidase. |
| ## SCOPE | This procedure applies to testing personnel authorized to perform testing. This group includes, but is not limited to Laboratory Technologists as well as leads and supervisory personnel. |
| ## RESPONSIBILITY | All the above personnel are responsible for following the DrySlide Oxidase procedure without exception. In addition, testing personnel are also responsible for evaluating the results and taking proper remedial action. |
| ## RELATED DOCUMENTS | MICR 6450 F Oxidase QC chart |

# REAGENTS

## BBL DrySlide Oxidase

- N,N,N’N’- tetramethyl-p-phenylenediamine dihydrochloride

- Ascorbic acid

Store at 15° to 30°C. Protect from light. DO NOT use the beyond the expiration date.  Each unsealed slide is good for one day’s use. Opened, resealed pouches are good for one week after opening.

# QUALITY CONTROL

Perform with each new lot or shipment.

- Positive Control = Escherichia coli ATCC 25922

- Negative Control = Pseudomonas aeruginosa ATCC 27853

# PROCEDURE

- Open the BBL Dry Slide oxidase pouch and remove the required number of slides.

- Using an appropriate inoculation device, pick isolated colonies or a sweep of confluent growth from a pure culture growing on nonselective and nondifferential media.

- Smear the organism directly onto the reaction area of the DrySlide Oxidase. Each reaction area can accommodate up to 4 tests. To ensure a proper reaction, spread the inoculum on the DrySlide reaction area 3 to 4 mm size.

- Examine the reaction area for the appearance of a dark purple color within 20 secs. Disregard color development after 20 seconds. (Mucoid Pseudomonas aeruginosa may require longer than 10 to 20 seconds to become positive.)

# REPORTING

- Positive = purple or dark color

- Negative = no color change or a change to light gray

Enter the result in the appropriate spot in the Meditech workcard.

### REFERENCES

BD BBL DrySlide Oxidase package insert, 2010/06