| ## Policy Statement | CAP requires that when more than one instrument/method is used to test for a given analyte, the instruments/methods are checked against each other at least twice a year for comparability of results. |
| --- | --- |
| ## Purpose | To ensure that results are reproducible between different instruments/methods within acceptable limits. |
| ## Scope | This policy applies to all instruments/methods in the Clinical and Pulmonary Laboratories. |
| Responsibility | All associates are responsible for following this procedure. The Lead Technologist or designee is responsible for completion of the study. Completed studies must be approved by the Laboratory Medical Director. |

General Requirements for all Clinical Laboratory Areas

Each month analyze at least two random or pooled patient samples for all reportable parameters/analytes performed by more than one instrument /methodology. Quantitative results must be less than or equal to acceptable limits for that analyte.  Semi-quantitative results must be within one gradient or dilution. Raw data is entered on an excel spreadsheet and retained per regulatory guidelines. The Medical Director of the section will review graphed data twice a year. Graph content is based on the type of testing and Medical Director Specifications.

Point of Care Testing

|  | Waived | Moderate |
| --- | --- | --- |
| Frequency | Once per year | Every six months |
| Meter to Meter | Glucose, Hemoglobin, BNP | ACT, PT/INR, Cardiac |
| Meter to Lab (5 specimens) | Glucose, Hemoglobin, Lipids, Hb A1C | PT/INR, Cardiac (Myo & CKMB only) |

Corrective Action Plan

The following steps will be taken if a result falls outside of the acceptable limits:

- Review quality control results from the previous 30 days.

- Review calibration results for specified lot number.

- Review maintenance records.

- Correct any issues discovered upon investigation

- Repeat testing with a new patient sample.

- If testing is still outside of the acceptable limits, call for technical service.

Acceptability Limits

| ANALYTE | LIMITS** | SOURCE |
| --- | --- | --- |
| CHEMISTRY |  |  |
| A1C | 8% | CAP |
| Acetaminophen | 10% | CAP |
| Albumin | 10% | CAP |
| Alkaline Phosphatase | 15% | CAP |
| ALT | 15% | CAP |
| Amylase | 15% | CAP |
| Amylase, Urine | 3SD | CAP |
| AST | 15% | CAP |
| Bilirubin, Direct | 15% | CAP |
| Bilirubin, Total | 15% | CAP |
| BUN | 15% | CAP |
| BUN, Urine | 15% | CAP |
| Calcium | 1.0 | CAP |
| Calcium, Urine | 15% | CAP |
| Cholesterol | 10% | CAP |
| CK | 15% | CAP |
| CKMB | 3SD | CAP |
| Chloride | 5% | CAP |
| Chloride, Urine | 15% | CAP |
| CO2 | 3SD | CAP |
| Creatinine | 15% | CAP |
| Creatinine, Urine | 15% | CAP |
| CRP | 3SD | CAP |
| Digoxin | 15% | CAP |
| ETOH | 15% | CLIA |
| Glucose | 10% | CAP |
| Glucose, Urine | 15% | CAP |
| Glucose, CSF | 10% | CAP |
| HCG | 3SD | CAP |
| HDL | 15% | CAP |
| Iron | 15% | CAP |
| Potassium | 0.5 | CAP |
| Potassium, Urine | 15% | CAP |
| Lactate | 3SD | CAP |
| LDL | 15% | CAP |
| Lipase | 15% | CAP |
| Microalbumin | 3SD | CAP |
| Magnesium | 15% | CAP |
| Magnesium, Urine | 15% | CAP |
| Myoglobin | 15% | CAP |
| Sodium | 4.0 | CAP |
| Sodium, Urine | 15% | CAP |
| Ammonia | 2.5SD | WLSH |
| Phenytoin | 15% | CAP |
| Phosphorus | 10.7% | CAP |
| Phosphorus, Urine | 15% | CAP |
| PTH | 3SD | CAP |
| Salicylate | 10% | CAP |
| Troponin T | 15% | CAP |
| Total Protein | 10% | CAP |
| Total Protein, Urine | 15% | CAP |
| Total Protein, CSF | 15% | CAP |
| Triglyceride | 9% | CAP |
| TSH | 3SD | CAP |
| UIBC | 15% | CAP |
| Valproic Acid | 15% | CAP |
| Vancomycin | 10% | CAP |
| COAGULATION |  |  |
| APTT | 15% | CAP |
| D-dimer | 3SD | CAP |
| Fibrinogen | 20% | CAP |
| INR | 20% | CAP |
| PT | 15% | CAP |
| HEMATOLOGY |  |  |
| WBC | 5% | Manufacturer’s Criteria |
| RBC | 2.5% | Manufacturer’s Criteria |
| HGB | 2.0% | Manufacturer’s Criteria |
| HCT | 2.5% | Manufacturer’s Criteria |
| MCV | 2.5% | Manufacturer’s Criteria |
| Platelet | 7% | Manufacturer’s Criteria |
| RDW | 3SD | CAP |
| Basophils | 1.0% | Manufacturer’s Criteria |
| Eosinophils | 2.0% | Manufacturer’s Criteria |
| Lymphocytes | 4.0% | Manufacturer’s Criteria |
| Monocytes | 3.0% | Manufacturer’s Criteria |
| Neutrophils/Granulocytes | 5.0% | Manufacturer’s Criteria |
| NRBC | 2.0% | Manufacturer’s Criteria |
| Reticulocyte | 20.0% | Manufacturer’s Criteria |
| URINALYSIS |  |  |
| Specific Gravity | 1 gradient | SAH |
| pH | 1 gradient | SAH |
| Protein | 1 gradient | SAH |
| Glucose | 1 gradient | SAH |
| Ketone | 1 gradient | SAH |
| Bilirubin | 1 gradient | SAH |
| Blood | 1 gradient | SAH |
| Nitrite | Exact | SAH |
| Urobilinogen | 1 gradient | SAH |
| Leukocyte Esterase | 1 gradient | SAH |
| BLOOD GAS |  |  |
| pCO2 | 5.0 | CLIA |
| pO2 (Blood Gas) | 3SD | CLIA |
| pH | .04 | CLIA |
| Calcium Total | 1.0 | CLIA |
| Calcium, Ionized | 3SD | CAP |
| FCOHb | 3.0 | R&D |
| FMetHb | 3.0 | R&D |
| FO2Hb | 3.0 | R&D |
| POCT |  |  |
| Glucose | 20% | CLIA |
| Hemoglobin | 7% | CAP |
| BNP | 7% | CAP |
| Cholesterol | 10% | CAP |
| HDL | 30% | CAP |
| Triglyceride | 2.0 or 9% | CAP |
| LDL | 30% | CAP |
| A1C | 8% | CAP |
| ACT (meter to meter only) | 20% | Manufacturer’s Criteria |
| INR | 20% | CAP |
| Myoglobin, CKMB and Trop I | Obtain the same interpretive outcome when compared to each method respective reference ranges. Interpretive Outcomes: Neg - ≤ Upper end of reference range Pos - > Upper end of reference range |  |

**Chemistry Limits: if CAP limit >15% SAH limit adjusted to current value of 15%

References

- College of American Pathologists Participant Summary 2010

- CLIA '88 Proficiency Testing Limits, U.S. Federal Register

- Wisconsin State Laboratory of Hygiene Proficiency Testing, http://www.slh.wisc.edu/pt

- ACT (ITC) Manufacturer’s stated guidelines

- Clinical and Laboratory Standards Institute.  Verification of Comparability of Patient Results Within One Health Care System; Approved Guideline.  CLSI document C54-P.  Wayne, PA; 2007

- Radiometer Research and Development