| ## Policy Statement | Core Laboratory Personnel are responsible for insuring the specimen submitted for testing is acceptable and the procedure for performing this assay is not violated. |
| --- | --- |
| ## Purpose | This procedure provides technical instruction for acceptable performance of the “Body Fluid Analysis” on the Sysmex XE-Series” analyzer. |
| ## Scope | This procedure applies to testing personnel authorized to perform testing using the Sysmex XE-5000 analyzer. |
| Responsibility | All authorized personnel are responsible for following procedural guidelines and insuring good laboratory practice is followed. |
| Related Documents | Sysmex Operators XE-5000 Main Manual Sysmex Operators XE-5000 IPU Manual Sysmex Training Manual |

BODY FLUID ANALYSIS

ON THE SYSMEX® XE-5000

# PRINCIPLE

The Body Fluid Analysis Mode of the Sysmex XE-5000 adds a quantitative automated procedure for analyzing cerebrospinal fluid, serous fluid, and synovial fluid.  The 4-DIFF scattergram utilizes fluorescent flow cytometry using lateral scattered light and lateral fluorescent light in a specialized analysis sequence to calculate and display the total nucleated (TC-BF) counts, mononuclear (MN) cells (includes lymphocytes, monocytes, and other mononuclear cell types) and polymorphonuclear (PMN) cells (includes neutrophils, eosinophils, and basophils) counts and percentages.  The direct current (DC) detection method is used to determine the RBC (RBC-BF) count.  Examination of the numerical and/or morphological findings in the body fluid are useful in the diagnosis of disease states, such as meningitis, hemorrhage, malignancy, inflammation, viral, bacterial, and parasitic infections.

# SPECIMEN

- Requirements

- Specimen collection:

- The recommended collection tube is potassium EDTA.

Plain Red to tube is acceptable if no visible clot is noted.

- Cerebrospinal (CSF) does not require anticoagulant and is usually collected in special CSF collection tubes.

- All specimens collected should be free of clots.

- Extremely viscous fluids can be treated with small amounts of hyaluronidase added to an aliquot of the sample.

- Specimen volume required

- A minimum of 1 ml of sample is required for analysis.

- Unacceptable specimens:

- Clotted samples should not be tested.  Clotted or sub-optimal specimens unsuitable for counting should be microscopically examined for WBC’s and RBC’s using a glass slide and coverglass.  Report as no, rare, few, moderate or many WBCs or RBCs present with the comment “clotted specimen, unable to perform an accurate cell count.” If the sample was submitted in the wrong tube type, the comment “sample collected improperly” should also be reported.

- Characteristics that may affect test results:

1.   Clotted samples

2.   Synovial samples that contain uric acid crystals or have a high viscosity.

E.	Stored Specimen Stability

- All fluids should be tested as soon as possible after collection according to CLSI H56-A.  CSF should be tested within 4 hours of collection.  All other body fluids should be tested within 24 hours of collection.

# SUPPLIES & REAGENTS

- Supplies

- De-ionized water

- Lint-free plastic lined lab wipes

- Gauze

- Test tubes

- Clorox™ bleach (use when CELLCLEAN is indicated)

- Sysmex reagents

- Tri-level commercial controls, e-CHECK(XE)

- Cell-Chex Body Fluid Cell count Control, L1-UC and L2

- Hyaluronidase

- Sysmex® Reagents

The same instrument is used for peripheral blood counting.  No additional reagents are required for body fluid testing.  See procedure Complete Blood Count of Whole Blood on the Sysmex XE-5000.

# OPERATING PROCEDURE

- Start-Up Procedure

- Check physical status.

- Check and drain pneumatic trap chamber. (Refer to section 9.2 of the          XE-5000 Instructions for Use for detailed information).

- Check reagent boxes for sufficient run volume.

- Check printer paper supply.

- Power-Up Sequence

- Press power switch on IPU – Information Processing Unit.  Must log on before powering up the Main unit.

- Log-on the IPU

- XE-5000 program log-on box appears.  Type User Name, HEME and press [ENTER].

- Press the power switch on Main Unit.  The Pneumatic Unit is controlled by the Main unit, so Pneumatic Unit power is left on.  The instrument automatically performs self check on the:

Microprocessor	Mechanical parts

Temperatures	Background counts

- Press the power switch on the printer.

- When the “Logon Name and Password” screen displays, log on to the Main Unit.

- Press [NUM/ALPH] on the Main Unit to display alpha characters (if alpha characters are used for Logon Name).

- Press the appropriate keys on the Main Unit keypad, HEME for Logon Name.

- Press [ENTER] for the password.

- Check background check and trap fluid level on maintenance log (See chart of background acceptable limits below).

| XE-5000 Acceptable Background Counts |  |
| --- | --- |
| Parameters | Acceptable Limit |
| WBC | 0.1 x 103/ μL |
| Diff – WBC | 0.2 x 103/ μL |
| IMI – Total | 0.3 x 103/ μL |
| IMI# | 0.005 x 103/ μL |
| NRBC-WBC | 0.2 x 103/ μL |
| RBC | 0.02 x 106/ μL |
| HGB | 0.1 g/dL |
| PLT | 5 x 103/ μL |
| PLT-O | 10 x 103/ μL |

- Analyze Quality Control if required.

- Patient Sample Processing

- MANUAL MODE (130 μL sample volume)

- Make sure the instrument is in the Ready status.  The Ready LED light should be lit.

- Press [MANUAL] key on the Main Unit keypad.

- Enter the specimen number (up to 15 alpha/numeric characters), using the Main Unit keypad or read the barcode using the handheld barcode scanner.

- Press [▼] to change the setting parameter to “Mode” and then [] or [] keys to set Manual if it is not already selected.

- Press [▼] to change the setting parameter to “Discrete” and then [] or [] keys to select the following:

● 3 - CBC + DIFF

- Press [▼] to change the setting parameter to “Sample” and then [] or [] keys to set to  3: Body Fluid.

- After all the settings are complete press [ENTER] to confirm.

- When [ENTER] key is pressed, a background check is initiated.  The background check analysis is repeated three times for the background check.  Results must be less than or equal to the values in the table below:

| WBC-BF | 0.001 x 103/µL |
| --- | --- |
| RBC-BF | 0.003 x 106/ µL |

- If the background does not fall below these limits, the message “Background Error” will be displayed and the sample explorer screen will show the “F” with a red background in front of the background run. Select [AUTORINSE] on the function menu of the Background Check screen to perform an automatic cleaning.  See Chapter 10 for additional information on resolving Background Errors.

- If the background check is acceptable, the READY LED lights and the Main Unit displays body fluid analysis ready status.

- Mix the contents of the test tube gently but thoroughly. Uncap the tube.

- Place sample under the aspiration pipette so that the tip of the pipette is at the bottom of the sample.

- Press [Start] switch.  Remove the sample when 2 beeps sound, or the green Ready LED stops blinking.  Patient results will print as samples are completed if auto-output is selected.

- After the sample results are transmitted to the Sysmex computer screen the results must be validated using the Validate icon. This can only be done if the Last 20 icon is turned OFF.

- The results are then transmitted to the Meditech computer system.

- When the analyzer RBC count is <1000, a blank will cross over to Meditech. Change the blank RBC field to “manual” and enter the manual RBC count result. A pop up will ask if you performed a manual count. If it is, then answer “yes”and verify. If you answered “no” the field will be left blank.

CAUTION:	      If either of the analysis parameters is high:

WBC-BF > 1000 /ul   or   RBC-BF > 1,000,000/ µL,

there is the possibility of carryover on the analysis results of the next sample and the “Execute Background Check” error is displayed.  Analyze the next sample after pressing the [B-check] function menu on the Main Unit.  Failure to do so will result with the sample being flagged with an “F” with a red background on the IPU in the Explorer and Browser screens.  Samples flagged with a red “F” may have counts which are falsely increased.

- Once body fluid analysis is complete, return the analyzer settings to

Sample No  ● 1

Mode  ● 1 Manual

Descrete  ●  7  - CBC + DIFF + NRBC + RET

Sample  ●  1:NORMAL

# CALCULATIONS

If a dilution was made prior to processing, multiply the results by the appropriate

dilution prior to reporting. Reportable result  =  Analyzer result  x dilution factor.

# REPORTING RESULTS

- REFERENCE RANGE

Cerebrospinal Fluid

| Parameter | Reference Range |  |
| --- | --- | --- |
|  | Neonate | Adult |
| WBC-BF | 0 – 30 / uL | 0 – 5 / uL |
| RBC-BF |  |  |
| MN % |  |  |
| PMN % | 0 - 8 % | 2 – 4 % |
| MN # |  |  |
| PMN # |  |  |
| TC-BF |  |  |

Synovial Fluid

| Parameter | Reference Range |
| --- | --- |
| WBC-BF | 0 – 200 / uL |
| RBC-BF | 0 / uL |
| MN % |  |
| PMN % | <25 % |
| MN # |  |
| PMN # |  |
| TC-BF |  |

Serous Fluid

| Parameter | Reference Range |  |
| --- | --- | --- |
|  | Pleural | Peritoneal |
| WBC-BF | <1000 / uL | <300 / uL |
| RBC-BF |  |  |
| MN % | 90 - 100 % |  |
| PMN % |  |  |
| MN # |  |  |
| PMN # |  |  |
| TC-BF |  |  |

# PROCEDURE NOTES

- Pay special attention to the RBC-BF result. The RBC counts are ONLY reported out in increments of 1000 (i.e. 1000, 2000, 3000, 4000, etc.). If the analyzer RBC count is <1000, a manual RBC count must be performed using the hemacytometer method.

- The result entry mode must be changed in Meditech to “manual”, when a manual count is performed.

- ALL clear CSF fluids MUST be enumerated for RBC by the classical hemacytometer method.

- Results analyzed in the Body Fluid Mode are not subject to automatic validation, regardless of the setting.  Check the display status of the “F” mark and perform validation carefully.

- Do not place samples on a mechanical rocker.  Excessive mixing may induce platelet clumping and alter white cell membranes resulting in false interpretive messages.

- Clorox, a filtered bleach, is recommended for use in cleaning.  If Clorox is not available, generic bleach may be used but must be 5% Sodium Hypochlorite concentration and be free of particles that may cause background contamination when used on the analyzer.  Clorox must be diluted to a 5% Sodium Hypochlorite concentration before preparing further dilutions recommended for maintenance.

- For troubleshooting specifics refer to the Sysmex XE-5000 Instructions for Use.

# LIMITATIONS OF PROCEDURE

- XE-5000 MANUFACTURER STATED LINEARITY

| Parameter | Range | Units |
| --- | --- | --- |
| WBC-BF | 0-10,000 | /μL |
| RBC-BF | 0-5,000,000 | /μL |

- Parameters that exceed these limits are flagged with @ beside the result.  The sample must be diluted, rerun and multiplied by the dilution factor.

- Note the use of dilution for linearity on the patient report.

- POSSIBLE SAMPLE INTERFERENCES

- Sample results with errors related to WBC and RBC parameters should be enumerated by the classical method on a hemacytometer. A smear should be made by Cytocentrifugation and a differential cell count performed. See procedure CORE 6555 R Body Fluid Cell Counts.

- Results may be compromised with clotted samples and synovial samples that contain uric acid crystals or have high viscosity.

- All fluids should be tested as soon as possible after collection.

- Further testing should be performed if more specific or sensitive limitation information is needed.

# CALIBRATION

The same instrument is used for peripheral blood counting.  No additional calibration is required for body fluid testing.  See procedure Complete Blood Count of Whole Blood on the Sysmex XE-5000.

# QUALITY CONTROL

Two Body Fluid Cell Count Control specimens, L1and L2, will be analyzed and results recorded daily. This will be performed by the day technologist assigned to the Sysmex bench.  The results will be recorded on the Body Fluid Control Sheet in the Sysmex maintenance Binder.

# MAINTENANCE

The same instrument is used for peripheral blood counting.  No additional

maintenance is required for body fluid testing.  See procedure for Complete

Blood Count of Whole Blood on the Sysmex XE-5000.

# REFERENCES

- Sysmex XE-5000 Instructions for Use (North American Edition), Sysmex Corporation, Kobe, Japan, July, 2007.

- Sysmex XE-5000 Software Guide (North American Edition), Sysmex Corporation, Kobe, Japan, July, 2007.

- Clinical and Laboratory Standards Institute (CLSI).  Laboratory Documents: Development and Control; Approved Guideline; Fifth Edition. (GP2-A5, 2006).

- College of American Pathologists (CAP) Hematology-Coagulation Checklist, October 2007.

- Body Fluids, 3rd ed.(1993), ed. by Kjedsberg and Knight, ASCP Press