PURPOSE

To define how to analyze patient whole blood samples on the GEM Premier 4000 instrument with Intelligent Quality Management.

BACKGROUND

The GEM Premier 4000 analyzer with iQM is a portable critical care system used to rapidly analyze whole blood samples to provide quantitative measurements of pH, blood gases, sodium, potassium, chloride, ionized calcium, glucose, lactate, hematocrit, and CO-Oximetry (tHb, O2Hb, COHb, MetHb, and HHb).  These parameters, along with derived parameters, aid in the diagnosis of a patient’s acid/base status, oxygen delivery capacity, and electrolyte and metabolite balance.

EQUIPMENT

GEM Premier 4000 Instrument

Blood Transfer Device

3 ml syringe with luer lock

REAGENTS

GEM Premier 4000 PAK® Cartridge with iQM

REAGENT STORAGE

Cartridges are stored at room temperature and are stable until the expiration date listed   on the package.  Once placed on the instrument the cartridge is stable for 30 days or the indicated number of tests, whichever is first.

CALIBRATION

The GEM Premier 4000 automatically performs three types of calibrations: one-point calibrations, two-point calibrations, and low oxygen calibrations.  During these calibrations, the instrument adjusts each sensor’s performance to ensure correct operation.  Calibrations will be performed after each analysis, periodically during idle time, and on demand when needed. No operator intervention is required.

SPECIMEN COLLECTION AND HANDLING

| TEST | VOLUME | TUBE | SAMPLE HANDLING |  | TIME LIMIT |
| --- | --- | --- | --- | --- | --- |
| Lactic Acid | 150 uL | Na Oxalate/ K Fluoride (Gray) | Whole Blood Room Temperature |  | Test within 15 minutes |
|  |  |  | Whole Blood Received on ice |  | Test within 2 HOURS |
| TEST | VOLUME | TUBE | SAMPLE HANDLING |  | TIME LIMIT |
| Ionized Calcium and pH | 150 ul | Lithium heparin | Whole Blood Room Temp or on ice |  | Test within 3 hours. |
|  |  |  | Spun and plasma refrigerated |  | Test plasma within 24 hours |
| Sodium Potassium Chloride TC02 (calculated) Hemoglobin Hematocrit Glucose | 150 ul | Lithium heparin | Whole Blood Room Temp | Received and tested within 15 minutes |  |
|  |  |  | Whole Blood on ice | Test within 3 hours |  |
|  |  |  | Spun and plasma refrigerated | Test plasma within 8 hours; note: Hgb/HCT invalid |  |
| Carboxyhemoglobin | 150 ul | Lithium heparin | Whole Blood Room Temp or on ice | Received and tested within 30minutes |  |
| Methemoglobin | 150 ul | Lithium heparin | Whole Blood Room Temp or on ice | Received and tested within 30minutes |  |

UNACCEPTABLE SPECIMENS

A specimen is considered unacceptable if it is not analyzed within the validated time frames listed above, is not labeled with two FHS patient identifiers, is in the incorrect anticoagulant,  or if specimen is clotted.  Specimens should not have been exposed to air prior to testing.  Blood transfer devices should be removed from the sample after testing to prevent air exchange.

Sample Preparation

Before analyzing the blood sample, it is important that it be mixed properly. Invert the sample tube gently for 5 seconds before sampling. Full tubes are preferred. Any samples in syringes, that are not tested immediately, should be inverted and rolled in between the palms before aspiration.

STEPS

Select Normal by pressing the “Normal/Micro” toggle button on the upper left corner of the screen.  Note: Micro sampling is not appropriate for whole blood venous samples.

SAMPLING STEPS WHEN CERNER INTERFACE UP

- From Main GEM Screen

- Scan the patient Cerner barcode label. The Order Detail screen should come up.

- Verify the patient information on the Order Detail screen.

- Press the “RUN THIS ORDER” button on the screen.

- Mix the sample thoroughly by gently inverting the tube 5-10 times. If clots are suspected, check the sample.  Obtain another sample if needed.

- Press “GO!”

- Sampler will emerge from its home position.

- The sample probe will emerge from its home position.

- Bring the opening of the syringe to the probe.  Ensure that the probe extends far enough into the syringe and sample to aspirate 150 uL of the sample. Press OK.

- The analyzer will aspirate the sample and then display “Remove Sample” when sampling is completed.

- Press “View Results” Acceptable results will be sent across the Cerner interface and verify the result in Cerner TSA menu for your facility’s workcenter and testsite.

SAMPLING STEPS WHEN THE CERNER INTERFACE IS DOWN

- Select test or panel of tests to be performed.

- Specify the sample type/container combination in section 2 of the menu.

- Mix the sample thoroughly by gently inverting the tube 5-10 times. If clots are suspected, check the sample.  Obtain another sample if needed.

- Place a blood transfer device onto the vacutainer tube.

- Invert the blood tube gently to mix.  Connect the syringe to the connector on the transfer device.

- Withdraw sample from the tube using the syringe and blood transfer device, taking care not to draw bubbles, or dead air volume, into the syringe.

- Press “GO!”

- The sampler will emerge from its home position.

- Present the syringe with blood sample to the Gem sampler, and press OK.

- The sampler should be inserted far enough into the syringe to allow 150 uL to be aspirated, but not so far that the sampler touches the bottom of the syringe.

- Make sure that the sampler is below the level of the blood in the syringe.

- The system will aspirate the sample and emit an audio prompt when aspiration is complete.  Remove the container promptly.  The probe will retract into the system.

- Select the “Enter Information” tab. Enter the Operator ID (Tech ID) and scan the LIS barcode label into the sample ID field.  Enter the full name for two patient identifiers.  Comments may also be entered if required.

9.   Dispose of sample syringes in a biohazard waste container once testing is completed.

- “View Results” cannot be accessed until all required fields are completed.  Press “Accept”. The results will then print. Staple printout to RQW and enter results into LIS manually in Cerner TSA menu for your facility’s workcenter and testsite.

RESULT NOTES

- St. Elizabeth Hospital--Use existing LIS Entry Processes

- Other exception flags may be found in the Operator’s Guide, section IX.

- Sample results that should not be included in the data base can be excluded by pressing     “Exclude”.  The operator then may rerun the sample or collect a new sample to run.

- To search the GEM 4000 database for patient history, review the Operator’s Guide section X.

NORMAL REFERENCE RANGE

| ANALYTE | AGE | NORMAL RANGE | CRITICAL |
| --- | --- | --- | --- |
| Lactic Acid, whole blood | Adult | 0.5-2.2 mmol/L | > 4.0 |
| NA, whole blood | Adult | 135-153 mmol/L | <120, >155 |
| K, whole blood | Adult | 3.5-5.3 mmol/L | <3.0, >6.0 |
| CL, whole blood | Adult | 98-109 mmol/L |  |
| CA, ionized whole blood | Adult | 1.11-1.33 mmol/L | <0.89, >1.51 |
| pH, blood | Adult | 7.30-7.45 |  |
| COHb | Adult | 0.5-1.5 % |  |
| MetHb | Adult | <3.0 |  |
| CO2, whole blood | Adult | 22-32 mmol/L | <10, >45 |
| GLU, whole blood | 2 Days | 40-90 | <40, >200 |
| GLU, whole blood | Adult | 65-99 | <50, >400 |
| HCT, whole blood | Adult- Male | 38-49% | <22, >65 |
| HCT, whole blood | Adult-Female | 36-43% | <22, >65 |
| HCT, whole blood | Pediatric ranges | See Cerner ITD |  |

LIMITATIONS

- Results that appear to be inconsistent with patient history, therapy, or condition should be viewed as questionable and the test should be repeated on the laboratory backup method.

- Hemolysis noted after reporting of results can be manually ordered in Cerner using test code  “HEMOL CHK”.  Follow normal department procedures and follow-up actions for hemolyzed samples.

FHS ANALYTICAL MEASUREMENT RANGE (AMR)

| WHOLE BLOOD ANALYTE | Reportable Range |
| --- | --- |
| NA+ | 106.0-174.0 mmol/L |
| K+ | 1.2-10.10mmol/L |
| CL- | 69.0-152.0 mmol/L |
| Glu | 16.0- 691.0 mg/dL |
| HGB | 6.3-20 g/dl |
| HCT | 18-61% |
| LAC | 0.4- 16.0 mmol/L |
| Ph | 6.90-7.74 |
| COHb | 0.1-24.5% |
| MetHb | 0.1-27.3% |
| pCO2 | 19.0-121.0 mmHg |
| pO2 | 37.0-572.0 mmHg |
| CA++ | 0.24-2.98 |
| tHb | 6.3-20.0 |

INTERFERING SUBSTANCES AND LIMITATIONS FOR GEM

Samples with questionable results on the GEM will be spun down, and immediately run on the DXC  for chemistries.

| Li Limitations | Description |
| --- | --- |
| Metabolic Changes | Errors can occur due to metabolic changes if there is a delay in the measurement of the samples. |
| Elevated White Blood Cells or Reticulocyte Counts | Samples will deteriorate more rapidly, even when kept in ice water. |
| Improper Mixing | Errors will be introduced for measurement of hematocrit and CO-Ox parameters if the sample is not properly mixed prior to measurement. |
| Exposure to Air | Tubes should be full to minimize the exposure to air ; Tubes should remain capped until testing; blood transfer devices should be removed after sampling to prevent air exposure. |

The following substances may show noticeable interference with certain channels on the GEM Premier 4000 analyzer, causing falsely elevated results.

| Substance | Affected Analyte | Substance Concentration Producing Interference |
| --- | --- | --- |
| Benzalkonium* | Ca++ | 5 mg/L |
| Bromide | Cl¯ | 10 mmol/L |
| Cyanomethemoglobin** | CO-Oximetry | >4% |
| Dopamine | Glucose, lactate | 5 mg/dL |
| Dobutamine | Glucose, lactate | 2 mg/dL |
| Fluoride | Cl¯, lactate | 500 mg/dL |
| Glycolic acid | Lactate | 1 mmol/L |
| Hemoglobin Based Oxygen Carriers (Hemopure***) | Hematocrit | 3.2 g/dL |
| Hydroxyurea | Glucose, lactate | 0.8 mg/dL |
| Iodide | Cl¯ | 3 mmol/L |
| Isoniazide | Glucose, lactate | 5 mg/dL |
| Oxalate | Cl¯, lactate | 500 mg/dL |
| Salicylate | Cl¯ | 4 mmol/L |
| Sulfhemoglobin** | CO-Oximetry | >3% |
| Thiopental* | PCO2, Ca++ | 30 mg/L |
| Turbidity** | CO-Oximetry | 5% based on turbidity created by Intralipid**** fat emulsion |
| Uric acid | Glucose | 20 mg/dL |

*The GEM Premier 4000 analyzer with iQM employs failure pattern recognition checks.  These checks include detecting the presence of positively charged lopophilic compounds (e.g., benzalkonium) and negatively lipophilic compounds (e.g., thiopental).  The GEM Premier 4000 analyzer offers the operator the ability to enable flagging of patient results if interference patterns for these compounds are detected by iQM.

**CO-Oximetry interference is detected and flagged by failure pattern recognition checks.

***Hemopure is a registered trademark of Biopure Corp.

****Intralipid is a registered trademark of Fresenius Kabi AB.

The following tested drugs may interfere with glucose determination, causing falsely low readings:

| Drug | Interference Observed |
| --- | --- |
| Oxalate | 1000 mg/dL |
| Fluoride | 500 mg/dL |

REFERENCES

Instrumentation Laboratory GEM Premier 4000 Reference Guide, October 2006.

Instrumentation Laboratory GEM Premier 4000 Training Guide, October 2006

Instrumentation Laboratory GEM Premier 4000 Configuration Guide, October 2006.

Instrumentation Laboratory GEM Premier 4000 Data Management Guide, October 2006.

Instrumentation Laboratory GEM Premier 4000 Operator’s Guide, October 2006.