PURPOSE

To provide work instructions for the performance of manual microscopic analysis of
urine sediment

BACKGROUND

Urine microscopic analysis is performed when blood, nitrite, leukocyte esterase, protein  are positive or if urine appearance is anything other than clear. It is also performed at
SJMC lab to confirm unusual microscopic elements that have been detected on the
iQ200 microscopic instrument or specifically requested by the physician. 

SPECIMEN COLLECTION

Minimum of 10ml of fresh urine, collected in a clean container or added to transport 
preservative tubes.

LIMITATIONS

1.  Testing must be done within 2 hours of receipt of specimen if specimen is not  
      refrigerated.  If there is a delay the specimen must be refrigerated or added to
      preservative, boric acid.  Significant delays (>24hrs) can cause degeneration of 
      certain cellular elements such as WBC's or casts and overgrowth of bacteria. 

2.  Urine volume can significantly affect microscopic quantification, urines with less
     than 5 ml, must be footnoted with a variance statement.  Using the TV code,

add the PF1/ P footnote. Urines that are QNS to centrifuge will have a comment 
     added that the microscopic was performed on unspun urine. 

SUPPLIES

Urisystem slides with coverslips or chambers.

STEPS

1.  Manual microscopy, either automated or manual is to be performed when the 
     following criteria has been met on the Urinalysis instrument:

- Blood, trace or more

- Leukocyte esterase, 1+ or more

- Nitrite is positive

- Protein, 1+ or more

- Clarity, anything other than clear

2.   Add 10-12mls of well mixed urine to a KOVA tube, cap and centrifuge at 2000RPM 
      for 5-10 minutes.

3.  After centrifuging is completed, place a KOVA pipette into the bottom of the tube
     and discard the supernatant above the bulb of the pipette into the sink.

4.  Mix the sediment well by squeezing the top of the pipette several times.

5.  Discard the first drop of sediment and deliver the next drop into the notched groove
     of one segment of the UriSystem slide.  The slide will fill automatically by capillary
     motion.  Add more drops if necessary to fill the slide segment.  Slight overfill will
     not affect the results.

6.  Use the urine microscopic guide in this work instruction to identify and quantitate
     urine microscopic elements. For identification of other elements that are unusual
     or not listed, refer to the resource books available at the urinalysis bench or
     the IRIS atlas.

7.  Normal urine sediment is not free of cells or casts but may contain a limited 
     number of formed elements.  A precise definition of normal is hard to define, but 
     the presence of one or two blood cells per high power fields, one or two leukocytes 
     and a few epithelial cells is not necessarily considered abnormal.  The urine of
     mature females may also contain large numbers of squamous epithelial cells
     from the vaginal walls.  An occasional hyaline cast may also be a normal

finding.

8.  Sperm seen on the microscopic exam will be reported regardless of age or
     sex of the patient.  On all females, if sperm is seen or suspected, a second
     tech must review the slide or iQ results before results are verified. First verify that the

specimen has been labeled correctly.

When a second tech is not immediately available and:

- If the patient is expected to be admitted to an inpatient bed, wait for a second tech opinion and call a preliminary verbal result if the patient is <16 years old
>80 years old or thought to be in a vulnerable environment.  Enter a final result if the second tech has confirmed the initial result.

- If the patient is not expected to be admitted and a second tech is due < 2hours, wait on the second opinion and call a preliminary verbal result (using age ranges above). Enter a final result if second tech has confirmed the initial result.

- If the patient is not expected to be admitted and a second tech is not due in <2 hours, consider sending the urine sample to SJMC for a stat review. Once the second opinion is complete and matches the first review, verify the result with a comment that indicates: Confirmed by two techs.

- If the second opinion does not match the first, request a manager/MTC or 
            pathologist to review ASAP.

9.  All urines that have a microscopic exam that is either ordered or reflexes
     to a microscopic, must enter 4 mandatory fields. These are: RBC's, WBC's
     Epithelial cells and bacteria.

10. Urine specimens that are positive for nitrite, leukocyte esterase and have greater
     than 4 WBC's/hpf will also be set up for a urine culture if a reflexive test has been
     ordered or culture if indicated has been ordered. 

11. Microscopic findings must always correlate with the macroscopic (chemistry)
     results.

- If there is a presence of protein, casts, RBC's and many epithelial cells can be seen on the microscopic exam.

- A positive occult blood on the chemistry test, RBC's are 
also seen on the microscopic unless there is hemolysis.

- If RBC's are seen on microscopic, but the chemistry result is negative or if the chemistry is positive and the microscopic is negative perform the Hemastix test or use Multistix to confirm the chemistry result.

- Positive leukocyte esterase would be an indication of WBC's and should
also be found on the microscopic.

- Positive nitrite would indicate the presence of bacteria.  The bacteria should
be seen on the microscopic exam.

- A urine that has a cloudy appearance may have cellular structures or bacteria
present on the microscopic.

- Low specific gravity (hypotonic) can cause RBC's to lyse.

- If the specimen is not centrifuged or decanted properly, there may be a
discrepancy between the chemistry strip and the microscopic results.

- The specimen is compromised.

If these microscopic/macroscopic correlations are not seen, the urinalysis will be
repeated.  A reason for the discrepancy must be noted under the footnote section. This is a free text section. Add comments about why there is a discrepancy, for example,  specimen too old or not refrigerated. If there is any question about the
quality of the specimen, the urine should be recollected.

| Sediment element | Magnification | Description | Quantification |
| --- | --- | --- | --- |
| RBC | 40X | Pale discs, approx. 7 microns in diameter | 0-2 3-5 6-10 11-25 26-50 TNTC |
| WBC | 40X | Round granular 7-12 microns in diameter, twice the size of RBC's | 0-4 5-10 11-25 26-50 TNTC |
| BACTERIA | 40X | Rods or cocci in chains (appear like beads in a row) | NONE Few 1+ 2+ 3+ 4+ |
| EPITHELIAL CELLS | 40X | Large irregular membrane and well defined central nucleus | 0-4 5-8 9-15 16-30 TNTC |
| Sediment element | Magnification | Description | Quantification |
| Transitional epithelial cells | 40X | Smaller than epithelial cells with a less irregular membrane | NONE 1-5 6-10 11-25 26-50 TNTC |
| Renal epithelial cells | 40X | Smaller than transitional cells, round and regular cell membrane | NONE 1-5 6-10 11-25 26-50 TNTC |
| Yeast, hyphae or budding | 40X | Egg shaped | NONE or PRESENT |
| Mucous | 40X | Appears as strands either individually or in clumps | NONE or PRESENT |
| Trichomonas | 40X | Same size as a WBC with motile flagella | NONE or PRESENT |
| Spermatozoa | 40X | Elongated head with long thin tail | NONE or PRESENT |
| CASTS | Count on 10X ID on 40X | Parallel sides, round or blunted ends, vary in shape and size | NONE 0-2 3-5 6-10 11-20 >20 |
| Hyaline Cast | Count on 10X ID on 40X | Colorless, homogeneous, semi-transparent | See above |
| Granular cast | Count on 10X ID on 40X | Contains fine or coarse granules | See above |
| WBC cast | Count on 10X ID on 40X | Cast with formed WBC's inside the cast | See above |
| RBC cast | Count on 10X ID on 40X | Cast with formed RBC's inside the cast May contain hemoglobin | See above |
| Broad, Waxy casts | Count on 10X ID on 40X | Waxy, yellowish, wide Can mean renal failure | See above |
| Cylindroid cast | Count on 10X ID on 40X | Tapers to a slender tail | See above |
| Sediment | Magnification | Description | Quantification |
| Crystals from Acid urine |  |  | NONE 1+ 2+ 3+ 4+ |
| Uric Acid | 40X | Yellow/brown, flat diamond, needle , hexagonal plates or rosettes | See above |
| Amorphous urates | 40X | Yellow-red granules | See above |
| Acid urates | 40X | Brown spheres | See above |
| Calcium oxalate | 40X | Refractile, colorless Octahedron or dumbbell shape. "Envelopes" | See above |
| Bilirubin | 40X | Red-brown needles or granules. Bilirubin is positive | See above |
| Crystals from Alkaline urine |  |  | NONE 1+ 2+ 3+ 4+ |
| Triple phosphate | 40X | Colorless, 3-6 sided prisms "coffin-lid" forms | See above |
| Ammonium biurate | 40X | Yellow spheres | See above |
| Calcium phosphate | 40X | Stellate prisms | See above |
| Calcium carbonate | 40X | Colorless spheres or dumbbells | See above |
| Amorphous | 40X | Small particulate sediment | None or present |
| Abnormal crystals |  |  | NONE 1+ 2+ 3+ 4+ |
| Cystine | 40X | Colorless, refractile, hexagonal plates | See above |
| Tyrosine | 40X | Yellow, sometimes black needles in sheaves,rosettes | See above |
| Leucine | 40X | Yellow, speres with radial striations | See above |
| Sulfapyridine | 40X | Arrowheads | See above |
| Sediment | Magnification | Description | Quantification |
| Sulfathiazole | 40X | Symmetrical crystals appearing as sheaves of wheat with central binding, rosettes or hexagonal plates | See above |
| Ampicillin | 40X | Long, slender needles | See above |
| Cholesterol | 40X | Large, flat plates with notched corner | See above |
| Radiographic dyes | 40X | Colorless, needles | See above |
| Starch granules | 40X | Donut shapes,maltese cross | See above |
| Oval fat bodies | 40X | Highly refractile,yellow brown, maltese crosses, sometimes found within other cells. | None or present |

REFERENCE
1.  Sister Laurine Graff, A Handbook of Routine Urinalysis, 1983. Philadelphia, PA. 
2.  Strasinger, Susan  Urinalysis and Body fluids, ed. 4.  Philadelphia, PA, 2001