|  | FORM |
| --- | --- |
|  | M-F-CH-1940-05 |
| DXC 600 (AMR) aNALYTICAL MEASUREMENT RANGE |  |
| St. Joseph Medical Center Tacoma, WA St. Clare Hospital Lakewood, WA St. Elizabeth Hospital Enumclaw, WA St. Francis Hospital Federal Way, WA St. Anthony Hospital Gig Harbor, WA PSC |  |

| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ACTM | 10 – 300 ug/mL | 10-600µg/mL | Saline | X2 (MAX dilution) |  | >600 µg/mL | <10 µg/mL |
| ALB | 1.0 – 7.0 g/dL | 1.0-14.0g/dL | Saline | X2 (MAX dilution) | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >14.0 g/dL | < 1.0 g/dL |
| ALP | 5 – 1650 IU/L Straight (5-1000) Ordac (800-1650) | 5-8250 IU/L | Saline | X5 (MAX dilution) |  | >8250 IU/L | < 5 IU/L |
| ALT | 5 – 2600 IU/L Straight (5-400) Ordac (350-2600) | 5-13000 IU/L | Saline | Dilute “OIR LO” X2 X5 (MAX dilution) |  | >13000 IU/L | < 5 IU/L |
| AMM | 9 – 1000 µmol/L | 9 – 1000 µmol/L | N/A | DO NOT DILUTE | 1+ lipemia: Run ultracentrifuged specimen. | >1000 µmol/L | < 9 µmol/L |
| AMY7 | 5 – 2000 U/L Straight (5–1200) Ordac (1000-2000) | 5-10000 U/L pl/serum 5-202,000 U/L urine | Saline | X5 (MAX dilution) for plasma/serum |  | >10000 U/L for plasma / serum | < 5 U/L plasma/ serum |
| AST | 5 – 2600 IU/L Straight (5-400) Ordac (350-2600) | 5-13000 IU/L | Saline | Dilute “OIR LO” X2 X5 (MAX dilution) |  | >13000 IU/L | < 5 IU/L |
| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| BUN | 5 – 100 mg/dL serum/plasma | 5 - 300mg/dL plasma and serum | Saline | X3 (MAX dilution) Plasma/serum | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >300 mg/dL plasma/ serum | < 5 mg/dL plasma & serum |
| CA | 2 - 20 mg/dL plasma/serum | 2-40 mg/dL plasma and serum | DI water (NERL) | X2 (MAX dilution) Plasma/serum | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >40 mg/dL plasma/ serum | <2 mg/dL plasma/ serum |
| CAR (TEG) | 2-20 µg/mL | 2-40 µg/mL | Saline | X2 (MAX dilution) | Reanalyze specimens reported out as SUPPRESSED due to RXN ERROR | >40 µg/mL | <2 µg/mL |
| CK | 5 – 4100 IU/L Straight (5-1200) Ordac (860-4100) | 5-20500 IU/L | Saline | X5 (MAX dilution) |  | >20500 IU/L | < 5 IU/L |
| CL | Serum/plasma/BF 50-200 mmol/L, Serum/plasma | 50-200 mmol/L serum/plasm | N/A | DO NOT DILUTE serum/plasma | 3+ lipemia or Lipemia Index >10: Run ultracentrifuge specimen. | >200 mmol/L serum/plasma | < 50 mmol/L Serum and plasma |
| CO2 | 5 – 50 mmol/L | 5-100 mmol/L | DI water (NERL) | X2 (MAX dilution) | 3+ lipemia or Lipemia Index >10: Run ultracentrifuge specimen. | >100 mmol/L | < 5 mmol/L |
| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| CREA | Serum/plasma - – 25 mg/dL Urine 10 – 400 mg/dL (for DSUs) | 0.3-50 mg/dL serum and plasma Urine 10-400 mg/dL (for DSUs) | Saline N/A | X2 (MAX dilution) serum/plasma DO NOT DILUTE URINE | 3+ lipemia: Run ultracentrifuge specimen. Send all urine creats except for DSUs to SJMC | >50 mg/dL serum/ plasma >400 mg/dl Urine (for DSUs) | < 0.3 mg/dL <10 mg/dL urine (for DSUs) |
| D BIL | - – 10.0 mg/dL | 0.1-20 mg/dL | Azide free human serum albumin | X2 (MAX dilution) |  | >20 mg/dL | < 0.1 mg/dL |
| ETOH | 0.01 – 0.6 g/dL plasma/serum/ urine | 0.01-1.2g/dL | ETOH Cal 1 | X2 (MAX dilution) | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. Centrifuge cloudy urines. | >1.2 g/dL plasma/ serum/urine | <0.01 g/dL plasma serum and urine |
| GGT | 5 – 3000 IU/L Straight (5 – 750) Ordac (550 – 3000) | 5-15000 IU/L | Saline | X5 (MAX dilution) |  | >15000 IU/L | < 5 IU/L |
| GLUC(m) | 3-1200 mg/dL CSF/BF/Urine Straight (3-600) Ordac (300-1200) | 3-2400mg/dL | Saline | X2 (MAX dilution) | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >2400 mg/dL Plasma/CSF/ BF/Urine | <3 mg/dL Plasma/CSF BF/Urine |
| K | 1.0-15.0 mmol/L Serum/plasma | 1.0-15.0 mmol/L | N/A | DO NOT DILUTE Serum/plasma | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >15.0 mmol/L Serum/ plasma | <1.0 mmol/L Serum and plasma |
| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| LACT | 0.3– 11.0 mmol/L Plasma/CSF | 0.3 - 22.0 mmol/L | Saline | X2 (MAX dilution) |  | >22.0 mmol/L | <0.3 mmol/L |
| LD | 5 – 2700 IU/L Straight (5-750) Ordac (600-2700) | 5 – 13500 IU/L | Saline | X5 (MAX dilution) |  | >13500 IU/L | < 5 IU/L |
| LIP LIPASE | 10 – 400 U/L Straight (10-200) Ordac (180-400) | 10-2000 U/L | Patient sample of known low lipase value | Dilute specimens with “out of range high values” with patient sample of known low lipase value and reanalyze. X5 (MAX dilution) | 3+ lipemia or Lipemia Index >10. Run ultracentrifuged specimen | >2000 U/L | <10 U/L |
| LITHIUM | 0.1- 3.0 mmol/L | 0.1 – 6.0 mmol/L | Saline | X2 (MAX dilution) | ORDAC NOT TURNED ON. EDTA the only acceptable anticoagulant. | >6.0 mmol/L | <0.1 mmol/L |
| M-TP | 6-300 mg/dL CSF | 6-600 mg/dL CSF | Saline | X2 (MAX dilution) CSF | IF serum protein carryover is suspected, run saline blanks prior to analysis of test specimen. | >600 mg/dL CSF | < 6.0 mg/dL CSF |
| MG | - – 7.0 mg/dL serum/plasma | - – 14.0 mg/dL | Saline | # X2 (MAX dilution) | 2+ lipemia: Run ultracentrifuged specimen. | >14.0 mg/dL | < 0.1 mg/dL |
| NA | 100 – 200 mmol/L Serum/plasma | 100-200 mmol/L serum and plasma | N/A | DO NOT DILUTE serum/plasma | 3+ lipemia or Lipemia Index >10: Run ultracentrifuged specimen. | >200 mmol/L serum/ plasma | < 100 mmol/L serum and plasma |
| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| PHY | 2.5 – 40 µg/mL | 2.5- 80 µg/mL | Saline | X2 (MAX dilution) | Reanalyze specimen reported out as SUPPRESSED due to RXN ERROR | >80 µg/mL | < 2.5 µg/mL |
| PHS | 1.0– 12.0 mg/dL Serum/plasma | 1.0 – 24.0 mg/dL | Saline | X2 (MAX dilution) | 3+ lipemia or Lipemia Index >10: Run ultracentrifuge specimen. Interference may occur with specimens from patients with abnormal immunoglobulin synthesis (e.g. multiple myeloma). Some of these specimens may precipitate out when mixed with reagent, causing “rxn noise” errors. In an event such as this: 1) Dilute specimen 1:2 with saline and reanalyze; 2) If “rxn noise” still occurs, dilute specimen 1:2 with 12% TCA (Trichloroacetic acid). Centrifuge and analyze the supernatant. | >24.0 mg/dL | < 1.0 mg/dL |
| SALY | 4 – 100 mg/dL | 4-200 mg/dL | Saline | X2 (MAX dilution) | 2+ lipemia: Run ultracentrifuged specimen. | >200 mg/dL | <4 mg/dL |
| T BIL | 0.1 – 30 mg/dL | 0.1 – 60 mg/dL | Azide free human serum albumin | X2 (MAX dilution) | 2+ lipemia: Run ultracentrifuged specimen. | >60mg/dL | < 0.1 mg/dL |
| TP | 3.0 – 12.0 g/dL Serum/plasma NO CSFs | 3.0- 24.0 g/dL | Saline | X2 (MAX dilution) |  | >24.0 g/dL | <3.0 g/dL |
| Test | Analytical Measurement Range (AMR) | Reportable Range | Acceptable Diluent | # Maximum Manual # Dilutions | Special Notes | Upper Limit report as | Lower Limit report as |
| TTCA | 0-800 ng/mL Serum/plasma | 0-800 ng/mL | N/A | DO NOT DILUTE | Semi-quantitative test reported as: 0-300 ng/mL 300-500 ng/mL 500-700 ng/ml 700-800 ng/mL >800 ng/mL | >800 ng/mL | Reported as 0-300 |
| URIC | 0.5 – 21.0 mg/dL Serum/plasma Straight (0.5-12.0) Ordac (9.0 – 21.0) | 0.5 – 42 mg/dL | Saline | X2 (MAX dilution) |  | >42 mg/dL | < 0.5 mg/dL |
| VANC | 3.5 – 60 ug/mL Straight (3.5 – 40) Ordac (30 – 60) | 3.5 – 120 ug/mL | See Manual Dilution Column | >60 ug/mL specimens: Dilute with saline X2 (MAX dilution) <0.1 ug/mL specimens: Dilute 1:2 with Synchron 2 & if value X2 does not match value of known sample, send to PAML, if does match, report <3.5 |  | >120 ug/mL | <3.5 ug/mL |

| DOCUMENT APPROVAL Purpose of Document / Reason for Change: |  |  |  |
| --- | --- | --- | --- |
| 11/8/12 – New header/format. Modifications made to ACTM, AMY7, GLUC, added LACT. Modified M-TP, PHY and VANC requirements. |  |  |  |
| No significant change to process in above revision. Per CAP, this revision does not require further Medical Director approval. |  |  |  |
| Committee Approval Date | Date: N/A – revision of department-specific document which is used at only one facility | Medical Director Approval (Electronic Signature) | 12/7/12 |