|  | WORK INSTRUCTION |
| --- | --- |
|  | M-W-CG-1172-00 |
| sta coag connexion db managEment |  |
| St. Joseph Medical Center Tacoma, WA St. Clare Hospital Lakewood, WA St. Elizabeth Hospital Enumclaw, WA St. Francis Hospital Federal Way, WA St. Anthony Hospital Gig Harbor, WA PSC |  |

PURPOSE

To provide instruction for maintaining the Connexion database for the STA Coag system.

BACKGROUND

Maintaining a computer's database is vital in keeping a computer running smoothly. End of Day Maintenance, to be performed nightly, just after midnight, consists of a Database Purge and a Sample ID Reset. The Database Purge deletes all samples from the All Complete status, along with Incidents, Spy Files, and samples in Pending Status older than one day (T – 1). A Sample ID Reset moves samples from Complete status to All Complete. Performing a Database Purge and SID Reset on a routine basis allows the database to run more efficiently. Maintenance is performed in Dictionary Mode; therefore patients cannot be running during the process. This daily procedure should take around 5 minutes or less.

STEPS

- Check the WCP (31 – SFH, 310 – SJMC, 3100 – SCH) pending prior to starting the maintenance and make sure there are no Stats on the pending list. Also, make sure the Stagos are idle, with no patients or QC currently running.

- Close out of Run Mode by selecting the “Review & Edit” screen from the Windows Taskbar (Pencil icon on the bottom line of the screen). Click on the Review & Edit tab to open the window, and then click on the Red “X” in the upper-right corner of the window to close it (do not close by right-clicking on the tab and selecting “Close”).

- You will be notified that closing out of Run mode will close communications with the analyzer and the LIS. Press “Yes” to continue.

- The windows should all start closing down and you should return to the Start Menu (looks like a stop light with three red buttons – Run, Dictionary, and Other Utilities).

- If the Start Menu does not appear within a minute or so, you will need to log-off of the computer and log back on. Do this by clicking the Windows Start button on the bottom left corner of the screen and click on Log-off User. Press Ctrl-Alt-Del keys to log back on. The password is “TDstago2010”. Click on OK until the Start Menu appears.  See how I did this section.  One major step with a bullet to explain more.

- From the Start Menu, double-click the Dictionary mode button (middle of the stop light). Once the application starts in Dictionary mode, select the Tools menu by clicking the tab on the Windows                Taskbar (bottom line).

- From the Tools menu, double-click on the End of Day option.

- In the Purge box in the bottom left of the End of Day window, check the selection boxes for Database, Incidents, and Spy Files.

- In the text box between “All Complete    4            to be purged”, enter the number from the box above it, “Total Samples”.

- Click the “OK” button in the upper-right corner of the window. A message box will inform you that Pending samples will also be purged. Click “OK” to confirm and continue.

- While the purge is executing, the “OK” button will be grayed out and the “Stop” button will be active. You will see the number of samples to be purged being updated in real time. Once the purge is complete, the “OK” button will be active again. A message box will appear giving the number of All Complete samples that have been deleted. Click “OK” to continue.

- Once the Database Purge has completed, check the selection box for SID Reset in the upper-left corner of the End of Day window. Make sure no other options are selected.

- Click the “OK” button in the upper-right corner of the End of Day window to execute the reset. The “OK” button will be grayed out and the “Stop” button will be active during the reset process. Once the SID Reset is complete, the “OK” button will be active again.

- Click on “Exit” in the far upper-left of the End of Day window (or on the Red “X” in the upper-right corner) to return to the Sample Control Panel.

- To close out of Dictionary mode, right-click on the Review & Edit tab located on the Windows Taskbar (Pencil icon on the bottom line of the screen). Choose the “Close” option.

- When the Start Menu appears, double-click the Run mode button (top of the stop light) to start the Coag ConneXion in Run mode and re-initialize the analyzer and LIS interfaces.

- New samples on the pending should now cross from Cerner to Coag ConneXion. The instrument is now ready to run.

| DOCUMENT APPROVAL Purpose of Document / Reason for Change: |  |  |  |
| --- | --- | --- | --- |
| 9/27/12 – New document |  |  |  |
| Committee Approval Date | Date: 9/27/12 NA – revision of department-specific document which is used at only one facility | Medical Director Approval (Electronic Signature) | 9/26/12 |