PURPOSE

To provide instruction for managing CBC specimens with marked lipemia.

BACKGROUND

Lipemic plasma may cause turbidity within the hemoglobin chamber and / or incomplete red cell lysis.  This may result in a falsely elevated HGB result and invalid RBC indices.

SPECIMEN

EDTA  whole blood sample

EQUIPMENT / SUPPLIES

Isoton diluent and / or 0.9% saline

Plastic Aliquot tube with cap

Transfer Pipettes

Centrifuge

STEPS

1.	Using a transfer pipette, remove a well-mixed aliquot of the original specimen.

Pipette the aliquot into the plastic tube and cap.

2.	Centrifuge the aliquot for 10 minutes at 3400 rpm.

3.	Remove the specimen from the centrifuge.  Mark a line on the exterior of the tube

indicating the top of the plasma level.

4.	Remove as much plasma as possible using a transfer pipette without disturbing

the buffy-coat layer.

5.	Add diluent up to the line marked on the exterior of the tube.

6.	Mix the aliquot well and cycle the specimen on the analyzer.

7.	Evaluate the printout.  Use the RBC result as a guide for assessment of proper

re-dilution of the specimen.

8.	If RBC results correlate with the original and HGB/HCT ratio is acceptable, the

RBC, HGB, HCT, MCV, MCH, MCHC AND RDW from the saline-replaced aliquot

may be reported.

9.	Evaluate if the WBC and PLT results correlate with the original results.  Results

might be falsely decreased if the buffy-coat layer has been disturbed.  If results

are in question, perform a slide estimate and correlate with results from the

original printout.  Report the WBC, PLT, and MPV.

10.	Review reportable results and determine if manual differential is indicated.

REFERENCE

Hematology Procedures for Abnormal Bloods, Beckman-Coulter Manual, 11, pp. 4.23-4.24.