## Principle

## Principles of the Procedure

The Access Cortisol assay is a competitive binding immunoenzymatic assay. A sample is added to a reaction vessel with rabbit antibody to cortisol, cortisol-alkaline phosphatase conjugate, and paramagnetic particles coated with goat anti-rabbit capture antibody. Cortisol in the sample competes with the cortisol-alkaline phosphatase conjugate for binding sites on a limited amount of specific anti-cortisol antibody. Resulting antigen: antibody complexes bind to the capture antibody on the solid phase. After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is inversely proportional to the concentration of cortisol in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

## Summary and Explanation

Cortisol is the major glucocorticoid produced and secreted by the adrenal cortex. It affects (a) the metabolism of protein, fat, and carbohydrates, (b) the maintenance of muscle and myocardial integrity, and (c) the suppression of inflammatory and allergic activities. (1)

Corticosteroid-binding globulin and albumin bind approximately 90% of the cortisol secreted by the adrenal cortex. Bound cortisol circulates in an available, but temporarily inactive state. The physiological activity of cortisol depends upon levels of the small fraction of circulating unbound cortisol. (2)

Adrenocorticotropic hormone (ACTH) activates the synthesis and release of cortisol from the adrenal cortex. The pituitary produces and releases ACTH upon stimulation by corticotropin-releasing hormone (CRH) secreted from the hypothalamus. Unbound cortisol acts in a negative feedback mechanism upon the hypothalamus-pituitary-adrenal cortex (HPA) axis at the levels of the pituitary and hypothalamus. (3,4) In addition, diurnal variation and stress such as pyrogenically induced fever, severe psychosis, or trauma influence the neuroendocrine regulation of the adrenal cortex. (1,3,5)

Abnormal changes in cortisol levels occur due to hypothalamic, pituitary, or adrenal malfunction. If undiagnosed and untreated, these disorders can lead to severe metabolic imbalance which may be life-threatening. The measurement of serum or plasma cortisol - utilizing morning and evening levels and/or stress tests such as ACTH stimulation or dexamethasone suppression - aids in the diagnosis of adrenal related disease. Excess cortisol levels are found in Cushing’s syndrome (adrenal cortical hyperfunction) while decreased levels are found in Addison’s Disease (adrenal cortical insufficiency). (6)

Cortisol bound to protein is protected from metabolism by the liver. (4) Unbound (or free) cortisol in serum is metabolized by the liver resulting in a wide variety of forms or metabolites. Many of these metabolites (conjugated, glucuronide and sulfate forms) are water–soluble and rapidly voided in the urine. A small amount (< 100 g/24 hours) of cortisol and other extractable metabolites are also excreted in the urine. Urine cortisol can be measured by performing an extraction step which removes some of the water–soluble

metabolites prior to analysis or by analyzing urine directly. Immunoassays measure urine cortisol as well as some immunoactive metabolites, therefore, representative expected values should be established for each immunoassay method. The measurement of urine cortisol reflects the amount of unbound (or free) serum cortisol and aids in the diagnosis of adrenal hyperactivity. An elevated level of urine cortisol is considered diagnostic for Cushing’s syndrome (adrenal cortical hyperfunction). (4,7,8,9,10)

## Specimen Collection

- Serum and plasma (heparin, EDTA) are the recommended samples.

- Observe the following recommendations for handling, processing, and storing blood samples: (11)

- Collect all blood samples observing routine precautions for venipuncture.

- Allow serum samples to clot completely before centrifugation.

- Keep tubes stoppered at all times.

- Within two hours after centrifugation, transfer at least 500 µL of cell-free sample to a storage tube. Tightly stopper the tube immediately.

- Store samples tightly stoppered at room temperature (15 to 30°C) for no longer than

- eight hours.

- If the assay will not be completed within eight hours, refrigerate the samples at 2 to 8°C.

- If the assay will not be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

- Thaw samples only once.

- Use the following guidelines when preparing specimens:

- 1.	Ensure residual fibrin and cellular matter has been removed prior to analysis.

  - Follow blood collection tube manufacturer’s recommendations for centrifugation.

- Each laboratory should determine the acceptability of its own blood collection tubes and serum separation products. Variations in these products may exist between manufacturers and, at times, from lot-to-lot.

Criteria for sample rejection as designated by this laboratory:*

See also Specimen Rejection Protocol

Special instructions for specimen handling as designated by this laboratory:*

Assemble supplies appropriate to perform test and don PPE as outlined in the FHS Safety Manual, Infection Control Plan and Policy.

## Reagents and Equipment

- R1: Access Cortisol Reagent Pack 
Cat. No. 33600: 100 determinations, 2 packs, 50 tests/pack.

Provided ready to use. Store upright and refrigerate at 2 to 10°C. Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument. Stable until the expiration date stated on the label when stored at 2 to 10°C. Stable at 2 to 10°C for 14 days after initial use. Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range. If the reagent pack is damaged (i.e., broken elastomer), discard the pack. All antisera are polyclonal unless otherwise indicated.

1.	R1a:	Cortisol-alkaline phosphatase (bovine) conjugate and paramagnetic particles coated with goat anti-rabbit IgG in TRIS buffered saline, with surfactant, BSA matrix, and < 0.1% sodium azide.

2.	R1b:	Rabbit antiserum to cortisol in TRIS buffered saline, with surfactant, BSA matrix, and 
< 0.1% sodium azide.

B.	Access Cortisol Calibrators
Cat. No. 33605: S0-S5, 4.0 mL/vial

Quantitative assay calibration is the process by which samples with known analyte concentrations (i.e. assay calibrators) are tested like patient samples to measure the response. The mathematical relationship between the measured responses and the known analyte concentrations establishes the calibration curve. This mathematical relationship, or calibration curve, is used to convert RLU (Relative Light Unit) measurements of patient samples to specific quantitative analyte concentrations.

Provided ready to use. Store upright and refrigerate at 2 to 10°C. Mix contents by gently inverting before use. Avoid bubble formation. Stable until the expiration date stated on the label when stored at 2 to 10°C. Signs of possible deterioration are control values out of range. Refer to calibration card and or vial labels for exact concentrations.

Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

- S0: Human serum with < 0.1% sodium azide, and 0.5% ProClin** 300. Contains 0 µg/dL (nmol/L) cortisol.

- S1–S5: Cortisol (purified chemical compound) in human serum at levels of approximately 2, 5, 10, 25 and 60 µg/dL (55, 138, 276, 690 and 1655 nmol/L), respectively, with < 0.1% sodium azide, and 0.5% ProClin 300.

- Calibration Card: 1

C.	Access Substrate
Cat. No. 81906: 4 x 130 mL

Provided ready to use. Refer to the following chart for storage conditions and stability. An increase in substrate background measurements may indicate instability.

| Condition | Storage | Stability |
| --- | --- | --- |
| Unopened | 2 to 8°C | Until expiration date stated on the label |
| Equilibration prior to use (unopened) | 15 to 30°C (room temperature) | Minimum 18 hours Maximum 14 days |
| In use (opened) | Internal substrate supply position | Maximum 5 days |
| In use (opened) | External fluids tray substrate position | Maximum 14 days |

R2 Substrate: Lumi-Phos 530 (buffered solution containing dioxetane Lumigen* PPD, fluorescer, and surfactant).

Refer to the appropriate system manuals and/or Help system for detailed instructions.

D.	Access®, Access 2, SYNCHRON LX®i:
Access Wash Buffer II, Cat. No. A16792

UniCel® DxI:

Unicel DxI: Wash Buffer II, Cat. No. A16793

Provided ready to use. Stable until the expiration date stated on the label when stored at room temperature (15 to 30°C). An increase in substrate background measurements or increased relative light units for the zero calibrators in “sandwich”-type assays may indicate instability.

R3 Wash Buffer: TRIS buffered saline, surfactant, < 0.1 sodium azide, and 0.1% ProClin 300.

Refer to the appropriate system manuals and/or Help system for detailed instructions.

E.	Quality Control (QC) materials: commercial control material

F.	Access Immunoassay System and supplies

G.	Warnings and Precautions

- For in vitro diagnostic use.

- Patient samples and blood-derived products may be routinely processed with minimum risk using the procedure described. However, handle these products as potentially infectious according to universal precautions and good clinical laboratory practices, regardless of their origin, treatment, or prior certification. Use an appropriate disinfectant for decontamination. Store and dispose of these materials and their containers in accordance with local regulations and guidelines.

- Sodium azide may react with lead and copper plumbing to form highly explosive metal azides. On disposal of liquids, flush with a large volume of water to prevent azide build-up.(12)

- Human source material used in the preparation of the reagent has been tested and found negative or non-reactive for Hepatitis B, Hepatitis C (HCV), and Human Immunodeficiency Virus (HIV-1 and HIV-2). Because no known test method can offer complete assurance that infectious agents are absent, handle reagents and patient samples as if capable of transmitting infectious disease. (13)

- Each serum/plasma pool used in the preparation of this product has been tested and found negative for the presence of fibrinogen.

- Xi. Irritant: 0.5% ProClin 300.
R 43: May cause sensitization by skin contact.
S 28-37: After contact with skin, wash immediately with plenty of soap and water. Wear suitable gloves.

- Substrate is sensitive to air exposure. Keep tightly closed at all times. Do not pool bottles of substrate.

- The Material Safety Data Sheet (MSDS) is available upon request.

## Calibration Details

An active calibration curve is required for all tests. For the Access Cortisol assay, calibration is required every 28 days. Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

The Access Cortisol Calibrators are provided at six levels - zero and approximately 2, 5, 10, 25, and 60 µg/dL - prepared gravimetrically from purified cortisol and human serum. Assay calibration data are valid up to 28 days.

Calibrators run in duplicate.

## Quality Control

Quality control materials simulate the characteristics of patient samples and are essential for monitoring the system performance of immunochemical assays. Because samples can be processed at any time in a “random access” format rather than a “batch” format, quality control materials should be included in each 24-hour time period.(14) Include commercially available quality control materials that cover at least two levels of analyte. Follow manufacturer's instructions for reconstitution and storage. Each laboratory should establish mean values and acceptable ranges to assure proper performance. Quality control results that do not fall within acceptable ranges may indicate invalid test results. Examine all test results generated since obtaining the last acceptable quality control test point for this analyte. Refer to the appropriate system manuals and/or Help system for information about reviewing quality control results. Refer to Chemistry Controls Forms and Quality Control policies. *

## Procedure

A.	Access Instrument

Refer to the appropriate system manuals and/or Help system for preparation and operation.

B.	Assay Procedure

- Refer to the appropriate system manuals and/or Help system for information on managing samples, configuring tests, requesting tests, and reviewing test results.

## Results

- Patient test results are determined automatically by the system software using a weighted four parameter logistic curve (4PLC) math model. The amount of analyte in the sample is determined from the measured light production by means of the stored calibration data. Patient test results can be reviewed using the appropriate screen. Refer to the appropriate system manuals and/or Help system for complete instructions on reviewing sample results.

B.	Expected Values

- Each laboratory should establish its own reference ranges to assure proper representation of specific populations.

- Single serum AM cortisol concentrations were measured in 130 human serum samples from apparently healthy adult male and female subjects using the Access Cortisol assay. The 95% confidence range of single serum AM cortisol concentrations was 6.7–22.6 µg/dL 
(185–624 nmol/L).

- Single serum PM Cortisol concentrations reported in the literature are generally < 10 µg/dL 
(< 276 nmol/L).(10)

- Serum Cortisol expected values in children over the age of 6 are considered to be equivalent to the adult population.(10)

## Reference ranges*

| Cortisol AM | Cortisol PM | Normal Post Stimulation |
| --- | --- | --- |
| 4.3 – 22.4 ug/dL | 3.0 – 16.0 ug/dL | >20.0 ug/dL |

## Procedural Comments

A.	Refer to the appropriate system manuals and/or Help system for a specific description of installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

B.	Mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument. Do not invert open (punctured) packs.

C.	Use (twenty-five) 25 µL of sample for each determination in addition to the sample container and system dead volumes. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

D.	The system default unit of measure for sample results is µg/dL. To change sample reporting units to the International System of Units (SI units), nmol/L, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply µg/dL by multiplication factor 27.59.

## Limitations of the Procedure

A.	Samples can be accurately measured within the analytic range of the lower limit of detection and the highest calibrator value (approximately 0.4–60.0 µg/dL [11–1655 nmol/L]).

- If a sample contains less than the lower limit of detection for the assay, report the results as less than that value (i.e. < 0.4 µg/dL [< 11 nmol/L]).

- If a sample contains more than the stated value of the highest Access Cortisol Calibrator (S5), report the result as greater than that value (i.e. > 60.0 µg/dL [> 1655 nmol/L]). Alternatively, dilute one volume of sample with one volume of Access Cortisol Calibrator S0 (zero) which is also available as Access Cortisol Calibrator S0 cat. no. 33606. Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution.

ReportING RESULTS OUTSIDE OF ANALYTICAL RANGE*

| Lower limit of detection | 0.4 ug/dL | Results below 0.4, report as <0.4 ug/dL |
| --- | --- | --- |
| Upper limit of detection | 60.0 ug/dL | Results >60.0 should be diluted, starting at X2, with Cortisol Calibrator S0 (zero), and reanalyzed. |

B.	For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce antibodies, e.g. HAMA, that interfere with immunoassays. Additionally, other heterophile antibodies such as human anti-goat antibodies may be present in patient samples. (15,16)

Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

C.	The Access Cortisol results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests and other appropriate information.

- Samples containing up to 10 mg/dL (171 µmol/L) bilirubin, hemolyzed samples containing up to 500 mg/dL hemoglobin, and lipemic samples containing the equivalent of 1800 mg/dL (20.32 mmol/L) triglycerides do not affect the concentration of cortisol assayed. In addition, samples containing up to 9 g/dL (90 g/L) total protein [normal sample + 3 mg/dL (0.03 g/L) albumin] do not affect the concentration of cortisol assayed.

- The following table describes the cross-reactivity of the assay with substances that are similar in structure to cortisol.

| Substance | Analyte Added (g/dL) | Cross-Reactivity (%) |
| --- | --- | --- |
| Corticosterone | 100 | 2.08 |
| Cortisone | 100 | 8.06 |
| 11-Deoxycorticosterone | 1000 | 0.91 |
| 11-Deoxycortisol | 100 | 17.80 |
| 17- Hydroxyprogesterone | 1000 | 5.33 |
| Progesterone | 1000 | 0.46 |
| Prednisolone | 20 | 7.60 |
| Tetrahydrocortisone | 1000 | 0.10 |
| Prednisone | 1000 | 3.05 |
| Dexamethasone | 1000 | 0.04 |

- The lowest detectable level of cortisol distinguishable from zero (Access Cortisol Calibrator S0) with 95% confidence is 0.4 µg/dL (11 nmol/L).

- Elevated cortisol levels may occur in patients receiving prednisolone or prednisone (which is converted to prednisolone in vivo) due to cross-reactivity to prednisolone.

Access, SYNCHRON LX, UniCel, DxI and the Beckman Coulter logo are trademarks of  
Beckman Coulter, Inc.

*Lumi-Phos and Lumigen are trademarks of Lumigen, Inc.

**ProClin is a trademark of Rohm and Haas Company, or its subsidiaries or affiliates.

## References

Beckman Coulter, Inc. Access Cortisol product insert, Fullerton, CA 92835, A33262.

Beckman Coulter, Inc. Access Substrate product insert, Fullerton, CA 92835, 386966.

Beckman Coulter, Inc. Access Wash Buffer II product insert, Fullerton, CA 92835, A16534 (Access), A16543 (UniCel DxI).

- Bondy PK. The adrenal cortex. In Metabolic control and disease, eighth edition, 1980; 1427-1499. Edited by Bondy, PK, Rosenberg LE, Philadelphia, PA: WB Saunders Co.

- Beisel WR, DiRaimondo VC, Forsham PH. Cortisol transport and disappearance. Annals of Internal Medicine, 1964; 60: 641-652.

- Travis JC (ed). Plasma cortisol. Rx: RIA for physicians, 1976; 1(8).

- Labhart F. Clinical Endocrinology, 1974; 290-294. New York: Springer-Verlag.

- Liddle GW, Melmon KL. The adrenals. In Textbook of endocrinology, fifth edition,1974; 233-283. Edited by Williams RH, Philadelphia, PA: WB Saunders Co., Philadelphia.

- Fujimoto WY. Disorders of glucocorticoid homeostasis. In Blue book of endocrinology, X edition, 1985; 43-59. Edited by Metz R, Larson E, Philadelphia, PA: WB Saunders Co.

- Murphy BEP, Okouneff LM, Klein GP, Ngo SK. Lack of specificity of cortisol determinations in human urine. Journal of Clinical Endocrin Metab, 1981; V53: 91-99.

- Schöneshöfer M, Fenner A, Dulce HJ. Interferences in the radioimmunological determination of urinary free cortisol. Clin Chem Acta,1980; V101: 125-134.

- Gough RM, Ellis G. The radioimmunoassay of cortisol in urine. Difficulties experienced in the development of an assay and problems of specificity observed with commercial reagents supplied as kits. Clin Biochem, 1981; V14 (2): 74-81.

- Moore A, Raitken R, Burke C, Gaskell S, Groom G, Holder G, Selby C, Wood P. Cortisol assays: guidelines for the provision of a clinical biochemistry service. Annals of Clinical Biochemistry, 1985; V22: 435-454.

- Approved Guideline – Procedures for the handling and processing of blood specimens, H18-A2. 1999. National Committee for Clinical Laboratory Standards.

- DHHS (NIOSH) Publication No. 78-127, August 1976. Current Intelligence Bulletin 13 - Explosive Azide Hazard. Available http://www.cdc.gov/niosh.

- HHS Publication, 4th ed., May 1999. Biosafety in Microbiological and Biomedical Laboratories. Available http://www.cdc.gov/od/ohs/biosfty/bmbl4/bmbl4toc.htm

- Cembrowski GS, Carey RN. Laboratory quality management: QC & QA. ASCP Press, Chicago, IL, 1989.

- Kricka, L. Interferences in Immunoassays – Still a Threat. Clin Chem 2000; 46: 1037.

- Bjerner J, et al. Immunometric Assay Interference: Incidence and Prevention. Clin Chem 2003; 48: 613–621.

* Indicates specific to FHS laboratories

Beckman Coulter, Inc. does not automatically distribute revised NCCLS procedures. If you receive a revised copy of the assay insert, call Technical Support to determine if the NCCLS procedure has also been revised.

Printed in U.S.A.