WINCHSTER HOSPITAL

Winchester, Massachusetts

# Section:	Centrifugation of Blood Specimens

Title:		Centrifuge Operation

Purpose:	To insure proper handling of specimens and correct operation of the centrifuge in order to maintain specimen integrity

Instrument:   Silencer 2110SP Centrifuge

Specimens:    Greiner plastic tubes

Note: Winchester Hospital Laboratory purchases Greiner plastic tubes.  On rare occasions, specimens are received in glass tubes from customers.

Serum tubes should be allowed to clot for a minimum of 10 minutes but not more than 2 hours before centrifugation.  All other tubes can be centrifuged immediately.

Set up:

- The centrifuge should be operated on a stable, level bench top.

- The environment should be between 50 degrees to 86 degrees F, with a relative humidity of 60%.  There should be 6 inch on all sides of the centrifuge.

- The lid should be left open between cycles to prevent overheating of samples.

- Power – Press once to bring the unit out of sleep mode and turn the displays on. Pressing the power switch while the power is on turns the displays off and puts the unit into a sleep mode.

- Speed – Use arrows to set speed. Speed should always be set at 4500 rpm to start a run. Speed display will show actual speed during the run cycle. Speed display will also show “HELP” when an error condition occurs. (See manual for any error codes).

Note: Winchester Hospital uses Greiner plastic tubes.  If any glass tubes are received from an outside source the speed must be adjusted to 3400 rpm.

- Time – Time should be set at 5 minutes.

- Start switch – Also doubles as a Stop switch. If pressed when the unit is in standby, it will start the run cycle. Pressing the start switch during a cycle stops the cycle.

- Lid Interlock Indicator– An LED light that is off when the lid is closed and locked and on when the lid is either open or unlocked.

Procedure:

- Load samples, taking care to maintain a balanced load.  Standard size tubes should be spun in the stainless steel inserts.  The plastic inserts should be used for pediatric tubes.

- Verify that speed and time are set correctly.

- Close lid and press the Start switch. The lidlock will engage and the centrifuge will accelerate to the set speed. Upon reaching the set speed the time will start counting down and the Time display will show the time remaining.

- When the rotor has come to a complete stop, the lidlock will release and the LED will be illuminated.

- For emergency lidlock release, only the supplied release tool can be used.  Unplug the centrifuge before performing this function.

- Please see the Silencer Instruction Manual for any operation or service issues.

Maintenance:

  - Weekly, the interior of the centrifuge should be wiped out with a bleach wipe.  This should be followed by a clear water rinse with a damp cloth.  The interior must be thoroughly dried before using the centrifuge.  Liquids should be used sparingly as the motor seal is not watertight.   Never use an aerosol or spry cleaner on the centrifuge.  The cleaning should be documented on the maintenance chart.

References:

Greiner bio-one Package Insert

Silencer 2110SP Service manual

Winchester Hospital Silencer Validation Study 6/2013

Procedure Adopted: March 1, 2012

Procedure Written By: Nancy Flaherty_________________________________

Procedure Reviewed and Approved By: ________________________________

Annual Review:

Reviewed By: _____________________________________________________

Reviewed Bt: ______________________________________________________

Reviewed By:______________________________________________________

Reviewed By:______________________________________________________

Reviewed By:______________________________________________________