Clinical Significance:

The OSOM hCG Urine Test is a rapid immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine for the early determination of pregnancy.  Human chorionic gonadotropin (hCG) is a glycoprotein hormone secreted by the developing placenta.  After fertilization, the concentration of hCG rapidly rises in the urine of pregnant women.

The detection of hCG in urine is an excellent marker for confirming pregnancy.   The OSOM hCG Urine Test is a rapitest, which can detect the presence of hCG in urine.  The test utilizes monoclonal antibodies specific to hCG.

POLICY:

OSOM hCG Urine Test is a solid phase, sandwich-format immunochromatographic assay for the qualitative detection of hCG.  Urine is added to the sample well of the Test Device using the fixed volume pipette provided.  The       sample migrates through reaction pads where hCG, if present in the sample, binds to a monoclonal anti-hCG dye conjugate.  The sample then migrates across a membrane towards the results window, where the labeled monoclonal antibody-hCG complex is captured at a test line region containing immobilized monoclonal anti-hCG.

Excess conjugate will flow past the test line region and be captured at a control line region containing an immobilized antibody directed against the anti-hCG dye conjugate with or without hCG complexed to it.  The appearance of 2 black bands in the results window - one at "T: Test" and the other at "C: Control" - indicates the presence of hCG in the sample.  If a detectable level of hCG is not present, only the control band will appear in the result window.

Collection of Specimen:

No filtration or centrifugation of urine specimen is required for testing with the OSOM hCG Urine Test.

Urine

Urine specimens may be collected in any clean, dry, plastic or glass container, for early determination of pregnancy.  The first morning specimen of urine is recommended since it usually contains the highest concentration of hCG.  Urine specimens may be stored at room temperature (15 to 30°C) for up to 8 hours, or refrigerated at 2 to soc for up to 72 hours.

Reagents and Materials:

- OSOM hCG Urine (Catalog Number 124) (25 tests)

- Clock or timer

- Sample collection cups or tubes

- Positive and Negative OSOM hCG Urine Control Set (Cat 3M43-l)

Storage:

New room temperature storage conditions, 15 to 30°C (59 to 86°F)

Store OSOM® hCG Urine Tests at room temperature, 15 to 30°C (59 to 86°F), out of direct sunlight.  Test devices are stable until the expiration date printed on the kit or foil pouch.  DO NOT FREEZE.  If the control band does not appear, the Test may have been stored or handled improperly or the foil pouch may not be intact.

Urine Control

Positive and Negative hCG Urine Control Set is ready to use and must be stored refrigerated at 2-8 °C until the expiration date on the label.

Procedural Notes:

- If specimen has been stored refrigerated, allow it to warm to room temperature before use.

- Several tests can be run at the same time.  Use a new AccuPipette with each test to avoid contamination errors.

Test Procedure:

- Remove the Test Device from the pouch.  Place the Device on a flat surface.

- Squeeze the bulb of the AccuPipette and drawup enough sample to completely fill the barrel.

- Expel the entire contents of the barrel (135ul) into the sample well of the Test Device. There is no drop counting is required.  Some samples may remain in the reservoir. Discard the AccuPipette in a suitable biohazardous waste container.

- Read the results at 3 minutes for urine.  Strong positive results can be observed sooner.  If no band appears at "T" the test can be considered negative.   Results are invalid after the stated read time.  The use of a timer is recommended.

- Any positive patient ≤ 18 years of age send hCG to main lab for testing.  Do not report results

Interpretation of Results:

- Positive

  - Two distinct black or gray bands- one at "T: Test" and the other at "C: Control" – are visible in the results window, indicating that the specimen contains detectable levels of hCG.   While the intensity of the test band may vary with different specimens, the appearance of 2 distinct bands should be interpreted as positive results.  All positive results must be sent to the Main Laboratory for Quantitative hCG Determination.

- Negative:

  - Only a black or gray band at the "C: Control" position.  Indicating that a detectable level of hCG is not present.

- Invalid:

  - If no band appears at the "C: Control" or incomplete or beaded bands appear at either the "T: Test" or "C: Control" the test is invalid.  The test should be repeated using another OSOM hCG Urine Test Device.

  - The test is valid if the control line appears by the stated read time, regardless of whether the sample has migrated all the way to the end of the sample window.

Warnings and Precautions:

- For in vitro diagnostic use.

- Do not use beyond the expiration date printed on the kit or foil pouch.

- Use appropriate precautions for the collection, handling and storage of specimens.  All human blood products should be treated as potentially infectious and handled with good laboratory practices using appropriate precautions recommended in the Center for Disease Control/National Institutes of Health Manual, Bio-safety in Microbiological and Biomedical Laboratories, AccuPipetes and specimens in suitable biohazardous waste containers.

- Test Devices are stable in the unopened foil pouches until the expiration date.  Do not remove the Test Device from the pouch until needed.

Quality Control:

- Internal Quality Control

  - Several controls are incorporated into each OSOM hCG. It is recommended that these procedural controls be documented for each sample as part of daily quality control.

  - The same-labeled conjugate antibody results in the appearance of both the test and the control bands.  The appearance of the control band in the results window is an internal positive procedural control, which validates the following:

    - Test System:

      - The appearance of the control band assures that the detection component of both the test line and control fine is intact, that adequate volume was added and that adequate capillary migration of the sample has occurred. It also verifies proper assembly of the Test Device.

    - Operator:

      - The appearance of the control band indicates that an adequate volume of fluid was added to the sample well for capillary migration to occur.  If the control band does not appear at the read time, the test is invalid.  Appearance of the control band should be documented with a "V" for valid as part of the daily quality control.

      - The clearing of the background in the results area may be documented as a negative procedural control. It also serves as an additional capillary flow control.  At the read time, the background should appear white to light gray and not interfere with the reading of the test.  The test is invalid if the background fails to clear and obscures the formation of a distinct control band.

      - If the control band fails to appear with a repeat assay, do not report patient results.

      - Notify Supervisor and immediately send specimens to the Main Laboratory for processing.  Contact Genzyme Diagnostics for technical service: Tel 617-252-7760/800-332-1042 (U.S. Customers only) / Fax 617-252-7759.

- External Quality Control

  - It is recommended that external hCG controls are run for Urine with each new box of devices, but not more than a week, whichever comes first, each new lot and with each new untrained operator.

Limitations in hCG Testing:

- This assay is capable of detecting only whole molecule (intact) hCG. It cannot detect the presence of free hCG subunits.  Therefore, this test should only be used for the qualitative detection of human chorionic gonadotropin in urine for the early determination of pregnancy.

- For diagnostic purposes hCG test results should always be used in conjunction with other methods and in the context of the patient's clinical information (e.g. medical history, symptoms, results of other tests, clinical impression, etc.). Ectopic pregnancy cannot be distinguished from normal pregnancy by hCG measurements alone.

- If the hCG level is inconsistent with, or unsupported by, clinical evidence, results should also be confirmed by an alternative hCG method.

- Interfering substances may falsely depress or falsely elevate results.  These interfering substances may cause false results over the entire range of the assay, not just at low levels and may indicate the presence of hCG when there is none. As with any immunochemical reaction, unknown interferences from medications or endogenous substances may affect results.

- Infrequently, hCG levels may appear consistently elevated and could be due to, but not limited to, the presence of the following:

  - heterophilic antibodies: Patients routinely  exposed to animals or to animal serum products, cab be prone to this interference  and anomalous  values may be observed.

  - trophoblasitic or non-trophoblastic neoplasms:  abnormal  physiological states that may falsely elevate hCG.levels. This test should not be used in the diagnosis of these conditions.

  - Nonspecific protein binding

  - HCG like substances

- Specimens from patients who have received preparations of Mouse Monoclonal Antibodies for diagnosis or therapy may contain Human Anti-Mouse Antibodies (HAMA).   Such specimens may demonstrate either falsely elevated for falsely depressed results when tested with assay kits, which employ Mouse Monoclonal Antibodies.  These specimens should not be tested with the OSOM hCG Urine/Serum Test Device.

- Because of the high degree of sensitivity of the assay, specimens tested as positive during the initial days after conception may later be negative due to natural termination of the pregnancy.   Overall, natural termination occurs in 22% of clinically unrecognized pregnancies and 31% of other pregnancies.  In the presence of weakly positive results, it is good laboratory practice to sample and test again after 48 hours.

- If the test band appears very faint, it is recommended that a new sample be collected 48 hours later and tested using another OSOM hCG Urine/Serum Test Device.

- Dilute urine specimens may not have representative levels of hCG.

- Detection of very low levels of hCG does not rule out pregnancy as low levels of hCG can occur in apparently healthy non-pregnant subjects.   Additionally post menopausal specimens may elicit weak positive results due to low hCG levels unrelated to pregnancy.   In a normal pregnancy, hCG values double approximately every 48 hours.  Patients with very low levels of hCG should be sampled and tested again after 48 hours or tested with an alternative method.

Expected Values:

hCG is not normally detected in the urine specimens of healthy men and non-pregnant  women.  In normal pregnancy 20ml U/ml hCG is reported to be present in both urine and serum 2 to 3 days before the first missed menstrual period.  The levels of hCG continue to increase up to 200,000mlU/ml at the end of the first trimester.

Performance Characteristics:

- Sensitivity

  - The OSOM hCG kit does not normally detected hCG in urine specimens with concentrations lower than 20 mIU/ml.

    - The expected sensitivity of urine samples at a read time of 3 minutes is 20 mlU/ml

Note: Samples containing minute quantities of hCG (below lO mlU/ml) may develop very faint test bands.

Interfering Substances:

The following substances were added to urine and serum specimens containing 0 to 20 mlU/ml hCG.  The substances at the concentrations listed below were not formed to affect the performance of the test.

Cross Reactivity:

The addition of leutinizing hormone (500 m!U/ml of LH), follicle stimulating hormone (1OOO mlU/ml of FSH), or thyroid stimulating hormone (1OOO mlU/ml of TSH) to negative urine and serum specimens gives negative results in the OSOM Urine Test.

FORMS:

None

ATTACHMENTS:

None

CHECKLISTS/JOB AIDS:

None

REFERENCES:

OSOM hCG Urine Kit package insert

CHRONOLOGICAL HISTORY:

Adopted:  	02/21/2014

Reviewed: 	02/21/2014	JR

Revised:

PREPARED BY:  Juan Roitman