##### WINCHESTER HOSPITAL			LABORATORY / MICROBIOLOGY

##### ______________________________________________________________________

##### TITLE:  PBP2A CULTURE COLONY TEST (Alere)

POLICY: Micro.03.0037.01 ______________________________________________________________________

PRINCIPLE

This is a High Complexity test. The AlerePBP2a Culture Colony Test is a qualitative, in vitro immunochromatographic assay for the detection of penicillin-binding protein 2a (PBP2a) in isolates identified as Staphylococcus aureus, as an aid in detecting methicillin-resistant Staphylococcus aureus (MRSA). .

AlerePBP2a Culture Colony Test is a rapid, immunochromatographic membrane assay that uses highly sensitive monoclonal antibodies to detect PBP2a from bacterial isolates. These antibodies and a control antibody are immobilized onto a nitrocellulose membrane as two distinct lines and combined with a sample pad, a blue conjugate pad, and an absorption pad to form a test strip.

SPECIMEN COLLECTION/TREATMENT

| Specimen: | Bacterial isolates of Staphylococcus aureus. |
| --- | --- |
| Culture media: | Acceptable: Tryptone Soy Agar (Tryptic Soy Agar) with 5% sheep blood (TSA blood), Columbia Agar with 5% sheep blood, and Mueller Hinton Agar. Other culture media should be validated prior to use with the AlerePBP2a Culture Colony Test. |
| Specimen Storage: | The use of fresh (18 - 24 hours) cultures is recommended. |
| Handling Precautions: | Culture isolates, controls, and test strips should be handled as though they could transmit disease. Observe established precautions against microbial hazards. |

MATERIALS

Reagents and Materials Provided

| - Component | - Content | - Quantity |
| --- | --- | --- |
| - Foil pouch | - A test strip covered with a clear, plastic overlay with directional arrows to indicate testing position, assay tube, workstation and desiccant packet. | - 25 |
| - Reagent 1 | - A clear, blue alkaline solution (irritant) | - 1 |
| - Reagent 2 | - A clear, slightly acidic solution containing sodium azide, buffer and proteins | - 1 |

Reagents and Materials not provided

- Clock, timer or stopwatch

- Applicator sticks

- Vortex mixer

- External positive and negative controls

STORAGE AND STABILITY

Store kit at (2-8ºC). Do not interchange or mix components from different kit lots.

The AlerePBP2a Culture Colony Test Kit and reagents are stable until the expiration dates marked on their outer packaging.

QUALITY CONTROL

###### Internal Procedural Controls

The AlerePBP2a Culture Colony Test has built-in positive and negative procedural controls.  Record these controls for each test run.

- The appearance of blue line at the ‘control line’ position can be considered an internal positive procedural control.  If capillary flow has occurred, this line will always appear.

- In comparison to the color of the control line, the background color on the test strip should be white within 5 minutes.

External Positive and Negative Controls:

Perform full Quality Control (QC) for each new lot or shipment, upon receipt using known stock organisms.  Results are recorded on the appropriate Quality Control sheet in the Quality Control binder.

Positive control: Staphylococcus aureus ATCC # 43300

Negative control: Staphylococcus aureus ATCC # 25923

LIMITATIONS, INTERFERANCES AND PRECAUTIONS

- For in vitro diagnostic use only.

- Allow all kit components to equilibrate to room temperature (15-30°C) before use.

- Leave test strip sealed in its foil pouch until just before use.

- Avoid skin and eye contact with reagents and test strip.

- Do not use kit past its expiration date.

- Do not interchange or mix components from different kit lots.

- As with all bacterial samples, controls and test strips may contain pathogenic organisms; handle with appropriate precautions and dispose of materials safely in biohazard waste receptacles.

- The AlereTM PBP2a Culture Colony Test should be performed only on isolates of Staphylococcus aureus species.  A coagulase or equivalent test must be performed to determine if the isolate is Staphylococcus aureus prior to testing with the AlereTM PBP2a Culture Colony Test.

- Reagent 1 is an irritant. Causes skin and eye irritation. Avoid contact with eyes, skin and mucus membranes.  Safety data sheet is available for professional user upon request.

- Reagent 2 contains sodium azide.

- There are no known cross reactants for the Aler PBP2a test, although other non-S. aureus samples that contain PBP2a may produce positive results (such as Staphylococcus epidermidis). That is why is it important to determine if S. aureus is present before running the Alere™ PBP2a test.

SPECIMEN IDENTIFICATION

All specimens are identified with the initial specimen label upon receipt.   Each test tube and cartridge is given a run number. This reference number is used on all tubes throughout the entire process and all phases of the assay.  At any point in the assay run the specimen can be positively identified.

PROCEDURE

Control Procedure:

Allow all reagents and test strips to equilibrate to room temperature (15-30°C) before testing.

- Subculture the control strain onto a blood agar plate. Incubate the plate overnight at 37°C ± 2°C for 18-24 hrs.

- Holding the dropper bottle vertically, add two drops of Reagent 1 to an assay tube.

- Lightly touch three well-isolated colonies with an applicator stick or inoculating needle, place into the  tube and thoroughly mix (Note: For testing from Mueller Hinton plates, three touches of the lawn will suffice).

- Holding the dropper bottle vertically, add two drops of Reagent 2 to the tube.

- Vortex briefly. The blue solution must turn a clear color (if the color does not change, add one more drop of Reagent 2 and mix until the sample turns clear).

- Insert the test strip into the assay tube with arrows pointed downward.

- At five minutes, withdraw the test strip from the tube and read the assay result.

Assay Procedure:

Allow reagents and test strips to equilibrate to room temperature (15-30°C) before testing.  The test can be performed from well-isolated colonies on the primary plate if there is sufficient growth, or from a subculture of the isolate.

- Holding the dropper bottle vertically, add two drops of Reagent 1 to an assay tube.

- Lightly touch three well-isolated colonies from the culture plate with an applicator stick or inoculating needle, place into the tube and thoroughly mix (Note: For testing from Mueller Hinton plates, three touches of the lawn will suffice).

- Holding the dropper bottle vertically, add two drops of Reagent 2 to the tube.

- Vortex briefly. The blue solution must turn a clear color (if the color does not change, add	one more drop of Reagent 2 and mix until the sample turns clear).

- Insert the test strip into the assay tube with arrows pointed downward.

- At five minutes, withdraw the test strip from the tube and read the assay result.

INTERPRETATION OF TEST RESULTS

Negative Sample:  A BLUE Control Line appears in the top half of the test strip.  No other line appears.

Positive Sample:   The BLUE Control Line appears AND a second BLUE Sample Line appears below it in the bottom half of the test strip.  Any Sample Line, even when very faint, is positive.  Note:  Strong positive results may cause a Control Line that is lighter in color than the Sample Line.  This is still a valid assay.  Any control line, no matter how faint in color, is considered valid.

Invalid Result:  The BLUE Control Line does not appear, whether a Sample Line is present or not. Repeat invalid tests with a new test strip.  Call Alere Technical Service if the problem persists.

REPORTING

If the test is positive, report as Staphylococcus aureus, with comment MRSAP = MRSA positive (by Antigen detection).  For cultures with a full sensitivity being performed, i.e. blood cultures, put in comment “Full Sensitivity pending”

If the test is negative, report as Staphylococcus aureus, with comment MRSAN = MRSA Negative (by antigen detection). For cultures with a full sensitivity being performed, i.e. blood cultures, put in comment “Full Sensitivity pending” ()

SAFETY

Follow Standard Precautions