Technical SOP

| Title | Thrombin Time (TT) By STA Compact |  |
| --- | --- | --- |
| Prepared by | Harvey Vandenburg | Date: 04/27/2006 |
| Owner | Ruben Layug | Date: 11/6/2012 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	6

7.	Equipment And Supplies	8

8.	Procedure	8

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	11

13.	Procedure Notes	11

14.	Limitations Of Method	12

15.	Safety	12

16.	Related Documents	13

17.	References	13

18.	Revision History	13

19.	Addenda	14

# TEST INFORMATION

| Assay | Method/Instrument | Order Code |
| --- | --- | --- |
| Thrombin Time | STA Compact | THR |

Synonyms/Abbreviations

TT

Department

Providence Hospital Coagulation

# ANALYTICAL PRINCIPLE

In the presence of a predetermined quantity of thrombin, normal tissue plasma will clot unless there is an abnormality of fibrinogen, or there is a form of antithrombin present.  After thrombin is added to citrated plasma, the time of clot formation is determined.  The instrument uses an electromagnetic clot detection system. The oscillation of a steel ball within the thrombin and plasma is monitored.  When the oscillation of the steel ball is stopped by clot formation, the sensor registers the time in seconds.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | none |
| Specimen Collection and/or Timing | If multiple tests are being drawn, the coagulation specimen must be done first. If only coagulation studies are being drawn, use the two-tube /syringe technique. Draw 5 mL into the first tube or syringe, discard, and then draw the specimen for coagulation studies. Vacutainer tubes must be filled to the blue line to ensure the proper ratio of blood to anticoagulant. When drawing blood from an indwelling line that may contain heparin, flush the line with 5 mL of saline. Draw off and discard the first 5 mL of blood before filling the coagulation tube. |
| Special Collection Procedures | N/A |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Whole Blood (LB) |  |
| -Other Acceptable | Plasma (LB) |  |
| Collection Container | Light blue top tube (3.2% sodium citrate) |  |
| Volume - Optimum | 4.5 mL whole blood 2.5 mL plasma |  |
| - Minimum | 3.0 mL whole blood 1.0 mL plasma |  |
| Transport Container | Light blue top tube (3.2% sodium citrate) |  |
| Stability & Storage Requirements | Room Temperature: | Plasma: 8 hours Whole Blood: 1 hour |
|  | Refrigerated: | Do not refrigerate |
|  | Frozen: | Do not freeze |
| Timing Considerations | After collection, whole blood for coagulation studies should be transported to the laboratory within 20 minutes. If there is a delay in transport, the specimen should be maintained at room temperature and transported as quickly as possible. Centrifuge within one hour of collection. |  |
| Unacceptable Specimens & Actions to Take | Clotted: test not performed, request a recollection Gross hemolysis: test not performed, request a recollection QNS: test not performed, request a recollection Short draw: Whole blood not drawn to blue line. Test not performed, request a recollection Too old: Whole blood received greater than 1 hour after collection. Test not performed, request a recollection |  |
| Compromising Physical Characteristics | Hemolysis: Avoid hemolysis Lipemia: Acceptable Icterus: Acceptable Refer to section 14.3 for additional information. |  |
| Other Considerations | Ensure that centrifugation yields platelet poor plasma (<10 x 109/L). |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| STA THROMBIN 10 | Diagnostica Stago- 00669 | 12 X 10 mL |
| NERL reagent grade H2O | NERL Diagnostics 3650 | 6 X 480 mL |
| STA Desorb U | Diagnostica Stago- 00995 | 12 per box |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | STA THROMBIN 10 |
| --- | --- |
| Container | Manufacturer Supplied Vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date on the vial Once reconstituted, remains stable for 7 days on STA |
| Preparation | Reconstitute the vial with 10 mL of distilled water. Allow the reconstituted reagent to stand at room temperature for 30 minutes. Swirl the reconstituted reagent gently to obtain a homogenous suspension without producing any bubbles, place it a STA-maxi Reducer, and place white perforated cap on the vial. |

| Reagent | STA Desorb U |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date indicated on the box label |
| Preparation | Ready for use |

| Reagent | NERL reagent grade H2O |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | Room temperature |
| Stability | Stable for 30 days after opening |
| Preparation | Ready for use |

# CALIBRATORS/STANDARDS

With each new lot of Thrombin 10, the operator must enter the reference time before the STA compact will allow QC to be run.  To enter the reference time:

- From the Main Menu, select CALIB/CONTROL.

- Select CALIBRATION, press Enter.

- Cursor to the PT test and press Enter.

- Press ESC for options

- Chose CHANGE REFERENCE TIME and press Enter

- Enter the reference time and save with F10.

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| STA Coag Control N + ABN | Diagnostica Stago- 00676 (N + ABN kit) |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | STA Coag Control N + ABN |
| --- | --- |
| Preparation | Reconstitute each vial with exactly 1.0 mL NERL reagent grade water. Allow the reconstituted vial to stand at room temperature for 30 minutes. Swirl gently |
| Storage/Stability | The reagents in unopened vials are stable until the expiration date indicated on the box label, when stored at 2-8 degrees C. Once reconstituted, the controls are stable for 8 hours on the instrument. |

  - Frequency

    - Both control N and control ABN are run daily at the beginning of each 8 hour   shift.

  - Tolerance Limits

    - The established QC ranges are lot specific, and are in the STA QC file.

    - The STA Compact monitors all control ranges automatically.  If any control exceeds + 2 SD, the instrument will audibly and visually alarm the operators. The 1-2s rule will be used for “rejection”.  If one QC result in the run exceeds 2 SD, run will be considered unacceptable.  In this case, the technologist must rerun the QC before running any patient sample

    - Rejected runs must be effectively addressed by corrective action.  Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed.

      - Troubleshooting QC.  If QC fails, take the following steps:

- Ensure that you are using the correct control lot number and the correct reagent lot number.

- Rerun the control using freshly prepared control material.

- If these steps fail to resolve the problem, attempt to run the assay on another analyzer.

- If QC is still unacceptable, immediately contact your supervisor.

    - Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, if the problem was resolved, how patient samples were handled and/or results impacted, and the date and initials of the person recording the information.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

  - Documentation

    - Control results are automatically filed in the STA Compact QC file. All results for a 24-hour period are reduced to a mean value, which is used in statistical analysis, and plotted on Levy-Jennings charts.

    - To print all QC data points, perform the following procedure prior to midnight. From the main menu under CALIB/CONTROL select QUALITY CONTROL.  Cursor to the TT test and press enter to view the Levy-Jennings chart. Press F1 to view the results in tabular form.  Press F6 then enter to print the individual values under current controls.  Press ESC to exit back to the graph.  Press F2 or F3 to view other levels and continue with F1 to view the result list.

    - Refer to the Record Retention SOP for record retention requirements.

  - Review of QC

    - Weekly and Monthly QC must be reviewed by the Supervisor or designee on a weekly and monthly basis, if the SD’s or CV’s are greater than the established ranges, investigate the cause for the imprecision and document implementation of corrective actions.

    - Monthly QC must be presented to Medical Director or designee for review and signature.

6.7	Quality Assurance Program

    - Refer to corporate and local polices and procedures for other quality assurance activities applicable to this procedure.

    - Each new lot number of reagent must be tested with external control materials and/or previously analyzed samples.  Performance of the new lot must be equivalent to the previous lot and must be document in the parallel check book.

    - All persons performing this assay must complete training and reviewed annually for competency, the result are documented on file.

    - This procedure must be incorporated into the departmental competency assessment program.

    - The laboratory participates in CAP proficiency testing.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

STA Compact

7.2	Equipment

Centrifuge

7.3	Supplies

Calibrated Pipettes

Pipette tips

Cuvette Rolls

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | Perform Daily Maintenance: - From the Main Menu, press STATUS, then SYSTEM and examine the screen for the daily maintenance checks. - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. - Record Maintenance in the maintenance log as soon as it is performed. - If any temperature check exceeds the tolerance limit, then check the peltier fluid, and the fluid circulation. If the fluid is low, replenish. If this does not solve the problem, call the STAGO service number on the instrument. |
| 2. | Perform Weekly, Monthly, or Four Month Maintenance (if required) - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. |
| 3. | Load cuvettes and cleaner/wash solution on the analyzer if needed. To load Desorb, request the reagent drawer to open through the main menu under Loading. Scan the barcode of the reagent. Scan the barcode of the reagent. Place the reagent into a position 40 or 41 in the product drawer. |
| 4. | For more details on Instrument set-up consult the operator’s manual. |

| 8.2 | Specimen / Reagent Preparation |
| --- | --- |
| 1. | Prepare the TT reagent as stated in the reagent preparation section. |
| 2. | Prepare controls as stated in the control preparation section. |
| 3. | Load reagents. Request the reagent drawer to open through the main menu under Loading. Scan the barcode of the TT reagent. Place the TT reagent into any position 19-24 in the product drawer. |
| 4. | Load controls. Place the N control into well 1 and the ABN control into well 2. |
| 5. | When all reagents and controls are loaded, press ESC. When the Quit box comes up, chose YES to exit the reagent-loading menu. The loading drawer will close. |
| 6. | Prepare samples by centrifuging citrated whole blood to produce a platelet poor sample. |
| 7. | For more details on specimen/reagent preparation, consult the operator’s manual. |

| 8.3 | Test Run |
| --- | --- |
| 1. | Request quality control. Through Main Menu under CALIB/CONTROL, select QUALITY CONTROL and press enter. Cursor to TT, and select by pressing F1 and then F10 to run QC. |
| 2. | If QC is acceptable, load patient’s samples. Access the sample drawer through the Main Menu under LOADING. After the drawer opens, identify the type of specimen. If micro sample, press F8. If STAT, press F12. Identify the sample by bar coding or manually entering the patient identification number on the keyboard. Place the sample into the sample drawer. Move the cursor to the appropriate RECORDED PROFILES to select the profile needed. If a single test is needed, move the cursor to selection and select the test needed. |
| 3. | Once samples are loaded, press ESC, and then Enter on the Quit menu. The sample drawer will close. Once you exit out of the test menu, the instrument will immediately start processing samples. |
| 4. | The TEST STATUS screen will appear. If there is not enough reagents to run the test, the reagent will appear in red with the amount of deficiency. This deficiency will BLOCK the sample pipettor. Add the deficient reagents to run samples, per the directions above. Go to the SYSTEM menu, and cursor down to STOP SAMPLE PIPETTING YES. Press Enter to change to NO. Press ESC and the instrument will begin to run. |
| 5. | All patient results are displayed on the TEST PANEL screen and automatically print out and transmit to the LIS if selected. |

| 8.4 | Special Handling |
| --- | --- |
| 1. | For results that require operation intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will bring up the FILE PROCESSING screen. Follow the options in the left-hand corner of the screen. |

# CALCULATIONS

None.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

Not applicable

10.2	Rounding

No rounding is needed.

10.3	Units of Measure

The TT is reported in seconds, to the nearest tenth.

10.4	Clinically Reportable Range (CRR)

10-58 seconds

  - Repeat Criteria and Resulting

| IF the TT result is … | THEN… |
| --- | --- |
| <10 seconds | Request a redraw. If the retest on the redrawn sample matches, then release the result as <10 and enter an HCOM that the result has been verified. |
| > 10 seconds and <58 seconds | Report the thrombin time to the nearest tenth of a second in the LIS. |
| >58 seconds | Check the sample for clots. If clotted request a redraw. If not, rerun. If the rerun matches within the instrument’s flagging limits, then release the result as “>58 seconds”. |

# EXPECTED VALUES

11.1	Reference Ranges

14.0 – 19.5 seconds

11.2	Priority1, 2 & 3 Results

Refer to the Critical Value Policy and Procedure

# CLINICAL SIGNIFICANCE

Prolongation of the TT indicates:

- An abnormality of fibrinogen

  - Qualitative: dysfibrinogenemia

  - Quantitative:

    - Severe hypofibrinogenemia or congenital afibrinogenemia

    - Acquired hypofibrinogenemia (DIC, Fibrinolysis, liver disease)

- The presence of antithrombins

  - Therapeutic:  heparin, hirudin, argatroban, etc.

  - Abnormal:  FDP, appearing during Myeloma, etc.

# PROCEDURE NOTES

FDA Status: Approved

Validated Test Modifications: None

Very sensitive to heparin contamination.

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

10-58 seconds.

14.2	Precision

|  | Intra-Assay Reproducibility |  | Inter-Assay Reproducibility |  |
| --- | --- | --- | --- | --- |
|  | Sample 1 | Sample 2 | Sample 3 | Sample 4 |
| n | 21 | 21 | 10 | 10 |
| mean | 19.1 | 32.2 | 17.9 | 33.4 |
| SD | 0.53 | 0.55 | 0.29 | 1.09 |
| CV | 2.79 | 1.69 | 1.63 | 3.25 |

14.3	Interfering Substances

    - Samples suspected of being clotted should not be tested.

    - Do not test samples that have been partially clotted.

    - Do not test samples that have been contaminated by heparin (in collection tubes, syringes, etc.)

14.4	Clinical Sensitivity/Specificity/Predictive Values

Often performed only if an initial panel of tests excludes more common disorders, because dysfibrinogenemia is uncommon. To increase specificity, use TT in concert with the Reptilase Time in dysfibrinogenemia workups.  The Reptilase time is not prolonged in the presence of heparin or hirudin.

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needle sticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- Warnings of other specific hazards are noted in this procedure.  Please comply with the requirements to reduce your risk of injury."

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

Providence Hospital Laboratory Panic Value Standard Operating Procedure

STA Compact Operator’s Manual

Current Package Insert

# REFERENCES

- Jacobs, D.S.; De Mott, W.R; and Oxley D.K.; Laboratory Test Handbook, 5th edition,. Lexi-Comp Inc., Cleveland, OH. 2001.

- THROMBIN 10, Determination of the Thrombin by STA Analyzers; Package Insert REF 0000669, 2003.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| .000 | 09/04/2008 | Page 1 | Changed local to laboratory approval. Updated owner. | Omar Dualeh | R. Levy |
| .000 | 09/04/2008 | 11.2 | Update section title to new terminology | Omar Dualeh | R. Levy |
| .000 | 09/04/2008 | 13 | Added FDA Status and Validated test Modifications | Omar Dualeh | R. Levy |
| .001 | 01/19/2010 | 19 | Updated Addendum (Deleted package insert) | Omar Dualeh | R.Levy |
| .002 | 10/14/10 | Page 1 | Update Owner | A Foster | R. Levy |
|  |  | 8.0 | Updated Procedure to reflect removal of package inserts |  |  |
|  |  | 16.0 | Updated related documents to include current product insert |  |  |
| .003 | 12/12/2011 | Page 1 | Updated owner | W. McMillan | R. Levy |
|  |  | 1 | Updated Department to “Providence Hospital Coagulation” |  |  |
|  |  | 3.1 | Removed Special Collection Procedures. No evidence of necessity. |  |  |
|  |  | 3.2 | Removed specific parameters for centrifuges. |  |  |
|  |  | 6.3 | Changed QC frequency to every 6 hours |  |  |
|  |  | 6.4.4 | Added requirement of noting patient impact after QC run failure. |  |  |
|  |  | 6.5.3 | Removed reference to “Administrative Manual” |  |  |
|  |  | 7.3 | Added requirement for Calibrated pipette. |  |  |
| .004 | 11/6/2012 | Page 1 | Updated owner. | R. Layug | R. Levy |
|  |  | 3.1 | Changed order of draw from last to first. |  |  |
|  |  | 3.2 | Modified to Light blue container only. |  |  |
|  |  | 6.3 | Changed QC frequency to every 8 hours |  |  |

# ADDENDA

| Addendum | Title |
| --- | --- |