Technical SOP

| Title | Activated Partial Thromboplastin Time (APTT) Mixing Study By STA Compact |  |
| --- | --- | --- |
| Prepared by | Omar Dualeh | Date: 09/16/2008 |
| Owner | Ruben Layug | Date: 11/9/2012 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	5

7.	Equipment And Supplies	8

8.	Procedure	8

9.	Calculations	9

10.	Reporting Results And Repeat Criteria	9

11.	Expected Values	10

12.	Clinical Significance	10

13.	Procedure Notes	11

14.	Limitations Of Method	11

15.	Safety	11

16.	Related Documents	12

17.	References	12

18.	Revision History	12

19.	Addenda	13

# TEST INFORMATION

| Assay | Method/Instrument | Order Code |
| --- | --- | --- |
| Mixing Studies | STA Compact | COCS |

Synonyms/Abbreviations

COCS

Department

Providence Hospital Coagulation

# ANALYTICAL PRINCIPLE

There are two types of mixing studies: the immediate mix and the incubated mix. In the immediate mix, the aPTT is performed immediately after mixing the patient plasma and normal plasma, without further incubation. Specimens with fast-reacting factor inhibitors will not correct the immediate mix.

In contrast, specimens with slow-reacting, or time-dependent, factor inhibitors may correct the immediate mix, giving the false impression of a factor deficiency. The majority of factor VIII inhibitors and eight percent of lupus anticoagulants are sufficiently time dependent that they correct the immediate mix. Therefore, it is essential that an incubated mix be performed to diagnose these conditions. In the incubated mix, the aPTT is performed after incubating the mixture of patient plasma and normal plasma for one hour at 37°.

The incubated mixing study is necessary for diagnosing coagulation factor inhibitor.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None |
| Specimen Collection and/or Timing | Please see APTT SOP. |
| Special Collection Procedures | Please see APTT SOP. |
| Other | NA |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Whole Blood (LB) |  |
| -Other Acceptable | Plasma (LB) |  |
| Collection Container | Light blue top tube (3.2% sodium citrate) |  |
| Volume - Optimum | 4.5 mL whole blood 2.5 mL plasma |  |
| - Minimum | 3.0 mL whole blood 1.0 mL plasma |  |
| Transport Container | Light blue vacutainer or a clean plastic (frozen screw-capped vial for plasma only). See Timing Considerations below. |  |
| Stability & Storage Requirements | Room Temperature: | Plasma: 8 hours Whole Blood: 4 hours |
|  | Refrigerated: | Plasma: 12 hours at 4°C Whole Blood: 8 hours |
|  | Frozen: | Plasma: Up to 24 hours at -20°C Whole Blood: Unacceptable |
| Timing Considerations | After collection, whole blood for coagulation studies should be transported to the laboratory within 20 minutes. If there is a delay in transport, the specimen should be maintained at room temperature and transported as quickly as possible. Centrifuge with one hour of collection. |  |
| Unacceptable Specimens & Actions to Take | Please see APTT SOP. |  |
| Compromising Physical Characteristics | Please see APTT SOP. |  |
| Other Considerations | Please see APTT SOP. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| STA PTT A 5 | Diagnostica Stago - 00595 | 12 X 5mL |
| NERL reagent grade H2O | NERL Diagnostics - 3650 | 6 X 480 mL |
| STA Desorb U | Diagnostica Stago - 00995 | 12 per box |
| STA CaCl2 0.025M | Diagnostica Stago - 00367 | 24 X 15 mL |
| Cryocheck Pooled Normal Plasma | Precision Biologic - CCN-40 | 81 x 4.0 mL |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | STA PTT A 5 |
| --- | --- |
| Container | Manufacturer Supplied Vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date on the vial Once reconstituted, remains stable for 24 hours on STA |
| Preparation | Reconstitute the vial with 5 mL of distilled water. Allow the reconstituted reagent to stand at room temperature for 30 minutes. Swirl the reconstituted reagent gently to obtain a homogenous suspension, and place white perforated cap on the vial. |

| Reagent | STA CaCl2 0.025M |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | 2-25 degrees C |
| Stability | Unopened: Stable until expiration date on the vial After opening, stable for 24 hours at 37 degrees C; and 72 hours on the STA compact. |
| Preparation | No reconstitution required. Allow the reagent to stand at room temperature for 30 minutes. Do not install either a STA-reducer or a perforated cap on the solution bottle. |

| Reagent | STA Desorb U |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date indicated on the box label |
| Preparation | Ready for use |

| Reagent | NERL reagent grade H2O |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | Room temperature |
| Stability | Stable for 30 days after opening |
| Preparation | Ready for use |

| Reagent | Cryocheck Pooled Normal Plasma |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | -20 to -70 deg C |
| Stability | Stable for 30 days after opening |
| Preparation | Ready for use |

# CALIBRATORS/STANDARDS

Please see APTT SOP.

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| STA Coag Control N + ABN | Diagnostica Stago- 00676 (N + ABN kit) |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | STA Coag Control N + ABN |
| --- | --- |
| Preparation | Reconstitute each vial with exactly 1.0 mL NERL reagent grade water. Allow the reconstituted to stand at room temperature for 30 minutes. Swirl gently |
| Storage/Stability | The reagents in unopened vials are stable until the expiration date indicated on the box label, when stored at 2-8 degrees C. Once reconstituted, the controls are stable for 8 hours on the instrument. |

  - Frequency

Please See APTT SOP.

  - Tolerance Limits

    - The established QC ranges are lot specific, and are in the STA QC file.

    - The STA Compact monitors all control ranges automatically.  If any control exceeds + 2 SD, the instrument will audibly and visually alarm the operators. The 1-2s rule will be used for “rejection”.  If one QC result in the run exceeds 2 SD, run will be considered unacceptable.  In this case, the technologist must rerun the QC before running any patient sample

    - Rejected runs must be effectively addressed by corrective action.  Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed.

      - Troubleshooting QC.  If QC fails, take the following steps:

- Ensure that you are using the correct control lot number and the correct reagent lot number.

- Rerun the control using freshly prepared control material.

- If these steps fail to resolve the problem, attempt to run the assay on another analyzer.

- If QC is still unacceptable, immediately contact your supervisor.

    - Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, potential patient impact and the date and initials of the person recording the information.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

  - Documentation

    - Control results are automatically filed in the STA Compact QC file. All results for a 24 hour period are reduced to a mean value, which is used in statistical analysis, and plotted on Levy-Jennings charts.

    - To print all QC data points, perform the following procedure prior to midnight. From the main menu under CALIB/CONTROl select QUALITY CONTROL.  Cursor to the APTT test and press enter to view the Levy-Jennings chart. Press F1 to view the results in tabular form.  Press F6 then enter to print the individual values under current controls.  Press ESC to exit back to the graph.  Press F2 or F3 to view other levels and continue with F1 to view the result list.

    - Refer to Record Retention SOP for record retention requirements, found in the Administrative Manual.

  - Review of QC

    - Weekly and Monthly QC must be reviewed by the supervisor on weekly and monthly basis, if the SD’s or CV’s are greater than the established ranges, investigate the cause for the imprecision and document implementation of corrective actions.

    - Monthly QC must be presented to Medical Director for review and signature.

6.7	Quality Assurance Program

    - Refer to corporate and local polices and procedures for other quality assurance activities applicable to this procedure.

    - Each new lot number of reagent must be tested with external control materials and/or previously analyzed samples.  Performance of the new lot must be equivalent to the previous lot and must be document in the parallel check book.

    - All persons performing this assay must complete training and reviewed annually for competency, the result are documented on file.

    - This procedure must be incorporated into the departmental competency assessment program.

    - The laboratory participates in CAP proficiency testing.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

STA Compact

7.2	Equipment

Centrifuge capable of producing platelet poor plasma

Water bath

7.3	Supplies

Calibrated Pipettes

Pipette Tips

Cuvette Rolls

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | Perform Daily Maintenance: Please See APTT SOP |
| 2. | For more details on Instrument set-up consult the operator’s manual |

| 8.2 | Specimen / Reagent Preparation |
| --- | --- |
| 1. | Please See APTT SOP |
| 2. | For more details on specimen/reagent preparation, consult the operator’s manual |

| 8.3 | Test Run |
| --- | --- |
| 1. | Please See APTT SOP |
| 2. | All patient results are displayed on the TEST PANEL screen and automatically print out and transmit to the LIS if selected. |

| 8.4 | Special Handling |
| --- | --- |
| 1. | Please See APTT SOP. |

| 8.5 | Mixing Studies |
| --- | --- |
| 1. | Mixing Studies is done by; Mixing equal volumes of patient plasma and Cryocheck Pooled Normal Plasma then repeat the aPTT on the mixture. Do the dilution as follow 1:1 dilution = mix 500 uL (0.5ml) patient plasma with 500 uL (0.5 ml) of Cryocheck Pooled Normal Plasma. Make sure the mixture is well mixed before testing. |
| 2. | In a small tube, pipette equal volumes of STA- Coag ABN control and Cryocheck Pooled Normal Plasma. |
| 3. | In another small tube, pipette equal volumes of patient plasma and Cryocheck Pooled Normal Plasma and Incubate at 37º for 1 hour. |
| 4. | Mix the tubes carefully, and transfer the incubated sample to labeled aspiration cups. |
| 5. | Load the control and patient samples into the sample drawer using the procedure defined in section 8.2 above. Note: PT incubated study is never performed. |

# CALCULATIONS

None.

# REPORTING RESULTS AND REPEAT CRITERIA

  - Interpretation of Data

Mixing Studies:  If the addition of Cryocheck Pooled Normal Plasma causes a correction of an elevated APTT toward the normal range, a factor deficiency is indicated. Specific factor assays may be indicated.  No correction may indicate the presence of an inhibitor.

Correction is defined as a 1:1 mixture result within the reference range. (Less than or equal to 11.5 sec)

10.2	Rounding

N/A

10.3	Units of Measure

Please see APTT SOP.

10.4	Clinically Reportable Range (CRR)

N/A

  - Repeat Criteria and Resulting

| IF the APTT result is … | THEN… |
| --- | --- |
| APTT is Normal | STOP |
| APTT is Abnormal | Order Mixing study |

Reporting Mixing Studies:

Enter results on the Labweb (TACTI) screen by using the code MIXR Enter results in screen as follows:

Mixing Study Results

Baseline Coagulation Studies		PT   (enter patients results) secs.

APTT (enter patients results) secs.

Mixing Studies:

Mixed 1:1 Dil with normal plasma	APTT (enter patients results) secs.

Mixed 1:1 Dil with normal plasma

Incubated at 37oC for 1 Hour  	APTT (enter patients results) secs.

# EXPECTED VALUES

11.1	Reference Ranges

N/A

11.2	Priority 1, 2 & 3 Results

Refer to the Critical Value Policy and Procedure

# CLINICAL SIGNIFICANCE

The mixing study is a common coagulation test used to distinguish between a coagulation factor deficiency, such as factor VIII deficiency, and a factor inhibitor such as.

Congenital Deficiencies:

- If the PT is normal, the following factors may be deficient:  VIII, IX, XI, and XII.

- If all of these factors are normal, a consider deficiency in prekalikrein (Fletcher factor) or HMW Kininogen (Fitzgerald factor).

Acquired Deficiencies and Abnormal Conditions

- Liver disease

- DIC

- Circulating anticoagulants, such as Lupus Anticoagulant

Therapeutic Monitoring:

- During heparin or oral anticoagulant therapy

- Treatment with thrombin inhibitors

- Note:  Platelet counts should be followed to ensure rapid diagnosis that if heparin-induced thrombocytopenia develops.

# PROCEDURE NOTES

# FDA Status: Approved

Validated Test Modifications: None

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

N/A

14.2	Precision

|  | Intra-Assay Reproducibility |  | Inter-Assay Reproducibility |  |
| --- | --- | --- | --- | --- |
|  | Sample 1 | Sample 2 | Sample 3 | Sample 4 |
| N | 21 | 21 | 10 | 10 |
| Mean | 30.8 | 52.7 | 31.6 | 54.7 |
| SD | 0.20 | 0.53 | 0.44 | 1.46 |
| CV | 0.7 | 1.0 | 1.4 | 2.7 |

14.3	Interfering Substances

Please see APTT SOP

14.4	Clinical Sensitivity/Specificity/Predictive Values

Please see APTT SOP.

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- Warnings of other specific hazards are noted in this procedure.  Please comply with the requirements to reduce your risk of injury."

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

Providence Hospital Laboratory Panic Value Standard Operating Procedure

STA Compact Operator’s Manual

Current Package Insert

# REFERENCES

- Lossing TS, Kasper CK, Feinstein DI. Detection of factor VIII inhibitors with the partial thromboplastin time. Blood. 1977; 49: 793-797.

2    STA-PTT A 5, Determination of Activated Partial Thromboplastin Time (APTT) by STA

Analyzers; Package Insert REF 0000595, 2004.

- Clyne LP, White PF. Time dependency of lupuslike anticoagulants. Arch Intern Med.

1988; 148: 1060-1063

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| .000 | 01/19/10 | 10.5 | Changed Result Entering Code To MIXR | O Dualeh | Dr V.Padmanabha |
| .000 | 01/19/10 | 19. | Updated Addendum (deleted package insert) | O Dualeh | R.Levy |
| .002 | 10/14/10 | Page 1 | Update Owner | A Foster | R. Levy |
|  |  | 8.0 | Updated Procedure to reflect removal of package inserts |  |  |
|  |  | 16.0 | Updated related documents to include current product insert |  |  |
| .003 | 12/12/2011 | Page 1 | Updated Owner | W.McMillan | R. Levy |
|  |  | 1 | Changed Department to Coagulation |  |  |
|  |  | 6.4.4 | Added “Potential Patient Impact” |  |  |
|  |  | 7.2 | Added “Capable of producing Patient Poor Plasma” |  |  |
|  |  | 7.3 | Added “Calibrated” |  |  |
| .004 | 11/9/2012 | Page 1 | Updated owner. | R. Layug | R. Levy |
|  |  | 4.1 and 4.2 | Added Cryocheck Pooled Normal Plasma to reagent list. |  |  |
|  |  | 8 and 10 | Replaced Normal pooled plasma with Cryocheck Pooled Normal Plasma. |  |  |

# ADDENDA

| Addendum | Title |
| --- | --- |