Non-Technical SOP

| Title | Proficiency Test Handling and Result Submission |  |
| --- | --- | --- |
| Prepared by | Karen Rupke, Pat Maloney, Christine Vernusky | Date: 12/15/2011 |
| Laboratory Approval | Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval date. |  |  |
| Local Issue Date: | Local Effective Date: |  |
| Review |  |  |
| Print Name and Title | Signature | Date |

| Corporate Approval |  |  |
| --- | --- | --- |
| Owner | Karen Rupke Director, Corporate Quality Management |  |
| Signature | Karen Rupke Electronic Signature | Date: 1/6/2012 |
| V. P and Chief Laboratory Officer | Stephen C. Suffin, M.D. |  |
| Signature | Stephen C. Suffin, M.D. Electronic Signature | Date: 1/6/2012 |
| Corporate Issue Date | 1//9/2012 |  |

| Retirement Date: |  |
| --- | --- |
| Reason for retirement/replacement: |  |

TABLE OF CONTENTS

1.	PURPOSE	2

2.	SCOPE	2

3.	RESPONSIBILITY	3

4.	DEFINITIONS	4

5.	OVERVIEW OF GENERAL REQUIREMENTS FOR PROFICIENCY TESTING	4

6.	ENROLLMENT IN PROFICIENCY TEST PROGRAMS	6

7.	PROGRAM REQUIREMENTS POINT OF CARE TESTING (POCT) WITHIN A MODERATE OR HIGH COMPLEXITY LABORATORY:	7

8.	ALTERNATIVE PERFORMANCE ASSESSMENT	7

9.	RECEIPT OF PROFICIENCY TEST SAMPLES	7

10.	TESTING PROFICIENCY SAMPLES	8

11.	DOCUMENTATION OF PROFICIENCY TEST RESULTS	9

12.	REVIEW AND SUBMISSION OF PT RESULTS	10

13.	RECORDS MAINTENANCE	10

14.	RELATED DOCUMENTS	10

15.	REFERENCES	10

16.	REVISION HISTORY	11

17.	ADDENDA	11

# PURPOSE

This document sets forth the procedure for receipt, handling, and testing of proficiency test (PT) samples and for submitting PT results to any external agency (federal, state, or other) for evaluation.

# SCOPE

- This procedure applies to proficiency testing irrespective of regulated status of the analyte designated by CMS (e.g., regulated and non-regulated analytes).

- This procedure applies to all proficiency material irrespective of the provider (e.g., CAP, AAB, WLSH, Pennsylvania, New York or other State agency).

- This procedure applies to all appropriate personnel (pre-analytical, analytical, and post-analytical) in all Quest Diagnostics laboratories, including main and esoteric laboratories (e.g., Quest Diagnostics Nichols Institutes, Focus, Athena, Berkeley HeartLab, etc.), RRLs, all remote accessioning locations (including off-site Quest Diagnostics employees that accession specimens), Quest Diagnostics managed sites or hospital labs with licenses in the Quest Diagnostics name except as detailed below:

- All staff (as appropriate for their job description) will be trained on this process at:

  - New Hire

  - Whenever changes are made to the process

EXCEPTIONS TO THIS PROCEDURE:

- Stand alone Point of Care Testing Sites (POCT). These sites must follow the standard POCT policy specific for this type of testing (Proficiency Test Handling and Results Submission for Point of Care Testing Sites, QDPS82).

# RESPONSIBILITY

- The Laboratory Director (CLIA license holder) has overall responsibility for:

- Local approval and implementation of this standard operating procedure (SOP), including all subsequent revisions.

- Ensuring laboratory enrollment in PT programs as required by Quest Diagnostics and regulatory agencies, including CAP e-LAB Solutions PT program (refer to www.cap.org for enrollment instructions).

- Ensuring the implementation of this SOP in all relevant departments.

- Ensuring compliance with this SOP.

- The Laboratory Director or designee is responsible for:

- The recurring review of this procedure.

- Ensuring appropriate handling of PT materials.

- Signing the attestation form (or copy of the completed form).

- The Technical Supervisor is responsible for:

- Implementing and ensuring compliance with this procedure in the analytical department(s) or laboratory(ies) for which he/she is responsible.

- Review of PT results, instrument/method information and test units for accuracy.

- Ensuring documented training in his/her department(s) for all employees including new hires (within 30 days of hire AND prior to handling PT samples).

- The Testing personnel are responsible for:

  - Testing PT samples in the same manner as patient samples, except where the nature of the PT material requires special handling; and

  - Documenting performance of all steps in the proficiency testing process.

  - Processing PT samples with normal workflow (if not, document why).

  - Utilizing the same repeat/dilution protocols and same calibration and quality control frequency.

  - Signing the attestation form (or copy of the completed form).

- Quality Assurance, as delegated by the Laboratory Director (in writing, see Addendum B ) is responsible for:

- Coordinating Proficiency Testing including oversight of the log-in process.

- Ensuring that reflex testing order codes that would normally trigger referral of the patient sample to another laboratory for further testing is not used for any PT.  A unique order code without reflex must be built for PT.

- Ensuring results are reviewed and reported by the deadline to the PT provider.

- Maintaining required enrollment in PT programs as required by Quest Diagnostics and regulatory agencies.  Appropriate enrollment must be reviewed annually.

- Managing the handling of all PT materials.

- Providing oversight of the PT program for affiliated Rapid Response Laboratories (or in the case of hospitals, providing oversight of the PT program for all departments in the laboratory).

- Administering the CAP e-LAB Solutions program.

- Ensuring Alternative Performance Assessment is performed on analytes that are not covered by external proficiency testing programs.  The need for alternative performance assessment must be reviewed annually.

# DEFINITIONS

- Alternative Performance Assessment (APA): A laboratory administered program similar to proficiency testing that is used to evaluate performance of assays (i.e., analytes that are not covered by CAP PT, other Commercial/State PT or National Quality Assurance (NQA) PT).

- CLIA:  Clinical Laboratory Improvement Amendments.

- CMS:  Centers for Medicare and Medicaid Services.   The Centers for Medicare & Medicaid Services (CMS) regulate all laboratory testing (except research, forensic and SAMHSA drug testing) performed on humans in the U.S. through the Clinical Laboratory Improvement Amendments of 1988 (CLIA).

- Proficiency Testing (PT): The process that uses samples from a PT provider to demonstrate the laboratory’s ability to provide accurate and reliable results in its patient testing system. The process incorporates elements of pre-analytical, analytical, and post-analytical areas. Educational challenges are to be handled the same way as routine proficiency testing.

# OVERVIEW OF GENERAL REQUIREMENTS FOR PROFICIENCY TESTING

- Laboratory Director delegation of proficiency testing responsibilities must be in writing (see Addendum B) and list the specific individual(s) to whom the responsibilities have been assigned.

PROGRAM COMPLIANCE:

- The laboratory must participate in a CMS-approved proficiency program for all CLIA specialties and sub-specialties included in the laboratory’s testing menu.

- If a proficiency test program is not available, the analyte must be challenged by an Alternative Performance Assessment at least twice per year.

- Where required, the laboratory must also enroll in state mandated PT programs.

- Quest Diagnostics MUST NOT ACCEPT* proficiency sample(s) from another laboratory (including another Quest Diagnostics laboratory).

- Quest Diagnostics MUST NOT TEST* any proficiency sample(s) received from another laboratory (including another Quest Diagnostics laboratory).

- Quest Diagnostics MUST NOT REFER* any portion of a proficiency test sample to another laboratory (including another Quest Diagnostics laboratory).

- Quest Diagnostics MUST NOT ENGAGE* in either Intra- or Inter-laboratory communication about proficiency testing sample(s) before formal evaluation of results by the proficiency testing provider (including communication with another Quest Diagnostics laboratory). If there are concerns about the assay, reagents or run containing a PT sample, contact the Technical Supervisor.

* See Procedure for Handling Inappropriate Referral of Proficiency Material or Inter/Intra-laboratory Communication of Proficiency Test Information (QDNQA712).

TESTING COMPLIANCE:

- Unless explicitly directed otherwise by the PT provider in the written instructions, PT samples must be treated and reported like a patient sample. Exception: Do Not Refer any PT sample as you might for a patient (e.g., Reflex / Confirmatory testing)

- Whenever possible, survey samples must be accessioned into the Laboratory Information System (LIS).

- The main BU and each affiliated site (RRL, hospital) must use unique accounts in the LIS.  Access to such accounts should be restricted to prevent simultaneous review of results from the same survey across multiple sites.

- PT samples must be examined, handled, and tested along with the laboratory’s regular workload by testing personnel using the laboratory’s routine methods. (Some special handling may be required due to the nature of the PT materials, but the PT samples must be treated in the same manner as patient samples to the extent possible.)

- If a PT sample exceeds the analytical measurement range (AMR) of the assay, it must be tested and reported like a patient sample. For example, if patient samples are diluted and retested, the PT sample is diluted and retested. If patient samples are reported as “greater than”, the PT sample is reported as “greater than”.

- If reflex testing would normally trigger referral of the patient sample to another laboratory for further testing, the PT sample must not be referred to another laboratory. Only the initial screening result generated by the enrolled laboratory can be tested and reported. (Note:  Any reflex tests automatically generated by the LIS must have a unique order code built specific for proficiency testing without reflex / confirm.)

- There is to be no discussion of any aspect of an active PT event with others outside of the testing lab, such communications cannot occur until after the results have been formally evaluated by the PT provider.

- If the laboratory is unable to perform PT because an instrument or method is down:

- Notify the QA department.

- Alert the PT provider OR follow instructions on the PT forms.

- Ensure the order codes for the affected PT tests are canceled.

- DO NOT allow PT material to be sent to a referral laboratory for testing.

- Perform PT only for the primary method when multiple methodologies exist for a single test.

- Do not perform PT using the “Second Instrument” material option from CAP.

- PT samples must not be tested more than once unless a repeat protocol for patient testing is specifically defined by the test SOP and the PT sample meets the repeat criteria.

- When multiple persons and/or instruments are routinely used for patient testing, PT materials must be rotated among testing personnel, shifts, and instruments.

- Normal calibration protocols and schedules must be followed.

- Limit access of PT results to employees as a requirement of their job function – do not access other sites PT results.

- All survey documents, including copies of forms returned to the PT provider, must be retained on site at the performing laboratory for at least two years and be readily available for review. Off site storage (beyond the two most current years) must comply with the Quest Diagnostics Records Management Program requirements.

# ENROLLMENT IN PROFICIENCY TEST PROGRAMS

- Quality Assurance or designee must ensure that an approved PT program covers all analytes, or where PT is not available, is covered by an alternative performance assessment.

- Enrollment in a PT program or the need for Alternative Performance Assessment must be reviewed:

  - Whenever new analytes are added to the test menu.

  - Whenever current analytes are deleted from the test menu, referred to another laboratory or the methodology is changed.

  - On an annual basis, as proficiency testing needs are reassessed and documented.

TYPES OF PROGRAMS:

- The College of American Pathologists (CAP) Proficiency Test Program is the primary provider of proficiency test surveys for all Quest Diagnostics laboratories.

- The laboratory must enroll in the CAP e-LAB Solutions program.

- The laboratory must ensure enrollment in a state program if required by state law.

- Other PT providers may be used to meet proficiency testing requirements for specific analytes that are not covered by CAP programs or at the direction of NQA.

PROGRAM REQUIREMENTS:

- Proficiency testing is required only for the primary method used to test an analyte. Secondary methods are evaluated internally by performing method comparison studies with the primary method (twice each year).

- If multiple instruments are used for the same test, testing of additional instruments may only be performed after the survey evaluation report has been generated from the survey provider.

- Proficiency testing materials must be specific for the specimen type tested.

- Different sample types with different reference ranges and physiological concentrations require separate proficiency test programs (e.g., serum sodium and urine sodium, plasma glucose and CSF glucose, molecular assays such as Aptima, SDA, and PCA methods requiring different transports).

- Samples types that are documented to be equivalent (same physiological concentration and reference range) do not require separate proficiency test programs (e.g., plasma glucose and serum glucose).

- For miscellaneous “fluids”, a sample type with a representative fluid matrix and similar physiological concentration is sufficient (e.g., the lab must enroll in all applicable, specific fluid surveys when available).  Enrollment in the CAP Body Fluid (FLD) survey is appropriate for synovial, thoracic, etc.  However, where fluid specific material is available (e.g., CSF) the laboratory must enroll as appropriate.

- In special cases, substitute specimen types such as lyophilized culture organisms or photomicrographs may be appropriate.

# PROGRAM REQUIREMENTS POINT OF CARE TESTING (POCT) WITHIN A MODERATE OR HIGH COMPLEXITY LABORATORY:

- Only one survey may be ordered for each type of POCT performed. Any testing of additional instruments may only be performed after the survey evaluation report has been generated from the survey provider.

- If additional proficiency testing is needed, extra surveys may be ordered from a different CLIA approved proficiency testing provider. (e.g., blood gases are tested on i-Stat's at 5 different locations within the hospital. The laboratory may order only one CAP-AQ survey for the hospital.)

  - The exception to this is the survey for waived whole blood glucose testing.  The survey has been designed to accept multiple testing sites.

- Proficiency testing must be performed by the personnel that routinely perform the testing (e.g., nursing, respiratory, etc.).  Ideally the results should be faxed or sent electronically directly to CAP (or other survey provider) from the unit that performs the testing.

- The POCT manager may perform the secondary review providing the person performing the review does not review the same survey for any other testing sites.

# ALTERNATIVE PERFORMANCE ASSESSMENT

For tests not covered by CAP PT, other Commercial/State PT or National Quality Assurance (NQA) PT, the laboratory must develop alternative performance assessment systems to determine the reliability of analytic testing. The Alternative Performance Assessment system must:

  - Mimic proficiency testing program to the degree possible.

  - Be performed at least twice yearly.

  - Contain at least 2 challenges (e.g., two levels, positive and negative, etc.).

  - Be evaluated against established (documented) grading expectations.

  - Be reviewed by the same individual(s) who review(s) PT performance.

  - Have documented corrective action when grading expectations are not met.

# RECEIPT OF PROFICIENCY TEST SAMPLES

- Quality Assurance or designee must track the scheduled mailing dates for PT materials and follow-up with the appropriate PT provider when materials are not received as expected.

- Personnel in departments that initially receive PT materials must be trained to recognize these shipments and to immediately deliver them to the designated contact.

- Survey materials must always be personally delivered to the responsible individual (e.g., not left in a mailbox or on a desk).

      - Note: If survey materials must be stored overnight or over a weekend, the designated recipient must ensure that the materials are stored at the proper temperature (as indicated on the shipping package) and issue a written communication to the individual(s) who will be responsible for initial processing of the survey kit.

- The survey kit must be visually checked to ensure it is complete and that components are received in good condition.

- Required paperwork must be initiated to track the survey through the analytical process and result reporting.

      - Note: A Survey Companion Document (Addendum A) is furnished to assist laboratories in tracking PT specimens throughout the entire process and may be modified for local use.

- Survey samples and paperwork must be delivered to designated individuals in the testing area who are responsible for coordinating or performing PT testing.

- All affected personnel must be notified of the survey receipt and result due date.

- Whenever possible, PT samples must be accessioned into the Laboratory Information System (LIS). Unusual sample types that do not resemble patient specimens (such as photomicrographs) are exempt from this requirement.

# TESTING PROFICIENCY SAMPLES

PRE-ANALYTICAL REQUIREMENTS:

- All proficiency samples must be prepared according to survey instructions and properly labeled.

- Unusual conditions (e.g., leakage, hemolysis, particulate matter, turbidity, failure of samples that require repeating to give consistent results, etc.) must be reported to a supervisor immediately and documented.

- If sample integrity problems are observed, another specimen may be requested from the PT provider.

- If PT samples cannot be bar-coded, the identity of the sample must be visually verified each time it is used.

ANALYTICAL REQUIREMENTS:

- PT samples must be tested according to the test SOP and in the same manner as patient samples.

- PT samples must be built onto a patient load (when applicable) and must be analyzed by the same individuals who perform testing on patient samples.

- If, for some reason, the PT samples cannot be tested with an actual patient load (e.g., rarely ordered tests, short stability PT materials, etc.), the reason for the special run containing only PT samples must be documented.

- If multiple departments use the same PT samples, the department supervisor of the primary testing department must coordinate the testing process with other areas to ensure that sample stability is not exceeded.

- If a test is performed on multiple shifts, proficiency testing must be rotated among all shifts during the course of the year.

- An instrument must not be specially calibrated immediately prior to running PT samples.

- Repeat testing can only be performed when required by the test SOP and must meet the same repeat requirements used for patient testing.

- Photomicrographs or other prepared reference materials must be given to a single technologist. Consensus identification by a group of technologists is not appropriate. (Note: Consultation regarding unusual findings may be done according to the existing written protocol used for actual patient specimens.)

- QC release requirements must be the same as those used for patient result release.

- Predictive Markers (HER2 and Other Immunohistochemical Stains)

  - If you ordinarily send specimens to an outside laboratory for immunohistochemical staining (but do your own interpretation), the CAP samples must be also be sent to the same reference lab for staining.

  - The laboratory that interprets HER2 or predictive marker slides stained by another facility must enroll in an accepted PT program and report the results of their interpretation following their usual methods.

  - If the laboratory routinely accepts unstained slides on patient samples for immunohistochemical staining only, it MAY accept unstained PT survey slides for staining.

POST-ANALYTICAL REQUIREMENTS:

- When PT samples are logged into the LIS, results must be entered and released at the same time as patients on the same run.

- When PT results are released, LIS reports must be printed (as applicable) and retained with the PT survey results.

- After testing, any remaining PT material may be retained under appropriate storage conditions until after the survey results are received from the PT provider for possible use in survey failure investigations (e.g., review of sample labeling or retesting of sample, when possible).

- PT material, slides, photomicrographs, etc. may be used for educational and/or competency, after results have been formally evaluated by the PT provider.

# DOCUMENTATION OF PROFICIENCY TEST RESULTS

- Test records must clearly identify the individual(s) who performed the proficiency test, as well as the instrument(s) used.

- The attestation form (or copy of the completed form) must be signed by the individual(s) performing the testing.

- Testing personnel or other designated individual(s) in the department must transfer all required information to the PT result form (or copy of the form). All required information regarding instrumentation, method, results obtained, units of measure, attestation statements, etc. must be recorded and reviewed.

- The Technical Supervisor, Department Supervisor, Manager, or other designated individual must perform a secondary review of the PT documentation to ensure that all information has been correctly entered.

- The laboratory must maintain a distinctive file that includes copies of all associated test records (worksheet, instrument printout, printed reports etc.) and other documentation directly associated with the PT testing event.

# REVIEW AND SUBMISSION OF PT RESULTS

- Quality Assurance or designee performs a final review of the paperwork to detect clerical errors, errors in methods, units, factors, etc. This final review may also involve transcription of information from a copy to the original survey form.

- The attestation form (or copy of the completed form) must be signed by the Laboratory Director or designee.  (Note: This signature does not have to be obtained prior to reporting results to the PT provider.)

- Results must be submitted to the PT agency on or before the due date, according to the PT provider’s instructions.

- If results are mailed, use a mailing method that ensures receipt will occur before the submission deadline and that can be tracked or verified.

- If results are faxed, confirm that the fax was successful, all pages were faxed, and that the fax was directed to the correct telephone number.  Maintain a record of receipt confirmation.

- If results are submitted electronically, verify that results are correct before final submission.

- For CAP surveys, the QA Manager or other designated individual(s) should review receipt and accuracy of survey results on their website (www.cap.org). Incomplete transmissions or incorrect data entry by CAP can be corrected prior to the survey evaluation.

# RECORDS MAINTENANCE

Records are maintained according to the requirements for Proficiency Testing records published in the Quest Diagnostics Records Management Program.

# RELATED DOCUMENTS

- Quest Diagnostics Procedure for Handling Inappropriate Referral of Proficiency Material or Inter/Intra-laboratory Communication of Proficiency Test Information (QDNQA712)

- Quest Diagnostics Process for Evaluation of Proficiency Test Results (QDNQA716)

- Quest Diagnostics Proficiency Test Handling and Results Submission for Point of Care Testing Sites (QDPS82)

- Quest Diagnostics Records Management Program

# REFERENCES

- CLIA Public Health 42 CFR Part 493 of CLIA Inspector Interpretive Guidelines.

- CAP Laboratory Requirements (www.cap.org).

# REVISION HISTORY

| Version | Date | Revision Purpose | Name |
| --- | --- | --- | --- |
| 1.0 | 3/2007 | New |  |
| 1.1 | 4/2007 | See Revision History in Version 1.1 | Ruthi Breazeale Virginia Sturmfels |
| 2.0 | 7/2008 | See Revision History in Version 2.0 | Patricia Maloney Karen Rupke Ruthi Breazeale |
| 3.0 | 11/2009 | See Revision History in Version 3.0 | Patricia Maloney Ruthi Breazeale Karen Rupke |
| 4.0 | 8/30/2010 | - Added “Pennsylvania” throughout document. - Section 5: Added information about “Test Down” - Section 6: Clarification of PT when multiple instruments are used. - Section 7. New section for Point of Care testing | Patricia Maloney Karen Rupke |
| 4.0 | 10/4/2010 | - Adopting corporate issued version 4 with above revisions. Local revisions as follows: - Section 9. All surveys will be written on the board in the Chemistry Department with the received date, person responsible for the analyzing and the due date to be delivered to the Administrative Director for review. - Section 11. All reported survey raw data will be kept in a labeled folder in the file cabinet in the Quality Assurance Office. - Section 12. All completed surveys will be reviewed by the Administrative Director and then forwarded to the Quality Assurance Specialist for review and submission of results to CAP online. | Nancy Johnson |
| 5.0 | 12/2011 | - Added clarification and simplification to Scope, Exceptions and Responsibility sections - Added requirements for signing of attestation and reflex testing unique order codes - Changed labeling of Appendices / Addendum | Patricia Maloney Karen Rupke Christine Vernusky |
| 4.0A | 5/30/2012 | Adopting corporate issued version 5 with above revisions dated 12/2011. Local revisions as follows: - Included version 4 revisions in revision history table. | Nancy Johnson |

# ADDENDA

| Addendum | File Name | Title |
| --- | --- | --- |
| A | N/A (at end of this file) | Survey Companion Document |
| B | N/A (at end of this file) | QA Manager – PT Delegation |

Testing Department: ____________________					Addendum A

SURVEY COMPANION DOCUMENT

| Survey Agency: | Date / Dept delivered to:: |
| --- | --- |
| Survey Name/Shipment: | Delivered to whom: |
| Date Shipped: (if known) | Date Results due to QA: |
| Date Kit Received: By whom: | Date Sent to Agency: |

QUALITY ASSURANCE/SUPERVISOR

Receipt of Shipment/Integrity of Sample: INITIALS

Received in good condition with no spills, breakage, or missing components? Y N NA

Any evidence of extreme temperature problems? Y N NA

Any evidence of particulate matter, hemolysis, or turbidity? Y N NA

Survey assigned to__________________________________________ Date:____________________

TECHNOLOGIST RESPONSIBILITY

Preparation and Handling of Specimens:

Are all materials identified properly? Y N NA

Are all materials reconstituted with correct diluent/amount per instructions? Y N NA

Were sample prep instructions reviewed and followed? Y N NA

Are all samples mixed well and identified again before testing? Y N NA

Have samples been stored at proper temperature? Y N NA

Comments:

Analytic Process (if testing is done in multiple departments, supply each department with a copy of this form)

If specimens cannot be bar-coded, are they run in proper order? Y N NA

Document unusual occurrences (e.g., PT could not be run with patient samples, failed QC, AMR exceeded, etc.):

SUPERVISOR/TECHNICAL SPECIALIST

Post Analysis:

Are method/instrument codes correct? Y N NA

Have units/decimals/correction factors been checked? Y N NA

Are results entered on the Survey Form correctly? Y N NA

Has secondary check of result transcription been performed? Y N NA

Comments:

##### Department Manager/Supervisor Review: _____________________________________ Date: ______________

QA Manager/Technical Supervisor Review: ____________________________________Date:  ______________

Other comments (continue on back if necessary):

_______________________________________________________________________________________________

Addendum B

## QUALITY ASSURANCE MANAGER

Proficiency Testing Delegation

I authorize the qualified individual(s) listed below to manage proficiency testing activities within this laboratory. I also reapportion the laboratory director responsibilities listed under 493.1445(e)(4) to this (these) individual(s). Specific duties include:

| Activity | Description |
| --- | --- |
|  | Coordinate proficiency testing (PT) activities within the laboratory. |
|  | Ensure that the laboratory is enrolled in an HHS approved PT program for all testing performed and is reviewed on an annual basis. |
|  | Ensure that alternative assessment is performed for tests that do not have an external PT program. |
|  | Ensure that PT samples are tested as required under Subpart H of 42 CFR Part 493. |
|  | Ensure that PT results are reviewed and returned within the timeframes established by the proficiency testing program. |
|  | Provide oversight of Rapid Response Laboratory (RRL) proficiency testing programs. |
|  | Administer the CAP e-LAB program within the laboratory. |
|  | Ensure that all PT results received are reviewed by the appropriate staff to evaluate the laboratory’s performance and to identify any problems that require corrective action. |
|  | Ensure that an approved corrective action plan is followed when any PT result is found to be unacceptable or unsatisfactory. |
|  | Ensure that the laboratory director is aware of the laboratory’s PT performance statistics, PT failures, and corrective actions. |
|  | Ensure that monthly summary statistics regarding PT performance and non-conformances are furnished to National Quality Assurance staff according to NQA established guidelines. |
|  | Assist the Laboratory Director with his/her formal response to Proficiency Test Exception Summaries (PTES) or other notifications of unsuccessful PT performance. (Note: the actual written response must come from the Laboratory Director.) |
|  | Ensure that the Laboratory Director notifies NQA, Medical Regulatory Affairs and Legal Operations whenever 2 of 3 consecutive proficiency testing events are unsuccessful or to report identification of any PT violation. |
|  | Ensures annual PT compliance training is conducted and documented in all appropriate departments as required. |

Laboratory Name and Location: _________________________________________

| Name of Individual | Activity |
| --- | --- |

| Laboratory Director Signature | Date |
| --- | --- |

Appropriate personnel include any staff involved in handling any aspect of the pre-analytic, analytic, and post-analytic laboratory processes that receive or log-in samples, test or report results of proficiency testing samples or refer samples for testing as well as departments that communicate with clients (e.g., Specimen Processing, Technical Operations, Referral Testing and Client Services & Solutions personnel).

CMS has ruled that IHC staining is a process, not an analysis and so the process of staining can be sent out but the interpretation, which is the analysis, must be done where the patient interpretations are routinely done.  The same rule does not apply to FISH probes. CMS has ruled that FISH probe application is analysis and therefore cannot be referred.