Dimension Vista Calibration List

Requested by: _____________						Department: Chemistry

| Date Calibration Request posted | Date Calibration performed | Vista 1 Analytes | Vista 2 Analytes | Vista 3 Analytes (SGAH only) | Recalibration * Current Lot, Include Lot # | New Lot / Correlation, Include Lot # | Tech Code |
| --- | --- | --- | --- | --- | --- | --- | --- |

*Recalibration of current lot must be done within 48 hours

Comments:  ________________________________________________________________________

__________________________________________________________________________________

__________________________________________________________________________________

DIMENSION VISTA CALIBRATION LOG

Analyte ________          Tech ______           Date _________ __

Room Temperature _________ (Range 22-28°C)   Acceptable Y/N _________

(Verify RM TEMP prior to reconstituting calibrators.  It may be necessary to place water and calibrators in a controlled temperature environment ie. Micro Hood, Styrofoam Container, if the RM TEMP is not within the acceptable range.  Document the Room Temperature above)

Reason for calibration (Check one)

New Lot                   Instrument Related            QC Problem                 Calibration Due

- Linear Slope Range: 		0.97 – 1.03

- Non-Linear Slope Range: 	0.95 – 1.05

- Enzymes Slope Range: 	0.90 – 1.10

| Calibrator Lot/Exp | QC Lot# | QC Exp. Date | Old Lot Result | New Lot Result | Expected Mean +/- 2SD | MEM QC SD +/- | Accept Y/N |
| --- | --- | --- | --- | --- | --- | --- | --- |

NEW LOT CORRELATION

| Specimen | Old Lot # |  | Average | New Lot # |  | Average |
| --- | --- | --- | --- | --- | --- | --- |
|  | Result 1 | Result 2 |  | Result 1 | Result 2 |  |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |

% DIFF FORMULA: [(High Result-Low Result)/High Result]*100

| Specimen | % Diff | TEA* | Acceptable Y/N |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

*See current TEA table for acceptable TEA.

FOR REVIEWER USE ONLY Acceptable? Yes No QC Needs Adjustment? Yes No Comment: ___________________________________________________________________________ Reviewed by _______________________Date____________________________________

(Attach Calibration Data to Reverse)