Technical SOP

| Title | Bilirubin by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | Ruben Layug | Date: 09/10/2012 |
| Owner | David Rajamonickam | Date: 02/21/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	7

7.	Equipment And Supplies	10

8.	Procedure	10

9.	Calculations	12

10.	Reporting Results And Repeat Criteria	12

11.	Expected Values	13

12.	Clinical Significance	14

13.	Procedure Notes	14

14.	Limitations Of Method	14

15.	Safety	15

16.	Related Documents	16

17.	References	16

18.	Revision History	16

19.	Addenda	17

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Bilirubin, Direct | Dimension Vista® System | DBIL |
| Bilirubin, Indirect |  | BILI |
| Bilirubin, Total |  | TBIL |

Synonyms/Abbreviations

BILI, TBIL, DBIL

Department

Chemistry

# ANALYTICAL PRINCIPLE

  - Total Bilirubin

Diazotized sulfanilic acid is formed by combining sodium nitrite and sulfanilic acid at low pH. Bilirubin (unconjugated) in the sample is solubilized by dilution in a mixture of caffeine/benzoate/acetate/EDTA. Upon addition of the diazotized sulfanilic acid, the solubilized bilirubin including conjugated bilirubins (mono and diglucoronides) and the delta form (biliprotein-bilirubin covalently bound to albumin) is converted to diazo-bilirubin, a red chromophore representing the total bilirubin which absorbs at 540 nm and is measured using a bichromatic (540, 700 nm) endpoint technique. A sample blank correction is used.

Solubilized bilirubin + Diazotized sulfanilic acid -------------------------->  Red chromophore

(absorbs at 540 nm )

  - Direct Bilirubin

Diazotized sulfanilic acid is formed by combining sodium nitrite and sulfanilic acid at low pH. The sample is diluted in 0.5M HCl. A sample blank reading is taken to eliminate interference from non-bilirubin pigments. Upon addition of the diazotized sulfanilic acid, the conjugated bilirubin is converted to diazo-bilirubin, a red chromophore which absorbs at 540 nm and is measured using a bichromatic (540, 700 nm) endpoint technique.

Conjugated bilirubin + Diazotized sulfanilic acid  -------------------------->  Red chromophore

(absorbs at 540 nm)

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing serum and plasma may be used for samples to be analyzed by this method. |
| Special Collection Procedures | Allow Red Top or SST to clot completely prior to centrifugation. |
| Other | N/A |

  - Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Plasma (Heparin) |  |
| -Other Acceptable | Serum, Cord Blood, Plasma (EDTA) |  |
| Collection Container | Plasma: Green top tube, Purple top tube Serum: Red top tube, Serum separator tube (SST) |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Plastic vial or spun barrier tube at room temperature, protect from light. |  |
| Storage Requirements | Protect from light |  |
| Stability Requirements | Room Temperature 22-28C: | 8 hours |
|  | Refrigerated: 2-8˚C | 5 days |
|  | Frozen: -20C to -70C | 3 months (Direct Bilirubin) 6 months (Total Bilirubin) |
|  | Instrument on board aliquot stability (T-BIL ONLY) | 2 hours (T-BIL ONLY) Not Applicable for D-BIL, for D-BIL use a fresh sample. |
| Timing Considerations | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Gross hemolysis. Reject sample and request a recollection. Credit the test with the appropriate LIS English text code explanation specimen markedly hemolyzed. |  |
| Other Considerations | Allow Red Top or SST to clot completely prior to centrifugation. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| Total Bilirubin | Siemens, Flex® reagent cartridge, Cat. No. K1167 |
| Direct Bilirubin | Siemens, Flex® reagent cartridge, Cat. No. K2125 |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | Total Bilirubin |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1 - 8 have been entered by the instrument, they are stable for 5 days. - Diazotized sulfanilic acid in wells 9 – 10 prepared by the automatic addition of sodium nitrite from wells 11 – 12 is stable for 15 days. |
| Preparation | All reagents are liquid and ready to use. |

| Reagent | Direct Bilirubin |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1 - 4 have been entered by the instrument, they are stable for 2 days. - Once wells 5 - 6 have been entered by the instrument, they are stable for 30 days. - Once wells 7 - 8 have been entered by the instrument, they are stable for 8 days. - Sulfanilic acid in wells 9 – 12 is used immediately to prepare the diazo reagent in wells 1 – 4. |
| Preparation | All reagents are liquid and ready to use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| BILI CAL | Siemens Dimension Vista®, Cat. No. KC210 |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | BILI CAL |
| --- | --- |
| Preparation | Calibrator is ready for use. No preparation is required. |
| Storage/Stability | - Store at 2 - 8° C - Unopened calibrator is stable until expiration date stamped on the box. - Opened Calibrator: once the stopper of the vial is punctured, assigned values are stable for 14 days when stored on board the Dimension Vista System. - Opened Calibrator: once cap is removed, assigned values are stable for 14 days when recapped immediately after use and stored at 2-8° C. Do not use this vial on board the instrument. Product exposure to light may cause significant decrease in bilirubin concentration. |

5.3	Calibration Parameter

| Criteria | Special Notations |
| --- | --- |
| Reference Material | BILI CAL |
| Assay Range | Total Bilirubin: 0.1 – 25.0 mg/dL Direct Bilirubin: 0.05 – 16.0 mg/dL |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in mg/dL |
| Frequency | - Every new reagent cartridge lot. - Every 90 days for any one lot - When major maintenance is performed on the analyzer. - When control data indicates a significant shift in assay. |
| Calibration Scheme | 2 levels, n = 5 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| LiquichekTM Unassayed Chemistry Control Levels 1 and 2 | Bio-Rad Laboratories Cat. No 691 and 692 |
| LiquichekTM Pediatric Control Level 2 | Bio-Rad Laboratories Cat. No. 355 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Liquichek Unassayed Chemistry Controls, Level 1 and 2 |
| --- | --- |
| Preparation | Allow the frozen control to stand at room temperature (22-28C) until completely thawed. Swirl the contents gently to ensure homogeneity. (Do not use a mechanical mixer) - Use immediately. After each use, promptly replace the stopper and return to 2-8C storage. |
| Storage/Stability | Once the control is thawed, bilirubin will be stable for 6 days at 2-8C. Unthawed controls are stable until the expiration date at -20 to -70oC. |

| Control | Liquichek Pediatric Control, Level 2 |
| --- | --- |
| Preparation | Allow the frozen control to stand at room temperature (22-28C) until completely thawed. Swirl the contents gently to ensure homogeneity. (Do not use a mechanical mixer) - Use immediately. After each use, promptly replace the stopper and return to 2-8C storage. |
| Storage/Stability | Once the control is thawed, bilirubin will be stable for 14 days at 2-8C. Unthawed controls are stable until the expiration date at -20 to -70oC. |

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| - 1 | Runs where QC rules are violated will be reviewed and the appropriate corrective action taken. |  |
| - 2 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. Patient samples in failed analytical runs must be reanalyzed and documented. - Corrective action documentation must include the following: the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated further, prior to reporting of any patient results. Courses of action might include: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| - 3 | All QC run on each shift must be attached to the run summary sheet. The Shift Lead tech will review and document all corrected reports on corrective report log. |  |
| - 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CV s are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SD’s and CV’s are sent to Bio-Rad for peer group review. |  |

6.5	Review Patient Data

Technologist must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

TBIL Flex® reagent cartridge Cat. No. K1167 and DBIL Flex® reagent cartridge Cat. No. K2125 is required to perform this test.

Bilirubin is performed on the Dimension Vista® System after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included as a Related Document.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

| Test Conditions |  |
| --- | --- |
| TBIL / DBIL |  |
| Sample Volume: | 5 µL / 5 µL |
| Reagent 1 Volume: | 125 µL / 12.5 µL |
| Reagent 2 Volume: | 23.5 µL / 25 µL |
| Reaction Time: | 7.5 minutes / 4.9 minutes |
| Test Temperature: | 37° C |
| Wavelength: | 540 & 700 nm |
| Type of measurement: | Bichromatic endpoint |

# CALCULATIONS

Indirect (unconjugated) Bilirubin= Total Bilirubin-Direct (conjugated) Bilirubin

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.  Instrument reports results up to one decimal point.

10.3	Units of Measure

mg/dL

10.4	Clinically Reportable Range (CRR)

Total Bilirubin: 	0.1 – 100.0 mg/dL

Direct Bilirubin: 	0.05 – 64.0 mg/dL

Indirect Bilirubin = TBIL-DBIL

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

Direct Bilirubin:

| IF the result is … | THEN… |
| --- | --- |
| < 0.05 mg/dL | Assure there is sufficient sample devoid of bubbles, cellular debris, and/or fibrin clots. Report as: < 0.05 mg/dL |
|  16.0 mg/dL | On Board Automated Dilution: Results  16.0 mg/dL will automatically have repeat testing performed into the instrument using dilution factor of 4. No multiplication is necessary. |
| > 64.0 mg/dL | If the recommended dilution does not give results within the clinically reportable range, report as: “> 64.0 mg/dL” Bring to the attention of your supervisor prior to releasing result. |

Total Bilirubin:

| IF the result is … | THEN… |
| --- | --- |
| < 0.1 mg/dL | Assure there is sufficient sample devoid of bubbles, cellular debris, and/or fibrin clots. Report as: < 0.1 mg/dL |
|  25.0 mg/dL | On Board Automated Dilution: Results  25.0 mg/dL will automatically have repeat testing performed into the instrument using dilution factor of 4. No multiplication is necessary. |
| > 100.0 mg/dL | If the recommended dilution does not give results within the clinically reportable range, report as:”>100.0mg/dL” Bring to the attention of your supervisor prior to releasing result. |

# EXPECTED VALUES

11.1	Reference Ranges

Total Bilirubin      0.2 – 1.0 mg/dL

Direct Bilirubin     0.0 – 0.2 mg/dL

Indirect Bilirubin    N/A (Calculated Value)

  - Priority 1 and 2

Total Bilirubin

< 2 months                     15.0 mg/dL

11.3	Priority 1 & 2 callback processing

| Inpatient | Technical laboratory personnel are to call all critical (panic) results on all inpatients 24 hours a day, seven days a week, as soon as they are resulted. Have the person receiving verbal results to repeat back those results and document the call in a call screen, to include the nurse’s or doctor first and last name. |
| --- | --- |
| Outpatient | The technical laboratory personnel are to notify the client service Department for all outpatient critical value and document on the PMSG screen on who they notify. |

11.4      Priority 3 Limit

None established.

# CLINICAL SIGNIFICANCE

Measurements of bilirubin are used in the diagnosis and treatment of liver, hemolytic hematological and metabolic disorders, including hepatitis and gall bladder disease.

There are at least four distinct bilirubin species that make up the total bilirubin in serum.

The direct reacting species are mono-and diconjugated bilirubin (β- and γ-bilirubin) and the delta fraction (δ-bilirubin), which is tightly bound to albumin. Unconjugated bilirubin (α-bilirubin) is water-insoluble and reacts only after addition of an accelerator such as caffeine.

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the operator of specific malfunctions. Any report slip containing such error messages should be held for follow-up.  Refer to you Dimension Vista Operator’s Guide.

The expected maximum observed standard deviations for repeatability using n = 5 replicates at the following bilirubin concentrations are:

TBIL Concentration		Acceptable S.D. Maximum

1.4 mg/dL				0.2 mg/dL

18.3 mg/dL				1.3 mg/dL

DBIL Concentration	         	Acceptable S.D. Maximum

1.0 mg/dL				0.1 mg/dL

14.8 mg/dL				0.9 mg/dL

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

Total Bilirubin: 	0.1 – 25.0 mg/dL

Direct Bilirubin: 	0.05 – 16.0 mg/dL

14.2	Precision

| Material | Mean mg/dL | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| TBIL, Serum Pool | 1.0 | 0.03 (2.7) | 0.03 (3.3) |
| TBIL, MAS bilirubin 3 | 19.3 | 0.31 (1.6) | 0.56 (2.9) |
| DBIL, Serum Pool 1 | 0.4 | 0.02 (5.7) | 0.02 (5.7) |
| DBIL, Serum Pool 2 | 13.8 | 0.22 (1.6) | 0.37 (2.7) |

14.3	Interfering Substances

HIL Interference:

The TBIL and DBIL method was evaluated for interference according to CLSI/NCCLS EP7-A2.   Bias, defined as the difference between the control sample (does not contain interferent) and the test sample (contains interferent), is shown in the table below. Bias exceeding 10% is considered “interference”.

| Substance tested | Substance Concentration | TBIL mg/dL | Bias % |
| --- | --- | --- | --- |
| Hemoglobin (hemolysate) | 1000 mg/dL | 1.1, 25 | <10 |
| Lipemia Intralipid® | 800 mg/dL 1000 mg/dL 3000 mg/dL | 1.1 1.2 1.4, 21 | 10 20 40, <10 |

| Substance tested | Substance Concentration | DBIL mg/dL | Bias % |
| --- | --- | --- | --- |
| Hemoglobin (hemolysate) | 50 mg/dL | <16 |  |
| Lipemia Intralipid® | 1000 mg/dL 3000 mg/dL | 0.4 0.4, 5.1, 14.1 | <10 50, <10 |

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not available

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert TBIL Flex® Reagent Cartridge K1167

- Current package insert DBIL Flex® Reagent Cartridge K2125

# REFERENCES

  - Ghoshal, Amit K. and Soldin, Steven J., Evaluation of the Dade Behring Dimension® RxL: Integrated chemistry system-pediatric reference ranges. Clinica Chimica Acta 2003; 331:144.

  - Package Insert, TBIL Flex® Reagent Cartridge K1167, Siemens Healthcare Diagnostics Inc., 04/14/2008.

  - Package Insert, DBIL Flex® Reagent Cartridge K2125, Siemens Healthcare Diagnostics Inc., 04/14/2008.

  - Package Insert, BILI CAL, Siemens Healthcare Diagnostics Inc., 03/2008.

  - Package Insert, Unassayed Liquichek Chemistry Controls, Bio-Rad Laboratories, 10/2010.

  - Package Insert, Liquichek Pediatric Control, Bio-Rad Laboratories, 10/2011.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 2-21-14 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 000 | 2-21-14 | 6.4 | Updated Tolerance Limits | David Rajamonickam | Wendell McMillan |
| 000 | 2-21-14 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |
| 000 | 2-21-14 | 1 | Test Information Local Code Updated | David Rajamonickam | Wendell McMillan |

# ADDENDA

None