Technical SOP

| Title | N-terminal Pro-Brain Natriuretic Peptide by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | Ruben Layug | Date: 9/26/2012 |
| Owner | David Rajamonickam | Date: 3/17/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	7

7.	Equipment And Supplies	9

8.	Procedure	9

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	11

11.	Expected Values	11

12.	Clinical Significance	12

13.	Procedure Notes	12

14.	Limitations Of Method	13

15.	Safety	14

16.	Related Documents	14

17.	References	14

18.	Revision History	15

19.	Addenda	15

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| N-terminal Pro-Brain Natriuretic Peptide | Dimension Vista® System | PBNP |

Synonyms/Abbreviations

PBNP, Pro-BNP

Department

Chemistry

# ANALYTICAL PRINCIPLE

The PBNP method is a one-step sandwich chemiluminescent immunoassay based on LOCI® technology. LOCI® reagents include two synthetic bead reagents and a biotinylated monoclonal antibody fragment which recognize an epitope located in the N-terminal part of proBNP. The first bead reagent (Sensibeads) is coated with streptavidin and contains photosensitive dye. The second bead reagent (Chemibeads) is coated with a second antibody specific for a second independent epitope on NT-proBNP and contains chemiluminescent dye. Sample is incubated with Chemibeads and biotinylated antibody to form a particle/NT-proBNP/biotinylated antibody sandwich. Sensibeads then are added and bind to the biotin to form a bead-aggregated immunocomplex. Illumination of the complex by light at 680 nm generates singlet oxygen from Sensibeads, which diffuses to the Chemibeads and triggers a chemiluminescent reaction. The resulting chemiluminescent signal is measured at 612 nm and is directly proportional to the concentration of NT-proBNP in the sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing serum and plasma may be used for samples to be analyzed by this method. |
| Special Collection Procedures | Allow Red Top or SST to clot completely prior to centrifugation. |
| Other | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Plasma (Heparin) |  |
| -Other Acceptable | Serum |  |
| Collection Container | Plasma: Green top tube Serum: Red top tube, Serum separator tube (SST) |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Collection container or Plastic vial at room temperature |  |
| Stability & Storage Requirements | Room Temperature 22-28 °C: | 8 hours |
|  | Refrigerated: 2-8˚C | 3 days |
|  | Frozen: -20C to -70C | 12 months |
|  | Instrument on board aliquot stability | 2 hours |
| Timing Considerations | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Gross hemolysis. Reject sample and request a recollection. Credit the test with the appropriate LIS English text code explanation specimen markedly hemolyzed. |  |
| Other Considerations | Allow Red Top or SST to clot completely prior to centrifugation. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| PBNP | Siemens, Flex® reagent cartridge, Cat. No. K6423A |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | PBNP |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1 - 12 have been entered by the instrument, they are stable for 5 days. |
| Preparation | All reagents are liquid and ready to use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| PBNP CAL | Siemens Healthcare Diagnostics, Cat. No. KC676A |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | PBNP Calibrator |
| --- | --- |
| Preparation | Before use, thaw at room temperature (22-28 °C), swirl and invert gently to mix. Do not use glass pipettes when transferring calibrators to sample cups. |
| Storage/Stability | - Store at -20C to -70C - Unopened calibrator (frozen) is stable until expiration date stamped on the box. - Unopened calibrator (thawed) is stable for 5 dayswhen stored at 2-8°C - Opened Calibrator: once the stopper of the vial is punctured, assigned values are stable for 5 dayswhen stored on board the Dimension Vista System. - Opened Calibrator: once cap is removed, assigned values are stable for 5 dayswhen recapped immediately after use and stored at 2-8°C. Do not use this vial on board the instrument. |

5.3	Calibration Parameter

| Criteria | Special Notations |
| --- | --- |
| Reference Material | PBNP Calibrator |
| Assay Range | 0.0 – 36,750 pg/mL |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in pg/mL |
| Frequency | - Every new reagent cartridge lot. - Every 30 days for any one lot - When major maintenance is performed on the analyzer. - When control data indicates a significant shift in assay. |
| Calibration Scheme | 5 levels, n = 3 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| LiquichekTM Cardiac Markers Plus Control Levels 1, 2 and 3 | Bio-Rad Laboratories Cat # 181, 182 and 183 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Liquichek Cardiac Markers Plus Controls, Level 1, 2 and 3 |
| --- | --- |
| Preparation | Allow the frozen control to thaw at room temperature (22-28 °C) until completely thawed. Swirl the contents gently to ensure homogeneity. (Do not use a mechanical mixer) - Use immediately. After each use, promptly replace the stopper and return to 2-8C storage. |
| Storage/Stability | Unthawed controls are stable until the expiration date at -20 to -70o C. Once the control is thawed and opened, NT-proBNP will be stable for 10 days when stored tightly capped at 2-8 oC. |

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| - 1 | Runs where QC rules are violated will be reviewed and the appropriate corrective action taken. |  |
| - 2 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. Patient samples in failed analytical runs must be reanalyzed and documented. - Corrective action documentation must include the following: the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated further, prior to reporting of any patient results. Courses of action might include: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| - 3 | All QC run on each shift must be attached to the run summary sheet. The Shift Lead tech will review and document all corrected reports on corrective report log. |  |
| - 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CV s are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SD’s and CV’s are sent to Bio-Rad for peer group review. |  |

6.5	Review Patient Data

Technologist must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

NT-proBNP Flex reagent, Cat. No. K6423A is required to perform this test.

NT-proBNP is performed on the Dimension Vista® System after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included as a Related Document.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

| Test Conditions |  |
| --- | --- |
| Sample Volume: | 8 µL |
| Chemibead Reagent Volume: | 20 µL |
| Biotinylated Antibody Volume: | 20 µL |
| Sensibead Reagent Volume: | 15 µL |
| Assay Buffer Volume: | 100 µL |
| Reaction Time: | 37.0 °C |
| Test Temperature: | 10 minutes |
| Wavelength: | 680 and 612 nm |
| Type of measurement: | Chemiluminescence |

# CALCULATIONS

The instrument automatically calculates the concentration of NT-proBNP in pg/mL.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.

10.3	Units of Measure

pg/mL

10.4	Clinically Reportable Range (CRR)

5 - 35,000 pg/mL

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

| IF the result is … | THEN… |
| --- | --- |
| < 5 | Report as < 5 |
|  | On Board Automated Dilution: Not recommended for this method |
| ≥ 35,000 | - Report as > 35,000 |

# EXPECTED VALUES

11.1	Reference Ranges

<75 Years Old:      125 pg/mL

>75 Years Old:      450 pg/mL

11.2	Priority 1 and 2 Limits

None established.

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

The significance of natriuretic peptides in the control of cardiovascular system function has been established. Initial studies reveal that natriuretic peptides can be used for diagnosing clinical problems associated with left ventricular dysfunction. The following natriuretic peptides have been reported in the literature: atrial natriuretic peptide (ANP), brain natriuretic peptide (BNP), and C–type natriuretic peptide (CNP).  ANP and BNP have natriuretic and diuretic properties. As antagonists of the renin-angiotensin-aldosterone system, they influence the electrolyte and fluid balance in an organism. In subjects with left ventricular dysfunction, serum and plasma concentrations of BNP increase, as do the concentrations of the biologically inactive prohormone, proBNP. ProBNP consists of 108 amino acids. It is secreted mainly by the left ventricle of the heart and, in this process, is cleaved into physiologically active BNP (77 – 108), and the N-terminal fragment NT-proBNP (1 – 76).  Studies indicate that NT-proBNP can be used in diagnostic and prognostic applications. The concentration of NT-proBNP in plasma indicates the prognosis for the left ventricular dysfunction. It is also useful in assigning symptoms to cardiac or non-cardiac causes.  Determination of NT-proBNP helps to identify subjects with left ventricular dysfunction. Changes in NT proBNP concentration can be used to evaluate the success of treatment in patients with left ventricular dysfunction. There are indications that NT-proBNP, due to its functions, is suitable for use in assessing vascular remodeling, and therefore contributes to the establishment of individualized rehabilitation procedures.  Fisher et al. found that congestive heart failure patients with elevated NT-proBNP values had a one year mortality rate of 53% compared to 11% in patients with lower values. In the Gusto IV study which involved more than 6800 patients it was shown that NT-proBNP was the strongest independent predictor of one year mortality in patients with acute coronary syndrome.  The European Society of Cardiology Task Force for the Diagnosis and Treatment of Chronic Heart Failure recommends in their guidelines that natriuretic peptides including NT-proBNP “may be most useful clinically as a rule out test due to consistent and very high negative predictive values.”  NT-proBNP levels are increased in patients with unstable angina and following myocardial infarction. Although not diagnostic for these conditions, several studies indicate NT-proBNP measurements provide prognostic information for the short- and long-term risk stratification of patients with unstable angina or myocardial infarction.

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the operator of specific malfunctions. Any report slip containing such error messages should be held for follow-up.  Refer to you Dimension Vista Operator’s Guide.

The expected maximum observed standard deviations for repeatability (within-run precision) using n = 5 replicates at the following analyte concentrations are:

| PBNP concentration | Acceptable SD Maximum |
| --- | --- |
| 125 pg/mL | 6.5 pg/mL |
| 450 pg/mL | 21.4 pg/mL |

A system malfunction may exist if the acceptable SD maximum is exceeded.

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

5 – 35000 pg/mL

14.2	Precision

| Material | Mean pg/mL | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| Serum Pool 1 | 120 | 1.72 (1.4) | 2.06 (1.7) |
| Serum Pool 2 | 438 | 4.75 (1.1) | 4.75 (1.1) |
| Serum Pool 3 | 880 | 9.05 (1.0) | 12.24 (1.4) |
| Serum Pool 4 | 5075 | 69.06 (1.4) | 91.97 (1.8) |
| Internal QC | 32633 | 763.39 (2.3) | 1140.27 (3.5) |

14.3	Interfering Substances

HIL interference

The PBNP method was evaluated for interference to hemoglobin, icterus and lipemia (HIL) according to 
CLSI/ NCCLS EP7-A2.31  Bias is the difference in the results between the control sample (without the interferent) and the test sample (contains the interferent) expressed in percent. Bias exceeding 10% is considered interference.

| Substance Tested | Substance Concentration | NT-proBNP pg/mL | Bias* % |
| --- | --- | --- | --- |
| Hemoglobin (hemolysate) | Hemoglobin (monomer) 1000 mg/dL [mmol/L]g | 134 2044 | < 10 |
| Bilirubin (unconjugated) | 60 mg/dL [1026 µmol/L] | 132 2019 | < 10 |
| Bilirubin (conjugated) | 60 mg/dL [1026 µmol/L] | 134 2066 | < 10 |
| Lipemia (Intralipid®) | 3000 mg/dL [33.9 mmol/L] | 131 2028 | < 10 |

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not available

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert CTNI Flex® Reagent Cartridge K6421

# REFERENCES

  - Package Insert, ProBNP Flex® Reagent Cartridge , Siemens Healthcare Diagnostics Inc., 03/19/2008.

  - Package Insert, ProBNP CAL, Siemens Healthcare Diagnostics Inc., 03/2008.

  - Package Insert, Liquichek Cardiac Markers PlusControl, Bio-Rad Laboratories, 11/2011.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 3-17-14 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 000 | 3-17-14 | 6.4 | Updated Tolerance Limits | David Rajamonickam | Wendell McMillan |
| 000 | 3-17-14 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |

# ADDENDA

None