Technical SOP

| Title | Vitamin B12 by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | Ruben Layug | Date: 4/18/2013 |
| Owner | David Rajamonickam | Date: 3/17/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	2

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	6

7.	Equipment And Supplies	9

8.	Procedure	9

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	12

13.	Procedure Notes	12

14.	Limitations Of Method	12

15.	Safety	15

16.	Related Documents	15

17.	References	15

18.	Revision History	16

19.	Addenda	16

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Vitamin B12 | Dimension Vista® System | VB12 |

Synonyms/Abbreviations

B12 or VB12

Department

Chemistry

# ANALYTICAL PRINCIPLE

The VB12 method is an in vitro diagnostic test for the quantitative measurement of vitamin B12 (B12) in human serum and plasma on the Dimension Vista® System. Measurements of vitamin B12 may be used in the diagnosis of vitamin B12 deficiency.

The vitamin B12 method (VB12) is a homogeneous, competitive chemiluminescent immunoassay based on LOCI® technology. LOCI® reagents include two synthetic bead reagents and biotinylated intrinsic factor (IF). The first bead reagent (Chemibead) is coated with a B12 derivative and contains a chemiluminescent dye. The second bead reagent (Sensibead) is coated with streptavidin and contains photosensitive dye. The patient sample is pretreated with sodium hydroxide (NaOH) and dithioerythritol (DTE) to release the serum B12 from its carrier proteins. Potassium cyanide (KCN) is added to convert all the forms of B12 into a single, cyanocobalamin form, and dicyanocobinamide is added to keep the B12 from rebinding with the carrier proteins. After the sample pretreatment, the biotinylated IF and chemibead reagents are added sequentially to the reaction vessel. Vitamin B12 from the sample competes with the B12-chemibead for a limited amount of biotinylated IF. Sensibead reagent is then added and binds to the biotin to form bead pair immunocomplexes. Illumination of the complex at 680 nm generates singlet oxygen from the Sensibeads which diffuses to the Chemibeads triggering a chemiluminescent reaction. The resulting signal is measured at 612 nm and is an inverse function of the concentration of vitamin B12 in the sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing serum may be used for samples to be analyzed by this method. |
| Special Collection Procedures | Allow Red Top or SST to clot completely prior to centrifugation. |
| Other | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Plasma (Heparin) |  |
| -Other Acceptable | Serum |  |
| Collection Container | Plasma: Green top tube Serum: Red top tube, Serum separator tube (SST) |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Collection container at room temperature. |  |
| Stability & Storage Requirements | Room Temperature 22-28C: | 4 days |
|  | Refrigerated: 2-8˚C | 7 days |
|  | Frozen: -20C to -70C | 6 months |
|  | Instrument on board aliquot stability | 2 hours |
| Timing Considerations | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Gross hemolysis. Reject sample and request a recollection. Credit the test with the appropriate LIS English text code explanation specimen markedly hemolyzed. |  |
| Other Considerations | Allow Red Top or SST to clot completely prior to centrifugation. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| VITAMIN B12 | Siemens, Flex® reagent cartridge, Cat. No. K6442A |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | VITAMIN B12 |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1 - 8 have been entered by the instrument, they are stable for 3 days. - Once wells 9 - 12 have been entered by the instrument, they are stable for 14 days. |
| Preparation | All reagents are liquid and ready to use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| VITAMIN B12 CAL | Siemens Dimension Vista®, Cat. No. KC640A |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | VITAMIN B12 CAL |
| --- | --- |
| Preparation | No preparation is required. |
| Storage/Stability | - Store at 2 - 8° C - Unopened calibrator is stable until expiration date stamped on the ampule. - Opened Calibrator: once the ampule is opened, it should be used immediately and any portion not used should be discarded. |

5.3	Calibration Parameter

| Criteria | Special Notations |
| --- | --- |
| Reference Material | VITAMIN B12 CAL |
| Assay Range | 5 – 1000 µg/dL |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in µg/dL |
| Frequency | - Every new reagent cartridge lot. - Every 90 days for any one lot - When major maintenance is performed on the analyzer. - When control data indicates a significant shift in assay. |
| Calibration Scheme | 5 levels, n = 3 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| LiquichekTM Unassayed Chemistry Control Levels 1 and 2 | Bio-Rad Laboratories Cat # 691 and 692 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Liquichek Unassayed Chemistry Controls, Level 1 and 2 |
| --- | --- |
| Preparation | Allow the frozen control to stand at room temperature (22-28C) until completely thawed. Swirl the contents gently to ensure homogeneity. (Do not use a mechanical mixer) - Use immediately. After each use, promptly replace the stopper and return to 2-8C storage. |
| Storage/Stability | Once the control is thawed, all analytes will be stable for 15 days at 2-8C. Unthawed controls are stable until the expiration date at -20 to -70oC. |

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| - 1 | Runs where QC rules are violated will be reviewed and the appropriate corrective action taken. |  |
| - 2 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. Patient samples in failed analytical runs must be reanalyzed and documented. - Corrective action documentation must include the following: the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated further, prior to reporting of any patient results. Courses of action might include: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| - 3 | All QC run on each shift must be attached to the run summary sheet. The Shift Lead tech will review and document all corrected reports on corrective report log. |  |
| - 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CV s are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SD’s and CV’s are sent to Bio-Rad for peer group review. |  |

6.5	Review Patient Data

Technologists must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

VITAMIN B12 Flex® reagent cartridge Cat. No. K6442A  is required to perform this test.

VITAMIN B12 is performed on the Dimension Vista® System after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included as a Related Document.

| 8.1 | Sample Processing |
| --- | --- |
| - 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| - 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| - 3. | The rack moves into the sample server and to the rack positioner. |
| - 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| - 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| - 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| - 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| - 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| - 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| - 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| - 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| - 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

| Test Conditions |  |
| --- | --- |
| Sample Volume (delivered to the vessel) | 12 µL |
| Reagent 1: DTE Reagent | 15 µL |
| Reagent 2: Extractant Reagent | 15 µL |
| Reagent 3: Blocking Neutralizer Reagent | 50 µL |
| Reagent 4: Chemibead Reagent | 15 µL |
| Reagent 5: Sensibead Reagent | 50 µL |
| Temperature | 37 °C |
| Reaction time | 32 minutes |
| Wavelength | Illumination 680 nm, emission 612 nm |
| Type of measurement | Chemiluminescence |

# CALCULATIONS

All calculations are performed by the instrument.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.  Instrument reports results as a whole number.

10.3	Units of Measure

pg/mL

10.4	Clinically Reportable Range (CRR)

60 – 6000 pg/mL

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

| IF the result is … | THEN… |
| --- | --- |
| less than 60 pg/mL | Reported as “less than 60 pg/mL” |
|  2,000 pg/mL | On Board Automated Dilution: Results  2,000 pg/mL will automatically have repeat testing performed into the instrument using dilution factor of 3. No multiplication is necessary. |
| >6,000 pg/mL | If the On Board Automated Dilution does not give results within the clinically reportable range, report as: “>6,000 pg/mL”. |

# EXPECTED VALUES

11.1	Reference Ranges

193 – 986 pg/mL

11.2	Critical Values

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

Vitamin B12, or cobalamin, is an essential vitamin with a molecular mass of 1355 Da that is found in a variety of foods such as fish, shellfish, meats and dairy products. Intrinsic factor (IF), transcobalamin II (TCII) and haptocorrin (HC) are binding proteins necessary for the assimilation, transport and delivery of B12 to the blood and body tissues. Vitamin B12 is primarily stored in the liver and released on demand. The body uses B12 very efficiently, reabsorbing B12 from the small intestine and returning it to the liver so little is excreted and nutritional deficiency is extremely rare.1 Vitamin B12 is necessary for DNA synthesis, normal red blood cell maturation and myelin sheath formation and maintenance. It is a coenzyme in the conversion of methylmalonic acid to succinic acid and in the synthesis of methionine.1

Vitamin B12 deficiency is one of the causes of megaloblastic anemia, a disease in which red blood cells are larger than normal and the ratio of nucleus size to cell cytoplasm is increased. Since folic acid deficiency can also cause megaloblastic anemia, measurement of serum B12 levels is an important part of the differential diagnosis.1 Vitamin B12 deficiency also causes macrocytic anemias which are characterized by abnormal red blood cell maturation and early release from the bone marrow. Pernicious anemia is a macrocytic anemia. In this disease, an absence of IF prevents normal absorption of B12. In both megaloblastic anemia caused by B12 deficiency and pernicious anemia, treatment with B12 is the therapeutic course.1

Vitamin B12 deficiency can also lead to abnormal neurologic and psychiatric symptoms such as ataxia, muscle weakness, dementia, psychosis and mood disturbances.2 Many patients show neurological changes without developing macrocytic anemia. Populations at risk for B12 deficiency include strict vegetarians, the elderly and populations with increased B12 requirements associated with pregnancy, thyrotoxicosis, hemolytic anemia, hemorrhage, malignancy and liver or kidney disease.

Early diagnosis of B12 deficiency is crucial because of the latent nature of this disorder and the risk of irreversible neurological damage. Recent studies suggest that in addition to serum B12 levels, folic acid, methylmalonic acid and homocysteine should be measured to improve the specificity of the diagnosis.  Elevated B12 levels are seen in hematological disorders (chronic myelogenous leukemia, promyelocytic leukemia, polycythemia vera) and in liver disorders (acute hepatitis, cirrhosis, hepatocellular carcinoma).

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The expected maximum observed standard deviations for repeatability (within-run precision) using n = 5 replicates at the following analyte concentrations are:

| B12 concentration | Acceptable SD Maximum |
| --- | --- |
| 200 pg/mL [148 pmol/L] | 35 pg/mL [26 pmol/L] |
| 1000 pg/mL [738 pmol/L] | 60 pg/mL [45 pmol/L] |

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

60 -2000 pg/mL

14.2	Precision

| Material | Mean µg/dL | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| Serum Pool |  |  |  |
| Level 1 | 106.2 | 1.4 (1.4) | 1.5 (1.4) |
| Level 2 | 38.4 | 1.1 (3.0) | 1.4 (3.7) |
| Level 3 | 715.7 | 5.3 (0.7) | 8.7 (1.2) |
| Multiqual Control |  |  |  |
| Level 3 | 228.4 | 1.9 (0.8) | 2.0 (0.9) |

14.3	Interfering Substances

## Hemolysis, Icterus, Lipemia (HIL) Interference

The VB12 method was evaluated for interference according to CLSI/NCCLS EP7-A2.  Bias Is the difference in the results between the control sample (without the interferent) and the test sample (contains the interferent) expressed in percent. Bias exceeding 10% is considered interference.

| Substance Tested | Substance Concentration | Vitamin B12 pg/mL [pmol/L] | Bias* % |
| --- | --- | --- | --- |
| Hemoglobin | Hemoglobin (monomer) |  |  |
| (hemolysate) | 500 mg/dL [0.31 mmol/L] | 200 [148] | 19.8 |
|  |  | 1000 [738] | < 10 |
|  | 300 mg/dL [0.19 mmol/L] | 200 [148] | < 10 |
|  |  | 1000 [738] | < 10 |
| Bilirubin (unconjugated) | 60 mg/dL [1026 µmol/L] | 200 [148] | < 10 |
|  |  | 1000 [738] | < 10 |
| Bilirubin (conjugated) | 60 mg/dL [1026 µmol/L] | 200 [148] | < 10 |
|  |  | 1000 [738] | < 10 |
| Lipemia (Intralipid®) | 3000 mg/dL [33.9 mmol/L] | 200 [148] | < 10 |
|  |  | 1000 [738] | < 10 |

Intralipid® is a registered trademark of Fresenius Kabi AG, Bad Homburg, Germany.

* Analyte results should not be corrected based on this bias.

# Non-Interfering Substances

The following substances do not interfere with the VB12 method when present in serum and plasma at the concentrations indicated. Inaccuracies (biases) due to these substances are less than 10 % at VB12 concentrations of 200 pg/mL [148 pmol/L] and 1000 pg/mL [738 pmol/L].

| Substance | Test Concentration | SI Units |
| --- | --- | --- |
| Acetaminophen | 20 mg/dL | 1324 µmol/L |
| Amikacin | 8 mg/dL | 137 µmol/L |
| Ampicillin | 5.3 mg/dL | 152 µmol/L |
| Ascorbic Acid | 6 mg/dL | 342 µmol/L |
| Biotin | 100 ng/mL | 406 nmol/L |
| Caffeine | 6 mg/dL | 308 µmol/L |
| Carbamazepine | 3 mg/dL | 127 µmol/L |
| Chloramphenicol | 5 mg/dL | 155 µmol/L |
| Chlordiazepoxide | 1 mg/dL | 33.3 µmol/L |
| Chlorpromazine | 0.2 mg/dL | 6.27 µmol/L |
| Cholesterol | 500 mg/dL | 12.9 mmol/L |
| Cimetidine | 2 mg/dL | 79.2 µmol/L |
| Creatinine | 30 mg/dL | 2652 µmol/L |
| Diazepam | 0.5 mg/dL | 17.6 µmol/L |
| Digoxin | 6.1 ng/mL | 7.8 nmol/L |
| Erythromycin | 6 mg/dL | 81.6 µmol/L |
| Ethanol | 400 mg/dL | 86.8 mmol/L |
| Ethosuximide | 25 mg/dL | 1770 µmol/L |
| Furosemide | 6 mg/dL | 181 µmol/L |
| Gentamicin | 1 mg/dL | 21 µmol/L |
| Heparin | 3 U/mL | 3000 U/L |
| Ibuprofen | 50 mg/dL | 2425 µmol/L |
| Immunoglobulin G (IgG) | 5 g/dL | 50 g/L |
| Lidocaine | 1.2 mg/dL | 51.2 µmol/L |
| Lithium | 2.2 mg/dL | 3.2 mmol/L |
| Nicotine | 0.1 mg/dL | 6.2 µmol/L |
| Penicillin G | 25 U/mL | 25,000 U/L |
| Pentobarbital | 8 mg/dL | 354 µmol/L |
| Phenobarbital | 10 mg/dL | 431 µmol/L |
| Phenytoin | 5 mg/dL | 198 µmol/L |
| Primidone | 4 mg/dL | 183 µmol/L |
| Propoxyphene | 0.2 mg/dL | 4.91 µmol/L |
| Protein: Albumin | 6 g/dL | 60 g/L |
| Protein: Total | 12 g/dL | 120 g/L |
| Rheumatoid Factor | 500 IU/mL | 500 IU/mL |
| Salicylic Acid | 60 mg/dL | 4.34 mmol/L |
| Theophylline | 4 mg/dL | 222 µmol/L |
| Triglycerides | 3000 mg/dL | 33.9 mmol/L |
| Urea | 500 mg/dL | 83.3 mmol/L |
| Uric Acid | 20 mg/dL | 1.2 mmol/L |
| Valproic Acid | 50 mg/dL | 3467 µmol/L |
| Vancomycin | 10 mg/dL | 69 µmol/L |

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Quest Diagnostics Records Management Procedure

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert VITAMIN B12 Flex® Reagent Cartridge.

# REFERENCES

  - Package Insert, VITAMIN B12 Flex® Reagent Cartridge, Siemens Healthcare Diagnostics Inc., 03/11/2008.

  - Package Insert, VITAMIN B12 CAL, Siemens Healthcare Diagnostics Inc., 03/2008.

  - Package Insert, Unassayed Liquichek Chemistry Controls, Bio-Rad Laboratories, 10/2010.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 2-28-14 | 1 | Updated Local Code | W. McMillan | R. Levy |
| 000 | 2-28-14 | 6.4 | Updated QC Tolerance Limits | W. McMillan | R. Levy |
| 000 | 2-28-14 | 10.5 | Update Dilution Protocol | W. McMillan | R. Levy |
| 000 | 3-17-14 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 000 | 3-17-14 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |

# ADDENDA

None