Technical SOP

| Title | CardioPhase® high sensitivity CRP by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | David Rajamonickam | Date: 05/13/2014 |
| Owner | David Rajamonickam | Date: 05/13/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	2

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	4

6.	Quality Control	6

7.	Equipment And Supplies	9

8.	Procedure	9

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	12

13.	Procedure Notes	12

14.	Limitations Of Method	12

15.	Safety	15

16.	Related Documents	16

17.	References	16

18.	Revision History	17

19.	Addenda	17

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| CardioPhase® high sensitivity CRP | Dimension Vista® System | CCRP |
| Synonyms/Abbreviations |  |  |
| hsCRP |  |  |
| Department |  |  |
| Chemistry |  |  |

# ANALYTICAL PRINCIPLE

Polystyrene particles coated with monoclonal antibodies specific to human CRP are aggregated when mixed with samples containing CRP. These aggregates scatter a beam of light passed through the sample. The intensity of the scattered light is proportional to the concentration of the respective protein in the sample. The result is evaluated by comparison with a standard of known concentration.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing serum and plasma may be used for samples to be analyzed by this method. As a trough, collect just before the next schedule dose. |
| Special Collection Procedures | Allow Red Top or SST to clot completely prior to centrifugation. |
| Other | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Serum |  |
| -Other Acceptable | Plasma (Heparinized) |  |
| Collection Container | Serum: Red top tube, Serum separator tube (SST) Plasma: Green top tube |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Collection container or Plastic vial at room temperature. |  |
| Stability & Storage Requirements | Room Temperature 22-28C: | 8 hours |
|  | Refrigerated: 2-8˚C | 7 days |
|  | Frozen: -20C to -70C | 8 Months |
|  | Instrument on board aliquot stability | 2 hours |
| Timing Considerations | Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Gross hemolysis. Reject sample and request a recollection. Credit the test with the appropriate LIS English text code explanation specimen markedly hemolyzed. |  |
| Other Considerations | Allow Red Top or SST to clot completely prior to centrifugation. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| hsCRP | Siemens, Flex® reagent cartridge, Cat. No. K7046 |
| PROT2 CAL | Siemens, PROT2 CAL, Cat. No. KC780 |
| System Diluent | Siemens, System Diluent, Cat. No. KS804 |
| N Diluent | Siemens, N Diluent, Cat. No. OUMT05 |
| hsCRP CON L | Siemens, hsCRP CON L, Cat. No. KC755 |
| hsCRP CON H | Siemens, hsCRP CON H, Cat. No. KC757 |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | hsCRP Flex® reagent cartridge, Cat. No. K7046 |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 90 days. - Once wells 1 - 12 have been entered by the instrument, they are stable for 21 days. |
| Preparation | All reagents are liquid and ready to use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| PROT 2 CAL | Siemens Dimension Vista®, Cat. No. KC780 |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | PROT 2 CAL |
| --- | --- |
| Preparation | Calibrator is ready for use. No preparation is required. |
| Storage/Stability | - Store at 2-8° C - Unopened calibrator is stable until expiration date stamped on the box. - Opened Calibrator: once the stopper of the vial is punctured, assigned values are stable for 15 days when stored on board the Dimension Vista System. |

5.3	Calibration Parameters

| Criteria | Special Notations |
| --- | --- |
| Reference Material | PROT 2 CAL |
| Assay Range | 0.016 – 19 mg/dL |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in mg/dL |
| Frequency | - Every new reagent cartridge lot. - Every 45 days for any one lot - When major maintenance is performed on the analyzer. - When control data indicates a significant shift in assay. |
| Calibration Scheme | 7 levels, n = 3 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| PROT2 CON L (Low) | Siemens Dimension Vista® Cat. No. KC785 |
| PROT2 CON H (High) | Siemens Dimension Vista® Cat. No. KC787 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | PROT2 CON L (Low) - Siemens Dimension Vista® Cat. No. KC785 PROT2 CON H (High) - Siemens Dimension Vista® Cat. No. KC787 |
| --- | --- |
| Preparation | - PROT2 CON L and PROT2 CON H is ready for use. No preparation is required. Invert gently to mix. Avoid vigorous shaking and foam formation. |
| Storage/Stability | - This product will be stable until the expiration date when stored unopened at 2 to -8 C. - Unopened: See Exp. Date on carton/vial. - Opened: Once the stopper is punctured assigned values are stable for 12 days when stored on board the Dimension Vista System. |

6.3	Frequency

Analyze a minimum of two levels of a Quality Control (QC) material with known C-reactive protein concentrations,  PROT2 CON L or H at least once each day of use.

Follow your laboratory internal QC procedures if the results obtained are outside acceptable limits. Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Tolerance Limits |  |
| --- | --- |
| 1. | Runs where QC rules are violated will be reviewed and the appropriate corrective action taken. |
| 2. | Corrective Action: - All rejected runs must be effectively addressed through corrective action. Patient samples in failed analytical runs must be reanalyzed and documented. - Corrective action documentation must include the following: the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated further, prior to reporting of any patient results. Courses of action might include: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |
| 3. | All QC run on each shift must be attached to the run summary sheet. The Shift Lead tech will review and document all corrected reports on corrective report log. |
| 4. | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CV s are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SD’s and CV’s are sent to Bio-Rad for peer group review. |
| Note: | Tolerance limits for SD must not exceed one-third Allowable Error (TEa/3) specifications for this test. Refer to the Total Allowable Error Table on the Quest Diagnostics intranet for the most current TEa. http://questnet1.qdx.com/Business_Groups/Medical/qc/quality_control_sops.htm |

6.5	Review Patient Data

Technologists must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Calibrated Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

hsCRP Flex® reagent cartridge, Cat. No. K7046 is required to perform this test.

hsCRP  is performed on the Dimension Vista® System after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included as a Related Document.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

NOTE:   In the event that the test system becomes inoperable, notify supervision or designee for further direction.  Patient specimens must be stored in a manner that maintains the integrity of the specimen.

| Test Conditions |  |
| --- | --- |
| Initial Sample Dilution | 1:1 (neat) |
| Diluted Sample Volume | 1.37 µL |
| Diluent Volume | 118 µL |
| hsCRP Reagent | 27.3 µL |
| Temperature | 37.0 °C |
| Reaction time | 5 minutes 50 seconds |
| Wavelength | 840 nm |
| Type of Measurement | Nephelometric |

# CALCULATIONS

The instrument automatically calculates the concentration of hsCRP  in mg/dL [mg/L].

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.

10.3	Units of Measure

mg/dL

10.4	Clinically Reportable Range (CRR)

0.016 – 19 mg/dL

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

| IF the result is … | THEN… |
| --- | --- |
| < 0.016 mg/dL | Assure there is sufficient sample devoid of bubbles, cellular debris, and/or fibrin clots. Report as: < 0.016 mg/dL |
|  0.95 mg/dL | On Board Automated Dilution: Results  0.95 µg/mL will automatically have repeat testing performed into the instrument using dilution factor of 20. No multiplication is necessary. |
| > 19.0 mg/dL | Manual Dilution: Dilute with N Diluent and bring the raw data within the AMR. Enter dilution factor as a whole number. Re-assay. Readout is corrected for dilution. Manual Dilution Steps: Dilution Factor 1:100 - Take 10 µL(microliter) of sample and add 990 µL (microliter) of N Diluent - Mix it thoroughly and load the sample manually on the instrument. - Under the manual dilution factor say 100 and submit the sample. Check to see if you get the result within the clinically reportable range and result it. If the recommended dilution does not give results within the clinically reportable range, report as: “> 19.0 mg/dL” Bring to the attention of your supervisor prior to releasing result. |

# EXPECTED VALUES

11.1	Reference Ranges

0.016 – 19 mg/dL

11.2	Priority 1 and 2 Limit(s)

None Established

11.3	Priority 1 & 2 callback processing

| Inpatient | Technical laboratory personnel are to call all critical (panic) results on all inpatients 24 hours a day, seven days a week, as soon as they are resulted. Have the person receiving verbal results to repeat back those results and document the call in a call screen, to include the nurse’s or doctor first and last name. |
| --- | --- |
| Outpatient | The technical laboratory personnel are to notify the client service Department for all outpatient critical value and document on the PMSG screen on who they notify. |

11.4      Priority 3

None established.

# CLINICAL SIGNIFICANCE

CRP is one of the ‘acute-phase’ proteins, the serum or plasma levels of which rise during general, nonspecific response to infectious and non-infectious inflammatory processes. CRP is synthesized in the liver and is normally present as a trace constituent of serum or plasma. In various disease states resulting in tissue injury, infection or acute inflammation, CRP values may rise above normal to 20 to 500 mg/L. As elevated CRP values are always associated with pathological changes, the CRP method provides useful information for the diagnosis, therapy and monitoring of inflammatory processes and associated diseases. Increases in CRP values are non-specific and should not be interpreted without a complete clinical history. Studies have shown that measurement of CRP by high sensitivity methods is a strong independent predictor of risk for future cardiovascular and peripheral vascular disease. High sensitivity CRP measurements have also been shown to add to the predictive value of other markers used to assess the risk of cardiovascular and peripheral vascular disease. Elevated CRP values determined by high sensitivity CRP methods may be indicative of the prognosis of individuals with acute coronary syndromes, and may be useful in the management of such individuals.

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the operator of specific malfunctions. Any report slip containing such error messages should be held for follow-up.  Refer to you Dimension Vista Operator’s Guide.

The expected maximum observed standard deviations (SD) for repeatability (within-run precision) using n = 5 replicates at the following nominal CRP concentrations are:

| CRP Concentration | Acceptable SD Maximum |
| --- | --- |
| 0.10 mg/dL [1 mg/L] 0.30 mg/dL [3 mg/L] | 0.023 mg/dL [0.23 mg/L] 0.057 mg/dL [0.57 mg/L] |

# 14. LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0.016 – 0.95 mg/dL

14.2	Precision

| Material |  | Mean |  | Standard Deviation mg/dL [mg/L] (% CV) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | mg/dL | [mg/L] | Repeatability |  |  |  | Within-Lab |  |  |
| hsCRP CON L |  | 0.115 | [1.15] | 0.006 |  | [0.06] | (4.8) | 0.006 | [0.06] | (5.4) |
| hsCRP CON H |  | 0.346 | [3.46] | 0.014 |  | [0.14] | (4.0) | 0.015 | [0.15] | (4.4) |
| Serum pool |  | 0.239 | [2.39] | 0.013 |  | [0.13] | (5.2) | 0.013 | [0.13] | (5.2) |
| Serum pool |  | 0.650 | [6.50] | 0.033 |  | [0.33] | (5.0) | 0.037 | [0.37] | (5.7) |
| Serum pool |  | 0.772 | [7.72] | 0.036 |  | [0.36] | (4.6) | 0.042 | [0.42] | (5.4) |

14.3	Interfering Substances

HIL Interference

The CardioPhase® hsCRP method was evaluated for interference according to CLSI EP7-A2.  Bias is the difference in the results between the control sample (without the interferent) and the test sample (contains the interferent) expressed in percent. Bias exceeding 10 % is considered interference.

| Substance Tested | Substance Concentration |  |  | CRP mg/dL [mg/L] |  | Bias* % |
| --- | --- | --- | --- | --- | --- | --- |
| Hemoglobin (hemolysate) | 1000 mg/dL | [0.155mmol/] |  | 0.093 [0.93] |  | +1 |
| Bilirubin (unconjugated) | 60 mg/dL | [1026µmol/L] |  | 0.097 [0.97] |  | -3 |
| Bilirubin (conjugated) | 60 mg/dL | [1026µmol/L] |  | 0.097 [0.97] |  | +1 |
| Lipemia | Refer to Specimen Collection and Handling section |  |  |  |  |  |

Non Interfering Substances

The following substances do not interfere with the CardioPhase® hsCRP method when present in serum and plasma at the concentrations indicated. Inaccuracies (biases) due to these substances are less than 10 % at CRP concentration of 0.023 mg/dL to 0.591 mg/dL [0.23 mg/L to 5.91 mg/L].

| Substance |  |  | Test Concentration | SI Units |
| --- | --- | --- | --- | --- |
| Acetaminophen |  |  | 0.025 mg/dL | 1.66 μmol/L |
| Amikacin |  |  | 15 mg/dL | 256 μmol/L |
| Ammonium heparin |  |  | 3 U/mL | 3000 U/L |
| Ampicillin |  |  | 5.3 mg/dL | 152 μmol/L |
| Ascorbic acid |  |  | 5 mg/dL | 227 μmol/L |
| Caffeine |  |  | 6 mg/dL | 308 μmol/L |
| Carbamazepine |  |  | 3 mg/dL | 127 μmol/L |
| Chloramphenicol |  |  | 5 mg/dL | 155 μmol/L |
| Chlordiazepoxide |  |  | 1 mg/dL | 33.3 μmol/L |
| Chlorpromazine |  |  | 0.2 mg/dL | 6.27 μmol/L |
| Cholesterol |  |  | 500 mg/dL | 12.9 mmol/L |
| Cimetidine |  |  | 2 mg/dL | 79.2 μmol/L |
| Creatinine |  |  | 30 mg/dL | 2652 μmol/L |
| Dextran 40 |  |  | 6000 mg/dL | 1500 μmol/L |
| Diazepam |  |  | 0.5 mg/dL | 17.6 μmol/L |
| Digoxin |  |  | 5 ng/mL | 6.15 nmol/L |
| Erythromycin |  |  | 6 mg/dL | 81.6 μmol/L |
| Ethanol |  |  | 400 mg/dL | 86.8 mmol/L |
| Ethosuximide |  |  | 25 mg/dL | 1770 μmol/L |
| Furosemide |  |  | 6 mg/dL | 181 μmol/L |
| Gentamicin |  |  | 12 mg/dL | 251 μmol/L |
| Ibuprofen |  |  | 50 mg/dL | 2425 μmol/L |
| Immunoglobulin G (IgG) |  |  | 5 g/dL | 50 g/L |
| Lidocaine |  |  | 1.2 mg/dL | 51.2 μmol/L |
| Lithium chloride |  |  | 2.3 mg/dL | 3.2 mmol/L |
| Lithium heparin |  |  | 3 U/mL | 3000 U/L |
| Nicotine |  |  | 0.1 mg/dL | 6.2 μmol/L |
| Penicillin G |  |  | 25 U/mL | 25000 U/L |
| Pentobarbital |  |  | 8 mg/dL | 354 μmol/L |
| Phenobarbital |  |  | 10 mg/dL | 431 μmol/L |
| Phenytoin |  |  | 5 mg/dL | 198 μmol/L |
| Primidone |  |  | 4 mg/dL | 183 μmol/L |
| Propoxyphene |  |  | 0.2 mg/dL | 4.91 μmol/L |
| Protein, Albumin |  |  | 6 g/dL | 60 g/L |
| Rheumatoid Factors |  |  | 500 IU/mL | 500 IU/mL |
| Salicylic acid |  |  | 60 mg/dL | 4.34 mmol/L |
| Sodium heparin |  |  | 3 U/mL | 3000 U/L |
| Theophylline |  |  | 4 mg/dL | 222 μmol/L |
| Urea |  |  | 500 mg/dL | 83.3 mmol/L |
| Uric acid |  |  | 20 mg/dL | 1190 μmol/L |
| Valproic acid |  |  | 50 mg/dL | 3467 μmol/L |

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not available.

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification  Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Therapeutic Drug Monitoring, (Chemistry SOP)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert CRP Flex® Reagent Cartridge K7032

# REFERENCES

  - Package Insert, hsCRP Flex® Reagent Cartridge K7046, Siemens Healthcare Diagnostics Inc., 03/2011.

  - Package Insert, PROT2 CAL, Siemens Healthcare Diagnostics Inc., 01/2008.

  - Package Insert, PROT2 Control L (Low), Siemens Healthcare Diagnostics Inc., 01/2008.

  - Package Insert, PROT2 Control H (High), Siemens Healthcare Diagnostics Inc., 01/2008.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |

# ADDENDA

None