Technical SOP

| Title | Hemoglobin A1c by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | Ruben Layug | Date: 8/16/2012 |
| Owner | David Rajamonickam | Date: 9/25/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	7

7.	Equipment And Supplies	9

8.	Procedure	10

9.	Calculations	11

10.	Reporting Results And Repeat Criteria	12

11.	Expected Values	13

12.	Clinical Significance	13

13.	Procedure Notes	13

14.	Limitations Of Method	14

15.	Safety	14

16.	Related Documents	15

17.	References	15

18.	Revision History	15

19.	Addenda	16

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Hemoglobin A1c | Dimension Vista® System | HA1C |
| Estimated Average Glucose | Dimension Vista® System | eAG |

Synonyms/Abbreviations

Glycohemoglobin, Glycosylated Hemoglobin, Glycated Hemoglobin/HA1C, HbA1c, GHB, Estimated Average Glucose/EAG, eAG.

Department

Chemistry

# ANALYTICAL PRINCIPLE

The Dimension Vista HbA1c assay measures both HbA1c and total hemoglobin. The HbA1c measurement is based on a turbidimetric inhibition immunoassay (TINIA) principle, and the measurement of total hemoglobin is based on a modification of the alkaline hematin

reaction. Using the values obtained for each of these two analytes, the relative proportion of the total hemoglobin that is glycated is calculated and reported. Pre-treatment to remove the labile fraction is not necessary as only the Amadori rearranged form of HbA1c is detected. All hemoglobin variants that are glycated at the beta-chain N-terminus and have epitopes identical to that of HbA1c are measured by this assay.

Total Hemoglobin Measurement: A sample of whole blood is added to a reaction vessel containing lysing reagent. This reagent lyses the red blood cells and simultaneously converts the released hemoglobin to a derivative that has a characteristic absorbance spectrum. An aliquot of the lysed whole blood is transferred from the reaction vessel to a cuvette where total hemoglobin concentration is measured at 405 nm and 700 nm.

Whole blood + lysing agent ------->  Released hemoglobin  -------> hemoglobin derivative

(measured at 405 nm)

Hemoglobin A1c Measurement: The same aliquot of the lysed whole blood that is transferred from the reaction vessel to the cuvette for the Hb measurement is also used for the measurement of HbA1c. The cuvette contains an anti-HbA1c antibody in a buffered reagent. Hemoglobin A1c in the sample reacts with anti-HbA1c antibody to form a soluble antigen-antibody complex. A polyhapten reagent containing multiple HbA1c epitopes is then added to this cuvette. The polyhapten reacts with excess (free) anti-HbA1c antibodies to form an insoluble antibody-polyhapten complex. The rate of this reaction is measured

turbidimetrically at 340 nm and blanked at 700 nm and is inversely proportional to the concentration of HbA1c in the sample.

hemoglobin A1c + anti-HbA1c antibody  ----------> hemoglobin A1c-anti-HbA1c antibody complex		anti-HbA1c antibody (excess) + polyhapten  ---------->  Ab/polyhapten complex  (absorbs at 340 nm)

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing plasma may be used for samples to be analyzed by this method. |
| Special Collection Procedures | N/A |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Whole blood treated with EDTA |  |
| -Other Acceptable | None |  |
| Collection Container | Lavender Top Tube |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Collection container or Plastic vial at room temperature |  |
| Stability & Storage Requirements | Room Temperature 22-28C: | 3 days (15 - 25º C) |
|  | Refrigerated: 2-8˚C | 7 days (2 - 8º C) |
|  | Frozen: -20C to -70C | Not recommended. Whole blood specimens should not be frozen. |
| Timing Considerations | No more than one rack should be loaded every 5 minutes. |  |
| Unacceptable Specimens & Actions to Take | Frozen Specimens. Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. Examples: Quantity not sufficient-QNS; Wrong collection-UNAC. Document the request for recollection in the LIS. |  |
| Compromising Physical Characteristics | Samples should be mixed gently by inversion (gently invert the tube ten times) or in a rocker mixer to obtain uniform distribution of the erythrocytes prior to testing. Avoid the formation of foam. |  |
| Other Considerations | Process no more than 6 samples at a time to prevent increased settling of the red cells. |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| HbA1c Kit | Siemens, Flex® reagent cartridge, Cat. No. K3105A Includes HbA1c flex reagent cartridge and calibrator |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | HbA1c Kit |
| --- | --- |
| Container | Reagent cartridge and calibrator (5 levels) |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1 - 12 have been entered by the instrument, they are stable for 5 days. |
| Preparation | All reagents are liquid and ready to use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| HbA1c CAL | Siemens Dimension Vista®, Cat. No. K3105A |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | HbA1c CAL |
| --- | --- |
| Preparation | - Remove the vials from the refrigerator and proceed to the next step. - Remove stopper avoiding loss of lyophilized material. - Volumetrically add 2.0 mL of reagent grade water to each vial. - Replace stopper, and let stand for 5 minutes. Do not invert the vials at this time. - Swirl vials gently for 30 seconds, then gently invert 10 times - Let vials stand for 30 minutes, then invert gently 10 times. |
| Storage/Stability | - Store at 2-8° C - Unopened calibrator is stable until expiration date stamped on the box. - Opened Calibrator: Reconstituted calibrators are stable for 8 hours at 25º C and 48 hours at 2-8º C. |

5.3	Calibration Parameter

| Criteria | Special Notations |
| --- | --- |
| Reference Material | HbA1c CAL |
| Assay Range | See Package Insert for specified assay ranges. |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in g/dL |
| Frequency | - Every new reagent cartridge lot. - Every 30 days for any one lot - When major maintenance is performed on the analyzer. - When control data indicates a significant shift in assay. |
| Calibration Scheme | 2 levels, n = 3 for Hb. 5 levels, n = 3 for HbA1c |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| Streck’s A1c-Cellular | Streck Cat. No. 211124 3-L1, 3-L2 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Streck’s A1c-Cellular Controls, Level 1 and 2 |
| --- | --- |
| Preparation | Before sampling, allow the control to reach room temperature (18-30C) for 15 minutes before use. - After each use, promptly replace the stopper and return to 2-10C storage. |
| Storage/Stability | Opened vials are stable for 30 days at 2-10C. Unopened vials of A1c-Cellular are stable through the expiration date at 2-10C. |

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| 1 | Acceptable ranges for QC are posted near each instrument and are programmed into the instrument Quality Control software system. |  |
| 2 | Rejection criteria: For all Rejection criteria see SOP PHQ.020. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the Westgard Rules |  |
| 3 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. Patient samples in failed analytical runs must be reanalyzed and documented. - Corrective action documentation must include the following rule, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated further, prior to reporting of any patient results. Courses of action might include: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| 4 | All QC run on each shift must be attached to the run summary sheet. Shift Lead tech will review and document all corrected reports on corrective report log |  |
| 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SD’s or CV’s are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SD’s and CV’s are sent to Bio-Rad for peer group review. |  |

6.5	Review Patient Data

Technologist must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–10C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included as a Related Document.

Note: 	HbA1c Kit Cat. No. K3105A is required to perform this test.

Note: 	Hemoglobin A1c is performed on the Dimension Vista® System after the method is 		calibrated (see Reference Material in Calibration section) and Quality Controls are 		acceptable.

Note:	1. HbA1c samples should be run only on the STAT RACK

2. DO NOT USE SMALL SAMPLE CONTAINERS to run HbA1c test. USE

PRIMARY TUBES (or) SAMPLE CUPS. (please see the below image)

3. DO NOT Run any other STAT samples or Controls on the same instrument when

running HbA1c Sample or Controls.

4. No more than one rack should be loaded every 5 min and process no more than 6

Samples at a time to prevent increased settling of the red cells.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

| Test Conditions |  |
| --- | --- |
| Vessel / Cuvette |  |
| Sample Volume: | 1 µL / 8 µL |
| Hemolyzing Reagent Volume: | 100 µL / 0 µL |
| Antibody/Buffer Volume: | 0 µL / 140 µL |
| Polyhapten Volume: | 0 µL / 18 µL |
| Diluent Volume: | 49 µL / 29 µL |
| Reaction Time: | 14.3 minutes |
| Test Temperature: | 37° C |
| Wavelength: | 340, 405 & 383 nm |
| Type of measurement: | Turbidimetric & Colorimetric |

# CALCULATIONS

Results reported in %HbA1c using the following equation:

%HbA1c = 100 x [HbA1c (g/dL) / Hb (g/dL)]

Calibrator values (g/dL) are required for hemoglobin A1c and total hemoglobin. These values are obtained from the Table of Assigned Values provided in the Dimension Vista HbA1c Calibrator Instructions for Use (IFU).

Subsequently the Dimension Vista® System calculates an NGSP standardized %HbA1c result based on the calculation shown below:

NGSP Standardized %HbA1c = [c x (%HbA1c)2] + (b x %HbA1c) + a

All calculations are performed by the instrument.

  - Calculating the Estimated Average Glucose

The Estimate Average Glucose is calculated automatically by the LIS using the 			standardized formula below;

eAG (mg/dl) = 28.7 x A1C -46.7.

The results are expressed (mg/dl) seen on glucose meters.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.  Instrument reports results up to one decimal point.

10.3	Units of Measure

HA1C = Percent (%); eAG = (mg/dL)

10.4	Clinically Reportable Range (CRR)

3.5 – 16.0 %

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (Tea) of the assay.  Refer to Tea listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

| IF the result is … | THEN… |
| --- | --- |
| < 3.5 % | A “Below Assay Range” error will be generated and the % HA1C result will not be printed. Dispense a new aliquot of whole blood and reassay the sample. If the errors persist contact your supervisor and/or Technical Assistance Center. |
| >16.0 | If the recommended results are greater than the CRR, report as: >16.0 %. Bring to the attention of the tech in charge and/or supervisor prior to releasing result. |

*******DO NOT USE DILUTION FACTOR.***

# EXPECTED VALUES

11.1	Reference Ranges

4.2 to 6.3%

| <5.7 % | Decreased risk of diabetes |
| --- | --- |
| 5.7 – 6.0% | Increased risk of diabetes |
| 6.1 – 6.4 % | Higher risk of diabetes |
| = or >6.5 % | Consistent with diabetes |

11.2	Priority 1 and 2 Limit(s)

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

HbA1c refers to the product of a non-enzymatic reaction between glucose and hemoglobin A1. The human erythrocyte is freely permeable to glucose, which can non-enzymatically combine with hemoglobin to form HbA1c. This non-enzymatic reaction between the alpha-amino group of the N-terminal valine of the hemoglobin beta-chain and glucose takes place to form an unstable aldimine or Schiff base intermediate (labile fraction). This reaction is slow and reversible and occurs at a rate that is proportional to the glucose concentration in the blood. The aldimine intermediate subsequently undergoes a non-reversible Amadori rearrangement to form the stable ketoamine 1-glucofrutovaline product. Since the reaction is

driven by the concentration of reactants, the degree of glycosylation (reported as HbA1c relative to the total hemoglobin) is proportional to the average concentration of blood glucose over the circulating life span of hemoglobin in the red cell (approximately 120 days).

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the operator of specific malfunctions. Any report slip containing such error messages should be held for follow-up.  Refer to you Dimension Vista Operator’s Guide.

The expected maximum observed standard deviations for repeatability using n = 5 replicates at the following Hemoglobin A1c concentrations are:

HbA1c Concentration	                  Acceptable S.D. Maximum

5.6 %					0.6 %

9.4 %					0.5 %

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

3.5 – 16.0 %

14.2	Precision

| Material | Mean % | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| Human EDTA Blood, Normal | 5.3 | 0.15 | 0.18 |
| Human EDTA Blood, Normal | 6.2 | 0.15 | 0.17 |
| Human EDTA Blood, High | 8.2 | 0.13 | 0.14 |
| Human EDTA Blood, High | 11.7 | 0.18 | 0.20 |

14.3	Interfering Substances

Various substances other than sugars can form aggregates with hemoglobin and potentially interfere with the assay causing false results. Examples include individuals with opiate addiction, lead poisoning and alcoholism.

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not available

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification  Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management  Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert HbA1c Kit K3105A

# REFERENCES

  - Package Insert, HbA1c Kit K3105A, Siemens Healthcare Diagnostics Inc., 03/04/2011.

  - Package Insert, Streck’s HbA1c Control., Catalog Number 211124 3-L1, 3-L2.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 1/22/2013 | 8 | Minor format change at beginning of section (changed position and “note” to 2 comments) for clarity purposes. | R. Layug | R. Levy |
| 1 | 6/27/2014 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 1 | 6/27/2014 | 6.1 | Control Material Changed to Streck A1c-Cellular | David Rajamonickam | Wendell McMillan |
| 1 | 6/27/2014 | 6.2 | Control Preparation and Storage | David Rajamonickam | Wendell McMillan |
| 1 | 6/27/2014 | 17 | Added Reference package insert and Catalog number | David Rajamonickam | Wendell McMillan |
| 2 | 9-15-2014 | 10.5 | Updated the Repeat Criteria and Resulting table | David Rajamonickam | Wendell McMillan |
| 2 | 9-15-2014 | 8 | Updated instructions regarding the procedure and loading the sample. | David Rajamonickam | Wendell McMillan |

# ADDENDA

None