Technical SOP

| Title | Total Prostate Specific Antigen by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | David Rajamonickam | Date: 05-02-2014 |
| Owner | David Rajamonickam | Date: 05-02-2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	ANALYTICAL PRINCIPLE	2

3.	SPECIMEN REQUIREMENTS	3

4.	REAGENTS	4

5.	CALIBRATORS/STANDARDS	4

6.	QUALITY CONTROL	6

7.	EQUIPMENT and SUPPLIES	9

8.	PROCEDURE	9

9.	CALCULATIONS	10

10.   REPORTING RESULTS AND REPEAT CRITERIA	10

11.   EXPECTED VALUES	11

12.   CLINICAL SIGNIFICANCE	11

13.   PROCEDURE NOTES	12

14.   LIMITATIONS OF METHOD	12

15.   SAFETY	16

16.   RELATED DOCUMENTS	17

17.   REFERENCES	17

18.   REVISION HISTORY	17

19.   ADDENDA	17

# TEST INFORMATION

| Assay | Total Prostate Specific Antigen |
| --- | --- |
| Method | Dimension Vista® System |
| Instrument | Dimension Vista® System |
| Synonyms | TPSA / PSA |
| Department | Chemistry |

| Order Code | Test Name |
| --- | --- |
| PSA | Total Prostate Specific Antigen |

# ANALYTICAL PRINCIPLE

The TPSA method is a homogeneous, sandwich chemiluminescent immunoassay based on LOCI® technology. The LOCI® reagents include two synthetic bead reagents and a biotinylated anti-TPSA monoclonal antibody F(ab′) fragment. The first bead reagent (Chemibeads) is coated with anti-TPSA monoclonal antibody and contains chemiluminescent dye; the second bead reagent (Sensibeads) is coated with streptavidin and contains a photosensitizer dye. Sample is incubated with biotinylated antibody and Chemibeads to form bead/TPSA/biotinylated antibody sandwiches. Sensibeads are then added and bind to the biotin to form bead pair immunocomplexes. When illuminated by light at 680 nm, Sensibeads convert dissolved oxygen in the reaction solution into singlet oxygen form (O2). In the bead pairs, the singlet oxygen diffuses (“channels”) into Chemibeads, triggering a chemiluminescent reaction. The resulting chemiluminescent signal is measured at 612 nm and is directly proportional the concentration of TPSA in the sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Normal procedures for collecting and storing serum and plasma may be used for samples to be analyzed by this method. |
| Special Collection Procedures | Serum and plasma can be collected using recommended procedures for collection of diagnostic blood specimens by venipuncture. |
| Other | Serum and plasma specimens should be separated from cells within 4 hours from the time of collection. |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Serum |  |
| -Other Acceptable | Plasma (Lithium heparin) |  |
| Collection Container | Plasma: Green top tube. Serum: Red top tube, Serum separator tube (SST) |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container & Temperature | Collection container or Plastic vial at room temperature. |  |
| Stability & Storage Requirements | Room Temperature: |  |
|  | Refrigerated: 2-8 °C. |  |
|  | Frozen: -80 °C |  |
| Timing Considerations | Sample should be analyzed within 8 hours. |  |
| Unacceptable Specimens & Actions to Take | Specimens that are unlabeled, improperly labeled, or those that do not meet the stated criteria are unacceptable. Request a recollection and credit the test with the appropriate LIS English text code for “test not performed” message. |  |
| Compromising Physical Characteristics | Gross hemolysis. Reject sample and request a recollection. Credit the test with the appropriate LIS English text code explanation specimen markedly hemolyzed. |  |
| Other Considerations | Allow Red Top or SST to clot completely prior to centrifugation. |  |

NOTE:  Labeling requirements for all reagents, calibrators and controls include: (1) Open date, (2) Substance name, (3) Lot number, (4) Date of preparation, (5) Expiration date, (6) Initials of tech, and (7) Any special storage instructions.  Check all for visible signs of degradation.

# REAGENTS

4.1	Reagent Summary

| Reagents / Kits | Supplier & Catalog Number |
| --- | --- |
| Total Prostate Specific Antigen | TPSA Flex® reagent cartridge, Cat. No. K6451 |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | Total Prostate Specific Antigen |
| --- | --- |
| Storage | Store at 2-8 °C. |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges - Sealed wells on the instrument are stable for 30 days. - Opened wells 1 – 12 are stable for 5 days. |
| Preparation | All reagents are liquid and ready for use. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier& Catalog Number |
| --- | --- |
| PSA CAL | Siemens Dimension Vista®, Cat. No. KC602 |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation

| Calibrator | PSA CAL, Cat. No. KC602 |
| --- | --- |
| Storage | Store at 2-8 °C. |
| Stability | - Unopened calibrator is stable until expiration date stamped on the box. - Opened Calibrator: once the stopper of the vial is punctured, assigned values are stable for 7 days when stored on board the Dimension Vista System. Opened Calibrator: Once cap is removed, assigned values are stable for 30 days when recapped immediately after use and stored at 2-8° C. Do not use this vial on board the instrument. |
| Preparation | PSA CAL is ready for use. No preparation is required. |

  - Calibration Criteria and Procedure

| Criteria | Special Notations |
| --- | --- |
| Reference Material | PSA CAL |
| Assay Range | 0.010 – 100 ng/mL |
| Suggested Calibration Level | See Reagent Package Insert for lot specific assigned values in ng/mL [µg/L] |
| Frequency | - Every 30 days for any one lot - Calibration interval may be extended based on acceptable verification of calibration. - For each new lot of Flex reagent cartridges - After major maintenance or service, if indicated by quality control results - As indicated in laboratory quality control procedures - When required by government regulations |
| Calibration Scheme | 6 levels, n = 3 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | Proceed with analysis |
| If result falls outside assay-specific specification, or QC values are out of Acceptable limits, | Troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Quality Control | Supplier & Catalog Number |
| --- | --- |
| LiquichekTM Immunoassay Plus Control Levels 1, 2 and 3 | Bio-Rad Laboratories Cat # 361, 362 and 363 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Liquichek Immunoassay Plus Controls, Level 1, 2 and 3 |
| --- | --- |
| Storage / Stability | - This product will be stable until the expiration date when stored unopened at -20 to -70 C. - Thawed and unopened: all analytes will be stable for 30 days. - Thawed and opened: all analytes will be stable for 14 days when stored tightly capped at 2-8 C. |
| Preparation | Before sampling, allow the control to reach room temperature (22-28 C) and swirl gently to ensure homogeneity. Promptly replace the stopper and return to 2-8C storage after each use. |

6.3	Frequency

| QC Frequency |  |
| --- | --- |
| 1 | - At least once each day of use, analyze two levels of a Quality Control (QC) material with known total prostate specific antigen concentrations. - Follow your laboratory internal QC procedures if the results obtained are outside acceptable limits. |

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| 1 | Acceptable ranges for QC are posted near each instrument and are programmed into the instrument software system. |  |
| 2 | Rejection criteria: For all rejection criteria see the Quality Control Rules and Problem Resolution SOP. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the quality control rules. |  |
| 3 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. When there is a failed Quality Control run any patient samples tested after the last successful Quality Control run need to be assessed according to Quality Control Rules and Problem Resolution SOP. - Corrective action documentation must include the following: rule, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated prior to reporting of any patient results. Courses of action might include but are not limited to: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| 4 | All Quality Control run on each shift must be attached to the run summary sheet. The shift lead technologist will review and document all corrected reports on corrective report log. |  |
| 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CVs are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SDs and CVs are sent to the Quality Control material manufacturer for peer group review. |  |

NOTE:  The laboratory director or designee may override rejection of partial or complete runs.  Justification for the override must be documented in detail.

6.5	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.6	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

Consult the Laboratory QC Program for complete details

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

TPSA Flex® reagent cartridge, Cat. No. K6451 is required to perform this test.

Total Prostate Specific Antigen Test is performed on the Dimension Vista® System after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:  For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

NOTE:   In the event that the test system becomes inoperable, notify supervision or designee for further direction.  Patient specimens must be stored in a manner that maintains the integrity of the specimen.

# CALCULATIONS

The instrument calculates the concentration of total prostate specific antigen in ng/mL [µg/L].

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.

10.3	Units of Measure

ng/mL [µg/L]

10.4	Clinically Reportable Range (CRR)

TPSA   0.010 – 100 ng/mL

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

Values that fall within the AMR or CRR may be reported without repeat.   Values that fall outside these ranges must be repeated.

| IF the result is … | THEN… |
| --- | --- |
| ≤ 0.010 ng/mL | Assure there is sufficient sample devoid of bubbles, cellular debris, and/or fibrin clots. Report as: < 0.010 ng/mL |
| ≥ 100 ng/mL | On Board Automated Dilution: Results ≥ 100 ng/mL will automatically have repeat testing performed into the instrument using dilution factor of 20. Do not release the results until the dilution is completed. |
| ≥ 2,000 ng/mL | If the recommended dilution does not give results within the clinically reportable range, report as: ≥ 2,000 ng/mL” |

.

# EXPECTED VALUES

11.1	Reference Ranges

0.00 – 4.00 ng/mL

11.2	Priority 1 & 2 Values

None established.

11.3	Priority 3 Values(s)

None established.

11.4     Standard Required Callback Processing

| Inpatient | Technical laboratory personnel are to call all critical (panic) results on all inpatients 24 hours a day, seven days a week, as soon as they are resulted. Have the person receiving verbal results to repeat back those results and document the call in a call screen, to include the nurse’s or doctor first and last name. |
| --- | --- |
| Outpatient | The technical laboratory personnel are to notify the client service Department for all outpatient critical value and document on the PMSG screen on who they notify. |

# CLINICAL SIGNIFICANCE

Prostate specific antigen (PSA) is a serine protease of approximately 30000 Daltons produced by the epithelial cells of the prostate gland. The level of PSA in serum and other tissues is normally very low. In malignant prostate disease (prostatic adenocarcinoma) and in non-malignant disorders such as benign prostate hypertrophy (BPH) and prostatitis, the serum level of PSA may become elevated.  In serum PSA exists primarily as three forms: free and complexed with either α1-antichymotrypsin (ACT) or α2-macroglobulin. The PSA protein associated with α2-macroglobulin is encapsulated and unavailable for measurement by current immunoassay systems. The TPSA assay on the Dimension Vista® System measures both the free and the ACT bound components of serum PSA equally, a property known as equimolarity.

The specificity of PSA to prostate tissue makes it a significant marker as an aid to the early detection and management of prostate diseases. Prostate cancer is the most common type of cancer found in men in the United States and the second leading cause of male cancer mortality, accounting for more than 30000 deaths in 2008. Prior to the use of PSA for early detection of prostate cancer, the traditional method of digital rectal examination (DRE) detected considerably fewer tumors. The most sensitive method for early detection of prostate cancer uses both DRE and PSA. The American Cancer Society and the American Urological Association (AUA) recommend that early detection of prostate cancer should be offered to asymptomatic men 50 years of age or older with an estimated life expectancy of more than 10 years. An abnormal DRE and/or an elevated PSA may suggest the presence of prostate cancer; however a prostate biopsy is required for final diagnosis.

PSA testing is also accepted as an adjunctive test in the management of prostate cancer. Serum levels of PSA are most useful when sequential values are obtained and monitored over time. After complete removal of the prostate gland (radical prostatectomy), PSA levels should decline to a very low or non-detectable level. A rise of the serum PSA level in prostatectomy patients indicates residual prostate tissue; recurrence or metastasis of the disease. Serum PSA levels during radiation treatment should decline and remain at baseline while the patient is in remission.

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the operator of specific malfunctions. Any report slip containing such error messages should be held for follow-up.  Refer to you Dimension Vista Operator’s Guide.

The expected maximum observed standard deviations for repeatability (within-run precision) using n = 5 replicates at the following analyte concentrations are:

| TPSA concentration | Acceptable SD Maximum |
| --- | --- |
| 4 ng/mL [µg/L] | 0.22 ng/mL [µg/L] |
| 20 ng/mL [µg/L] | 1.21 ng/mL [µg/L] |

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0.010 – 100 ng/mL

14.2	Precision

| Material | Mean ng/mL | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Repeatability | Within-Lab |
| Serum Pool | 0.46 | 0.013 (2.93) | 0.24 (5.18) |
| Liquichek™ 1 | 2.33 | 0.03 (1.19) | 0.07 (3.10) |
| Liquimmune® 2 | 14.6 | 0.21 (1.46) | 0.34 (2.31) |
| Serum Pool | 76.7 | 1.48 (1.99) | 2.62 (3.53) |

14.3	Interfering Substances

Hemolysis, Icterus, and Lipemia (HIL) Interference:

The TPSA method was evaluated for interference according to CLSI/NCCLS EP7-A2 at a Total Prostate Specific Antigen concentration of approximately 4.0 ng/mL [µg/L]. Bias is the difference in the results between the control sample (without the interferent) and the test sample (contains the interferent) expressed in percent. Bias exceeding 10% is considered interference.

| Substance Tested |  |  | Substance Concentration [SI Units] |  | Bias* % |
| --- | --- | --- | --- | --- | --- |
| Hemoglobin (hemolysate) |  |  | 500 mg/dL [0.31 mmol/L] (monomer) |  | < 10 |
| Bilirubin (unconjugated) |  |  | 20 mg/dL [342 µmol/L] |  | < 10 |
| Bilirubin (conjugated) |  |  | 20 mg/dL [342 µmol/L] |  | < 10 |
| Lipemia (Intralipid®) |  |  | 3000 mg/dL [33.9 mmol/L] |  | < 10 |

Intralipid® is a registered trademark of Fresenius Kabi AG, Bad Homburg, Germany.

*Analyte results should not be corrected based on this bias.

Non Interfering Substances

The following substances do not interfere with the TPSA method when present in serum and plasma at the concentrations indicated.  Inaccuracies (biases) due to these substances are less than 10% at total prostate specific antigen concentration of approximately 5 ng/mL [µg/L].

| Substance | Test Concentration | SI Units |
| --- | --- | --- |
| Acetaminophen | 20.0 mg/dL | 1324 µmol/L |
| Amikacin | 8 mg/dL | 137 µmol/L |
| Aminoglutethimide | 6 ng/mL | 25.8 nmol/L |
| Ampicillin | 5.3 mg/dL | 152 µmol/L |
| Ascorbic Acid | 6 mg/dL | 342 µmol/L |
| Biotin | 100 ng/mL | 409 nmol/L |
| Caffeine | 6 mg/dL | 308 µmol/L |
| Carbamazepine | 3 mg/dL | 127 µmol/L |
| Chloramphenicol | 5 mg/dL | 155 µmol/L |
| Chlordiazepoxide | 1 mg/dL | 33.3 µmol/L |
| Chlorpromazine | 0.2 mg/dL | 6.27 µmol/L |
| Cholesterol | 500 mg/dL | 12.9 mmol/L |
| Cimetidine | 10 mg/dL | 396 µmol/L |
| Ciprofloxacin | 46 µg/mL | 138 µmol/L |
| Creatinine | 30 mg/dL | 2652 µmol/L |
| Cyclophosphamide | 25 mg/dL | 957 µmol/L |
| Dextran 40 | 6000 mg/dL | 1500 µmol/L |
| Diazepam | 0.5 mg/dL | 18 µmol/L |
| Diethylstilbestrol | 0.02 mg/dL | 745 nmol/L |
| Digoxin | 6.1 ng/mL | 7.8 nmol/L |
| Doxorubicin-HCL | 7 mg/dL | 121 µmol/L |
| Erythromycin | 6 mg/dL | 81.6 µmol/L |
| Estramustine phosphate | 20 mg/dL | 372 µmol/L |
| Ethanol | 400 mg/dL | 86.8 mmol/L |
| Ethosuximide | 25 mg/dL | 1770 µmol/L |
| Finasteride | 0.2 mg/dL | 5.37 µmol/L |
| Fluoxetine HCL | 0.55 µg/mL | 1.59 µmol/L |
| Flutamide | 1 mg/dL | 36.2 µmol/L |
| Furosemide | 6 mg/dL | 181 µmol/L |
| Gentamicin | 1 mg/dL | 21 µmol/L |
| Goserelin acetate | 0.01 mg/dL | 78.8 nmol/L |
| Heparin | 3 U/mL | 3000 U/L |
| Human plasma kallikrein | 100 µg/mL | 1.40 µmol/L |
| Ibuprofen | 50 mg/dL | 2425 µmol/L |
| Immunoglobulin G (IgG) | 5 g/dL | 50 g/L |
| Ketoconazole | 7 mg/dL | 131 µmol/L |
| Leuprolide acetate | 10 mg/dL | 78.8 µmol/L |
| Lidocaine | 1.2 mg/dL | 51.2 µmol/L |
| Lithium | 2.2 mg/dL | 3.2 mmol/L |
| Megestrol acetate | 2 mg/dL | 52.0 µmol/L |
| Methotrexate | 300 mg/dL | 6.60 µmol/L |
| Nicotine | 0.1 mg/dL | 6.2 µmol/L |
| Penicillin G | 25 U/mL | 25000 U/L |
| Pentobarbital | 8 mg/dL | 354 µmol/L |
| Phenobarbital | 10 mg/dL | 431 µmol/L |
| Phenytoin | 5 mg/dL | 198 µmol/L |
| Prednisone | 1.65 µg/mL | 4.60 µmol/L |
| Primidone | 4 mg/dL | 183 µmol/L |
| Propoxyphene | 0.16 mg/dL | 4.91 µmol/L |
| Protein: Albumin | 6 g/dL | 60 g/L |
| Protein: Total | 12 g/dL | 120 g/L |
| Salicylic Acid | 60 mg/dL | 4.34 mmol/L |
| Sildenafil | 0.2 mg/mL | 421 µmol/L |
| Sulfamethoxazole | 117 µg/mL | 462 µmol/L |
| Terazosin HCL | 1.45 mg/mL | 5.59 mmol/L |
| Theophylline | 4 mg/dL | 222 µmol/L |
| Urea | 500 mg/dL | 83.3 mmol/L |
| Uric Acid | 20 mg/dL | 1190 µmol/L |
| Valproic Acid | 50 mg/dL | 3467 µmol/L |
| Vancomycin | 10 mg/dL | 69 µmol/L |

14.4	Clinical Sensitivity/Specificity/Predictive Values

Analytical Sensitivity:

0.008 ng/mL [µg/L]

The analytical sensitivity represents the lowest concentration of TPSA that can be distinguished from zero. This sensitivity is defined as the mean value (n = 20) plus two standard deviations of the level A (0 ng/mL [µg/L]) PSA Calibrator.

Functional Sensitivity:

≤ 0.033 ng/mL [µg/L]

The Functional sensitivity represents the lowest concentration with an observed 20% coefficient of variation.

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management  Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Therapeutic Drug Monitoring (Chemistry SOP)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert TPSA Flex® Reagent Cartridge K6451

# REFERENCES

  - Package Insert, TPSA Flex® Reagent Cartridge K6451, Siemens Healthcare Diagnostics Inc., 09/22/2010.

  - Package Insert, PSA CAL  KC602, Siemens Healthcare Diagnostics Inc., 01/2014.

  - Package Insert, Liquichek Immunoassay Plus Control, Bio-Rad Laboratories, 11/2013.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 0 | 1-13-15 | 6.4 | Updated Tolerance Limits | David Rajamonickam | Wendell McMillan |
| 0 | 1-13-15 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |

# ADDENDA

None