Technical SOP

| Title | D-Dimer (LIATEST) By STA Compact |  |
| --- | --- | --- |
| Prepared by | Harvey Vandenburg | Date: 04/21/2006 |
| Owner | David Rajamonickam | Date: 02/02/2015 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information…………………………………………………………………………….	2

2.	Analytical Principle………………………………………………………………………...	3

3.	Specimen Requirements……………………………………………………………………	3

4.	Reagets……………………………………………………………………………………..	4

5.	Calibrators/Standards………………………………………………………………………	5

6.	Quality Control……………………………………………………………………………..	6

7.	Equipment And Supplies…………………………………………………………………...	8

8.	Procedure…………………………………………………………………………………...	9

9.	Calculations…………………………………………………………………………………	10

10.	Reporting Results And Repeat Criteria……………………………………………………..	10

11.	Expected Values…………………………………………………………………………….	11

12.	Clinical Significance………………………………………………………………………..	11

13.	Procedure Notes…………………………………………………………………………….	12

14.	Limitations Of Method……………………………………………………………………..	12

15.	Safety……………………………………………………………………………………….	13

16.	Related Documents…………………………………………………………………………	13

17.	References………………………………………………………………………………….	13

18.	Revision History……………………………………………………………………………	14

19.	Addenda…………………………………………………………………………………….	14

# TEST INFORMATION

| Assay | Method/Instrument | Order Code |
| --- | --- | --- |
| D-Dimer | LIATEST STA Compact | DDQ |

Synonyms/Abbreviations

D-Dimer

Department

Providence Hospital Coagulation

# ANALYTICAL PRINCIPLE

The assay is based on the change in turbidity of a microparticle suspension that is measured by photometry.  A suspension of latex microparticles, coated by covalent bonding with monoclonal antibodies specific for d-Dimer, is mixed with the patient’s sample.  If d-Dimer is present, the resulting aggregates from antigen-antibody reaction increase the turbidity of the solution. The increased turbidity is measured as an increase in absorbance, which is proportional to the amount of d-Dimer in the patient’s sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None |
| Specimen Collection and/or Timing | If multiple tests are being drawn, the coagulation specimen must be done first unless blood cultures are being collected at the same time. Vacutainer tubes must be filled to the blue line to ensure the proper ratio of blood to anticoagulant. |
| Special Collection Procedures | N/A |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |  |
| --- | --- | --- | --- |
| Type -Preferred | Plasma (LB) |  |  |
| -Other Acceptable | None |  |  |
| Collection Container | Light blue top tube (3.2% sodium citrate) |  |  |
| Volume - Optimum | 4.5 mL Plasma |  |  |
| - Minimum | 3.0 mL Plasma |  |  |
| Transport Container and Temperature | Light blue top tube (3.2% sodium citrate). Room temperature. |  |  |
| Stability & Storage Requirements | Room Temperature: | Plasma: 8 hours. 20 ± 5 ºC |  |
|  | Refrigerated: | Do not refrigerate |  |
|  | Frozen: | Do not freeze |  |
| Timing Considerations | After collection, whole blood for coagulation studies should be transported to the laboratory and centrifuged ASAP. |  |  |
| Unacceptable Specimens & Actions to Take | Clotted: |  | Test not performed, request a recollection. |
|  | Gross hemolysis: |  | Test not performed, request a recollection. |
|  | QNS: |  | Test not performed, request a recollection. |
|  | Short draw: |  | Whole blood not drawn to the blue line. Test not performed, request a recollection. |
|  | Old specimen: |  | Whole blood received greater than 2 hours after collection. Test not performed, request a recollection. |
| Compromising Physical Characteristics | Hemolysis: Avoid hemolysis Lipemia: Acceptable Icterus: Acceptable |  |  |
| Other Considerations | None |  |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents / Kits | Supplier & Catalog Number |
| --- | --- |
| STA LIATEST D-DI | Diagnostica Stago- 00515 |

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| NERL reagent grade H2O | NERL Diagnostics 3650 | 6 X 480 mL |
| STA Desorb U | Diagnostica Stago- 995 | 12 per box |
| STA OWREN KOLLER | Diagnostica Stago- 00360 | 24 X 15 mL |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Assay Kit |  |
| --- | --- |
| Reagent 1 | Tris Buffer |
| Reagent 2 | Anti-D-Dimer coated latex microparticles |
| Container | Manufacturer supplied vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date on the vial Once opened, remains stable for 15 days on STA |
| Preparation | Allow Reagents 1 & 2 to stand at room temperature for 15 minutes before use. Mix the reagents by gently swirling the vials without creating any bubbles. Then place a STA-mini Reducer and the white perforated cap on each vial. |

| Reagent | STA Owren Koller |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | 2-25 degrees C |
| Stability | Unopened: Stable until expiration date on the vial After opening, stable for 72 hours on the STA compact. |
| Preparation | Ready to use. Do not install either a STA-reducer or a perforated cap on the solution bottle. |

| Reagent | STA Desorb U |
| --- | --- |
| Container | Manufacturer supplied vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date indicated on the box label |
| Preparation | Ready for use |

| Reagent | NERL reagent grade H2O |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | Room temperature |
| Stability | Stable for 30 days after opening |
| Preparation | Ready for use |

# CALIBRATORS/STANDARDS

# 5.1 Kit reagents are pre-calibrated; this pre-calibration is identical for all the reagents of each lot.

5.2 To enter the data for the calibration curve:  Scan the bar-code printed on the bar- code insert across the bar-code reader.

5.3 The calibration values for the lot of reagents being used will subsequently be validated after the Liatest D-Dimer controls have been run.  If the validation controls are outside of the assayed range, the instrument will not run patient samples.

5.4 To examine the calibration curve on the screen:  Through the Main menu under CALIB/CONTROL select CALIBRATION.  Move the cursor to D-Dimer and enter.

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| STA Liatest Control N + P | Diagnostica Stago- 00515 ( N + P kit) |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | STA Liatest Control N + P |
| --- | --- |
| Preparation | Reconstitute each vial with exactly 1.0 mL NERL reagent grade water. Allow the reconstituted to stand at room temperature for 30 minutes. Swirl gently. |
| Storage/Stability | The reagents in unopened vials are stable until the expiration date indicated on the box label, when stored at 2-8 degrees C. Once reconstituted, the controls are stable for 8 hours on the instrument. |

  - Frequency

    - Both control N and control P are run automatically when a new lot is calibrated.

    - Both control N and control ABN are run daily at the beginning of each 8 hour shift.

  - Tolerance Limits

    - The established QC ranges are lot specific, and are in the STA QC file.

    - The STA Compact monitors all control ranges automatically.  If any control exceeds + 2 SD, the instrument will audibly and visually alarm the operators. The 1-2s rule will be used for “rejection”.  If one QC result in the run exceeds 2 SD, run will be considered unacceptable.  In this case, the technologist must rerun the QC before running any patient sample

    - Rejected runs must be effectively addressed by corrective action.  Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed.

      - Troubleshooting QC.  If QC fails, take the following steps:

- Ensure that you are using the correct control lot number and the correct reagent lot number.

- Rerun the control using freshly prepared control material.

- If these steps fail to resolve the problem, attempt to run the assay on another analyzer.

- If QC is still unacceptable, immediately contact your supervisor.

    - Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, and the date and initials of the person recording the information.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

  - Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

  - Review of QC

●  Weekly and Monthly QC must be reviewed by the supervisor on weekly and monthly basis, if the SD’s or CV’s are greater than the established ranges, investigate the cause for the imprecision and document implementation of corrective actions.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

- Weekly and Monthly QC must be reviewed by the supervisor on weekly and monthly basis, if the SD’s or CV’s are greater than the established ranges, investigate the cause for the imprecision and document implementation of corrective actions.

- Monthly QC must be presented to Medical Director for review and signature.

  - Quality Assurance Program

- Monthly QC must be presented to Medical Director for review and signature.

- Each new lot number of reagent must be tested with external control materials and/or previously analyzed samples.  Performance of the new lot must be equivalent to the previous lot and must be document on the Parallel check Book

- All persons performing this assay must complete training and reviewed annually for competency, the result are documented on file.

- This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

STA Compact

7.2	Equipment

Mini-Reducer

Centrifuge

7.3	Supplies

Pipettes

Pipette tips

Cuvette Rolls

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | Perform Daily Maintenance: - From the Main Menu, press STATUS, then SYSTEM and examine the screen for the daily maintenance checks. - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. - Record Maintenance in the maintenance log as you perform it. - If any temperature check exceeds the tolerance limit, then check the peltier fluid, and the fluid circulation. If the fluid is low, replenish. If this does not solve the problem, call the STAGO service number on the instrument. |
| 2. | Perform Weekly, Monthly, or Four Month Maintenance (if required) - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. |
| 3. | Load cuvettes and cleaner/wash solution on the analyzer if needed. To load Desorb, Request the reagent drawer to open through the main menu under Loading. Scan the barcode of the reagent. Place the reagent into a position 40 or 41 in the product drawer. |
| 4. | For more details on Instrument set-up consult the operator’s manual |

| 8.2 | Specimen / Reagent Preparation |
| --- | --- |
| 1. | Prepare the D-Dimer reagents as stated in the reagent preparation section |
| 2. | Prepare controls as stated in the control preparation section |
| 3. | Load reagents. Request the reagent drawer to open through the main menu under Loading. Scan the barcode of the reagent. Place the reagents into any of two positions 29-35 in the product drawer. Scan the barcode of the Owren-Koller buffer. Press F1 for diluent. Load the buffer into the sample drawer. |
| 4. | Load controls. Place the controls into any of two positions 3-20 in the product drawer. |
| 5. | When all reagents and controls are loaded, press ESC. When the Quit box comes up, chose YES to exit the reagent-loading menu. The loading drawer will close. |
| 6. | Prepare samples by centrifuging citrated whole blood for 1500 G for 10 min or using (Stat Spin Express 4) at 5100/4100 RPM/ RCF for 5 min to produce a platelet poor sample. |
| 7. | For more details on specimen/reagent preparation, consult the operator’s manual |

| 8.3 | Test Run |
| --- | --- |
| 1. | Request quality control. Through Main Menu under CALIB/CONTROL, select QUALITY CONTROL and press enter. Place the cursor on D-Dimer. Select D-Dimer by pressing F1 and then F10 to run QC. |
| 2. | If QC is acceptable, load patient’s samples. Access the sample drawer through the Main Menu under LOADING. After the drawer opens, identify the type of specimen. If micro sample, press F8. If STAT, press F12. Identify the sample by bar coding or manually entering the patient identification number on the keyboard. Place the sample into the sample drawer. Move the cursor to the appropriate RECORDED PROFILES to select the profile needed. If a single test is needed, move the cursor to selection and select the test needed. |
| 3. | Once samples are loaded, press ESC, and then Enter on the Quit menu. The sample drawer will close. Once you exit out of the test menu, the instrument will immediately start processing samples. |
| 4. | The TEST STATUS screen will appear. If there is not enough reagents to run the test, the reagent will appear in read with the amount of deficiency. This deficiency will BLOCK the sample pipettor. Add the deficient reagents to run samples, per the directions above. Go to the SYSTEM menu, and cursor down to STOP SAMPLE PIPETTING YES. Press Enter to change to NO. Press ESC and the instrument will begin to run. |
| 5. | All patient results are displayed on the TEST PANEL screen and automatically print out and transmit to the LIS if selected. |

| 8.4 | Special Handling |
| --- | --- |
| 1. | For results that require operation intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will bring up the FILE PROCESSING screen. Follow the options in the left-hand corner of the screen. |
| 2. | If the patient’s results fall outside the assay reportable range (4.0 μg/mL), order a 1:5 dilution by selecting the LOADING menu, scanning the barcode on the patient’s sample, and then selecting D-DI 1:5. On the newer instrument, the dilution is performed automatically when the patient’s results exceed the AMR. |

# CALCULATIONS

None required. The instrument automatically corrects for dilutions.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

Not applicable

10.2	Rounding

No rounding is needed.  Results are reported to the nearest hundredth.

10.3	Units of Measure

μg/mL FEU (fibrinogen equivalent unit)

An FEU is defined as the quantity of fibrinogen initially present that leads to the observed level of D-Dimer.

10.4	Clinically Reportable Range (CRR)

0.22 – 20.0 μg/mL

  - Repeat Criteria and Resulting

| IF the D-Dimer result is … | THEN… |
| --- | --- |
| < V min | Report less than 0.22 μg/mL in the LIS. |
| > V max | Repeat using the D:DI 1:5 test. If the D-DI test result is >V max, enter results as “greater than 20.00 μg/mL”. If the D-DI test results gives a numerical answer <20.00 μg/mL, report the value in the LIS. |
| 0.22 μg/mL – 4.00 μg/mL | Report the value in the LIS |

# EXPECTED VALUES

11.1	Reference Ranges

0.0 - 0.5 μg/mL FEU.

11.2	Priority 1, 2 and 3 Limits

Not established.

# CLINICAL SIGNIFICANCE

Coagulation                                Plasminogen

Activation                                                          tissue Plasminogen activator

Fibrinogen                            Fibrin                                         Fibrin

Urokinase

Plasmin

D-Dimer

On the fibrin clot, plasmin degrades fibrin into various products, of which D-Dimer is one.

The D-Dimer level is increased in DIC. A normal D-Dimer level can rule out deep vein thrombosis and pulmonary embolism. In some thrombotic states, the D-Dimer level may increase after anticoagulant withdrawal, in the postoperative period, cancers, hemorrhages, and severe infections.

# PROCEDURE NOTES

.    FDA Status: Approved

.    Validated Test Modifications: None

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0.22-4.00 μg/mL

  - Precision

|  | Intra-Assay Reproducibility |  | Inter-Assay Reproducibility |  |
| --- | --- | --- | --- | --- |
|  | Sample 1 | Sample 2 | Sample 3 | Sample 4 |
| N | 21 | 21 | 10 | 10 |
| Mean μg/mL | 0.29 | 2.71 | 0.31 | 2.78 |
| SD μg/mL | 0.04 | 0.08 | 0.05 | 0.14 |

  - Interfering Substances

  - Samples suspected of being clotted should not be tested.

  - Cloudy samples may lead to a spuriously low D-Dimer level.  If the absorbance value at 540 nm of the plasma diluted 1:6 with Owren-Koller is < 0.35. If sample is too cloudy, try to obtain a fasting sample, or send the sample to Chantilly for testing.

  - FDP concentrations greater than 15 μg/mL may lead to an overestimation of the D-Dimer level.

  - RF at > 50 IU/mL may lead to an overestimation of the D-Dimer level.

  - The presence of anti-bovine albumin and/or anti-mouse antibodies in certain subjects may lead to an over estimation of D-Dimer level.

  - Clinical Sensitivity/Specificity/Predictive Values

The D-Dimer level increases during pregnancy, and also rises with age (>70 years).  The test is insensitive to fibrinogen and the E-fragment. Cross-reactivity is observed with the D-fragment. This has no effect because the presence of α2-antiplasmin precludes the production of FDP-D from fibrinogen.  The test is insensitive to unfractionated heparin (up to 2 IU/mL) and LMW heparin (up to 2 anti-Xa IU/mL).

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needle sticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- Warnings of other specific hazards are noted in this procedure.  Please comply with the requirements to reduce your risk of injury.”

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

Providence Hospital Laboratory Panic Value Standard Operating Procedure

STA Compact Operator’s Manual

Current Package Inserts

# REFERENCES

- Jacobs, D.S.; De Mott, W.R; and Oxley D.K.; Laboratory Test Handbook, 5th edition, Lexi-Comp Inc., Cleveland, OH. 2001.

- STA-LIATEST D-DI, Immuno-turbidimetric Assay of D-Dimer by STA Analyzers; Package Insert REF 00515, 2002.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 07/25/08 | Page 1 | Changed local to laboratory approval. Updated owner. Corrected site in header. | O Dualeh | R. Levy |
| 000 | 07/25/08 | 11.2 | Update section title to new terminology | O Dualeh | R. Levy |
| 000 | 07/25/08 | 13 | Added FDA Status and Validated test Modifications | O Dualeh | R. Levy |
| .001 | 01/19/10 | 19 | Updated Order Code Removed Current package insert from Addendum | O Dualeh | R.Levy |
| .002 | 10/14/10 | Page 1 | Update Owner | A. Foster | R.Levy |
|  |  | 8.0 | Updated Procedure to reflect removal of package inserts |  |  |
|  |  | 16.0 | Updated related documents to include current product insert |  |  |
| .003 | 12/12/2011 | Page 1 | Updated Owner | W. McMillan | R. Levy |
|  |  | Page 1 | Removed “Refer to the cover page for the Local Effective Date” |  |  |
|  |  | 1 | Updated Department to Coagulation |  |  |
|  |  | 3.1 | Removed Special Collection Procedures |  |  |
|  |  | 6.3.2 | Changed QC frequency to every 6 hours |  |  |
| .004 | 11/6/2012 | Page 1 | Update owner. | R. Layug | R. Levy |
|  |  | 6.3.2 | Changed QC frequency to every 8 hours |  |  |
| .005 | 02-02-2015 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| .005 | 02-02-2015 | 3.1 | Updated Patient Preparation | David Rajamonickam | Wendell McMillan |
| .005 | 02-02-2015 | 3.2 | Updated Specimen Type & Handling | David Rajamonickam | Wendell McMillan |
| .005 | 02-02-2015 | 8.2 | Updated Specimen / Reagent Preparation | David Rajamonickam | Wendell McMillan |

# ADDENDA

| Addendum | Title |
| --- | --- |