Technical SOP

| Title | Body Fluid Analysis |  |
| --- | --- | --- |
| Prepared by | Solomon Tsehay | Date: 06/25/2007 |
| Owner | David Rajamonickam | Date: 07/18/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	ANALYTICAL PRINCIPLE	3

3.	SPECIMEN REQUIREMENTS	3

4.	REAGENTS	4

5.	CALIBRATORS/STANDARDS	5

6.	QUALITY CONTROL	5

7.	EQUIPMENT and SUPPLIES	7

8.	PROCEDURE	8

9.	CALCULATIONS	12

10.	REPORTING RESULTS AND REPEAT CRITERIA	12

11.	EXPECTED VALUES	13

12.	CLINICAL SIGNIFICANCE	13

13.	PROCEDURE NOTES	15

14.	LIMITATIONS OF METHOD	15

15.	SAFETY	15

16.	RELATED DOCUMENTS	16

17.	REFERENCES	16

18.	REVISION HISTORY	16

19.	ADDENDA	16

# TEST INFORMATION

| Assay | Method/Instrument | Order Code | Local Code |
| --- | --- | --- | --- |
| Cell Count, Body Fluids | Hemacytometer, Microscopic/Manual |  | BFCC |
| Crystals, Analysis | Microscopic/ Manual |  | BFCR |

Synonyms/Abbreviations

Body fluids

Department

Hematology

# ANALYTICAL PRINCIPLE

The total nucleated cell count in body fluids is performed manually using a hemacytometer.  A differential cell count is performed via cytospin.  The color and appearance of the fluid is also reported.

Crystals are first observed microscopically with polarizing lenses, and if present, are identified with a first-order red compensator via birefringence.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | Not applicable |
| Specimen Collection and/or Timing | None defined |
| Special Collection Procedures | Fluid is collected in sterile vacuum bottle or other collection container and then aliquoted as needed into lavender top (EDTA), green top (Heparin), or red top tube (unpreserved). |
| Other | Not applicable |

3.2	Specimen Type & Handling

| Criteria |  |  |  |
| --- | --- | --- | --- |
| Type -Preferred | Site specified on collection 3 mL Fluid in EDTA for Count, Diff 1 mL Fluid (unpreserved) for Crystal |  |  |
| -Other Acceptable | 3 mL fluid in Heparin or Plastic Vial |  |  |
| Collection Container - Alternate | Lavender Top Tube Heparin (Green Top Tube) or Plastic Vial Plain red top or sterile container |  |  |
| Volume - Optimum | 3.0mL |  |  |
| - Minimum | 1.0mL If less than 1.0mL is received, perform testing in the following order: count, diff, crystal. & cancel those portions of the test that cannot be performed due to QNS. |  |  |
| Transport Container and Temperature | EDTA (Lavender) or Heparin (Green). Room Temperature |  |  |
| Stability & Storage Requirements | Room Temperature: |  | 48 hours |
|  | Refrigerated: |  | 48 hours |
|  | Frozen: |  | Unacceptable |
| Timing Considerations | Not applicable |  |  |
| Unacceptable Specimens & Actions to Take | Since Synovial fluid specimens are more difficult to obtain and are frequently “one time only” samples, the criteria for an acceptable sample should not be as rigid as those for peripheral blood. Samples with small clots should not be rejected, but the presence of clot should be noted. |  |  |
|  | Condition | Then |  |
|  | Small clots noted | Perform test, put a comment stating- Cell count may be inaccurate due to presence of clots in fluids. |  |
|  | Solid clot | Test not performed, delete in as: Test not performed. Specimens received clotted, please resubmit. Test has been cancelled. |  |
|  | Age of specimen | Since each specimen deteriorates at unpredictable ranges, aged specimens are to be tested and evaluated for significant deterioration of WBC. Only results deemed valid will be reported. All others should be deleted: Unsuitable for analysis due to the age of the specimen. Test has been cancelled. |  |
| Compromising Physical Characteristics | None defined |  |  |
| Other Considerations | None defined |  |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| Diluting Fluid, 0.2%(w/v)Crystal Violet | RICCA Chemical, Cat #7996-4 |

4.2	Reagent Preparation and Storage

NOTE:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

| Reagent | Diluting Fluid, 0.2%(w/v)Crystal Violet |
| --- | --- |
| Container | Glass or Plastic |
| Storage | Manufacturer’s expiration date. |
| Stability | Manufacturer’s expiration date. |
| Preparation | No preparation needed |

# CALIBRATORS/STANDARDS

Not applicable

# QUALITY CONTROL

6.1	Controls Used

| Control | Supplier & Catalog Number |
| --- | --- |
| Cell-Chex - 2x2mL each of Level 1&2 | Streck Laboratories, Inc. Cat. #212419 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

- Store upright at 2-8ºC.

- Closed-vial stability 60 days.

- Open-vial stability 30 days.

- Ready for use

  - Frequency

QC should be tested once a day.

|  | Minimum Number of QC Samples |
| --- | --- |
| 1. | Each Technologist, for each day of testing, must perform two levels of QC |
| 2. | Prior to reading differentials on patient specimens stained on the Hema-Tek slide stainer, a smear must be reviewed to verify that the staining is adequate for differential of the various cells. |

- Diluting fluid must be checked daily for contamination and documented.

6.4	Tolerance Limits

a)	Cell count by Manual Hemacytometer

QC values for Manual Hemacytometer are lot specific so check package insert.

b)	Differential %

QC values for Differential % are lot specific so check package insert.

- If both QC values are within 2 SD, patient results may be released.

- If a control value is >2SD, repeat the control before running patient samples.  If the repeat control is within 2SD, patient samples may be run.

- If the repeat of the control value is still >2SD, further investigation is required before running patient samples.

c)	Run Rejection Criteria:  Specify the QC Limits

- The QC procedure for cell count assay will employ a 2-SD QC rules. Runs where the values for any of the 2 SD values are violated will be rejected.

- Each time one control exceeds the above limits for rejection, the run is considered out of control (failed) and patient results must not be reported.  The testing personnel must employ the departmental remedial action protocol.

e)	Corrective Action:

- All rejected runs must be effectively addressed through corrective action. Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed.  Supervisors may override rejection of partial or complete runs only with detailed documentation and criteria for overrides that are approved by the Laboratory Director.

- Corrective action documentation must include the following:  The QC rule(s) (or specific QC criteria) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, and the date and initials of the person recording the information.

f)	Review of QC

- See the Corporate QC SOP entitled “QC Responsibilities” for more details.

- Upon weekly and monthly review of QC, if the SD’s or CV’s are greater than the above maximums, investigate the cause for the imprecision and document implementation of corrective actions.

  - Review Patient Data

Since only a few patient samples may be tested in one day, daily review for trends may not be applicable.

  - Documentation

Refer to local policies and procedures for QC documentation and to Quest Diagnostics records management program for record retention requirements.

  - Quality Assurance Program

Refer to corporate and local polices and procedures for other quality assurance activities applicable to this procedure.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Not applicable.

  - Equipment

| Equipment | Supplier |
| --- | --- |
| Cytospin 3 Centrifuge | Not Specified |
| Microscope | Not specified |
| Slide-Stainer Cytocentrifuge | Wescor |

  - Supplies

| Supply | Supplier & Catalog Number |
| --- | --- |
| Microhematocrit Tubes | Fisher Scientific - Cat. #22-315790 |
| Hemacytometer | Fisher Scientific |
| Serological Pipette | Fisher Scientific |
| Disposable Tubes | Fisher Scientific |
| Cover glass | Fisher Scientific, Cat. #12-544-10 |
| Microscope Slides | Fisher Scientific, Cat. #12-550-13 |
| Petri Dish | Fisher Scientific, Cat. #08-757-12 |
| Alcohol Swabs | Fisher Scientific, Cat. #06-669-62 |
| Filter Paper | Fisher Scientific, Cat. #11-996 |
| Applicator Sticks | Fisher Scientific, Cat. #01-340 |
| Cytopro | Wescor, Cat. #SS-113 |

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

8.1	Site: Result using the following translation keys

| IF | THEN | IF | THEN |
| --- | --- | --- | --- |
| Ankle | A | Not Given | NG |
| Elbow | E | Right Ankle | RA |
| Knee | K | Right Elbow | RE |
| Left Ankle | LA | Right Knee | RK |
| Left Elbow | LE | Unspecified | UN |
| Left Knee | LK |  |  |

  - Color:    Determine the color of the Synovial fluid and report using translation keys listed in the following table.

| IF | THEN | IF | THEN |
| --- | --- | --- | --- |
| Amber | A | Gray White | GW |
| Blue | BL | Light Yellow | LY |
| Brown | BR | Orange | O |
| Colorless | CS | Pale Yellow | PY |
| Dark Yellow | DY | Red | R |
| Green | GR | Straw | S |
|  |  | Yellow | Y |

8.3	Appearance: Determine the appearance of the synovial fluid and report using translation keys listed in the following table.

| IF | THEN | IF | THEN |
| --- | --- | --- | --- |
| Bloody | B | Clotted | CLT |
| Bloody, cloudy | BC | Hazy | HZ |
| Clear | C | Turbid | T |
| Cloudy | CD |  |  |

8.4	Concentration

This is procedure should be done to confirm all positive BFCCs from automated

testing.

| Step | Specimen Preparation |
| --- | --- |
| 1. | Inspect specimen to determine the appropriate dilution. |
| 2. | For bloody specimens, dilute specimen with lysing fluid, (dilution should be optimal for counting nucleated cells, as RBCs have been lysed). |

| Step | Undiluted Count |
| --- | --- |
| 1. | Clean a hemacytometer with an alcohol swab and then wipe dry. Place coverslip on hemacytometer |
| 2. | Fill a micro hematocrit tube or disposable pipette with well-mixed Synovial fluid. |
| 3. | Charge the two chambers of the hemacytometer by touching the tip of the capillary to the coverslip edge where it meets the chamber floor. The chamber will fill by capillary action if the hemacytometer and coverslip are clean. |
| 4. | If the hemacytometer is overcharged, clean and recharge the hemacytometer |
| 5. | Place the charged hemacytometer in a humidified Petri dish for 10 minutes to allow the cells to settle. |
| 6. | Place the hemacytometer on the microscope and count the total number of nucleated cells present on both sides. The sides should agree within 20%. |
| 7. | Calculate the total nucleated cells. See Section 9 for calculation formula. |
| 8. | All calculations should be recorded on worksheet. |

| Step | Diluted Specimen 1:10 |  |
| --- | --- | --- |
| 1. | Mix specimen well. Using a 100μL pipette, add 100μL of body fluid to 0.9mL of diluting fluid. Mix dilution well. |  |
| 2. | Charge a counting chamber (one pipette per side), using proper technique. |  |
| 3. | Place in a petri dish for about 10 minutes to let the cells settle. |  |
| 4. | Count nucleated cells. |  |
| 5. | IF | THEN |
|  | The four WBC squares are counted | Multiply the cell count by 25. |
|  | The decision is made to count a different number of squares | Calculate the total number of nucleated cells. See Section 9 for calculation formula. |
| 6. | Step 5 needs be repeated for both sides of the hemacytometer. The results should agree within 20%. |  |
| 7. | Show all calculations on worksheet. |  |

| Step | Diluted Specimen 1:20 |  |
| --- | --- | --- |
| 1. | Mix specimen well. Using a 50μL pipette, add 50μL of body fluid to 0.95mL of diluting fluid. Mix dilution well. |  |
| 2. | Charge a counting chamber (one pipette per side), using proper technique. |  |
| 3. | Place in a petri dish for about 10 minutes to let the cells settle. |  |
| 4. | Count nucleated cells. |  |
| 5. | IF | THEN |
|  | The four WBC squares are counted | Multiply the cell count by 50. |
|  | The decision is made to count a different number of squares | Calculate the total number of nucleated cells. See Section 9 for calculation formula. |
| 6. | Step 5 needs be repeated for both sides of the hemacytometer. The results should agree within 20%. |  |
| 7. | Show all calculations on worksheet. |  |

8.5	Cytospin

| Step | Cytospin |  |
| --- | --- | --- |
| 1. | Assemble sample chamber: With cytoclip slide clip in open position, fit glass microscope slide and disposable cytofunnel against the metal back. Bring spring clip up and fasten under the two retaining hooks. |  |
| 2. | Pipette Sample: 1) If the cell count is >100 and <300, place 2 drops of specimen plus 1 drop of albumin into a disposable cytofunnel and place into the Cytospin centrifuge. The albumin is used to make the cells adhere to the slide better before the staining procedure. 2) If the cell count is >300, make a pre-dilution according to the following chart and then place 2 drops of the diluted specimen plus 1 drop of albumin into a disposable cytofunnel and place into the Cytospin centrifuge. |  |
|  | Cells/mm³ | Dilution |
|  | 301-700 | 1:2 (5 drops fluid = 5 drops Isoton) |
|  | 701-1500 | 1:5 (2 drops fluid = 10 drops Isoton) |
|  | 1501-3000 | 1:10 (2 drops fluid = 20 drops Isoton) |
|  | >3000 | 1:20 (2 drops fluid = 40 drops Isoton) |
|  | Place Plastic cap on cytofunnel. |  |
| 3. | Centrifuge Sample: 1) Remove the sealed head from the Cytospin centrifuge and place on counter. 2) Release the button in the center of the sealed head by holding the lid with one hand and pulling the release button with the other. Remove lid. 3) Place assembly containing patient sample into rotor bowl and balance with cytofunnel marked as balance in rotor bowl. 4) Replace the lid and lock in place by pushing down on center button. 5) Return sealed head to centrifuge and close centrifuge lid. 6) Press start. Cytospin centrifuge is preprogrammed to spin samples for 10 minutes at 1000rpms. An audible beep will signal when centrifuge has complete spin cycle. 7) Open lid and remove sealed head. 8) Open unit and remove patient assembly. 9) Open cytoclip and remove cytofunnel being careful not to disturb cell button on slide. If needed, draw a circle around cell button with a sharpie to help in locating it after it air dries for minimum of 15 minutes. |  |
| 4. | Stain slide using Hema-Tek slide stainer. |  |

8.6	Differential Count

| 1 | Perform a 2-part differential cell count of 100 cells on a cytocentrifuged specimen using Wright’s stain. The nucleated cells are classified and reported as a percentage |
| --- | --- |
| 2 | Exam smears for the presence of immature or abnormal cells. Refer to a Pathologist if abnormal or immature cells are noted. |

8.7	Crystal Examination

| Step | Crystal Examination |
| --- | --- |
| 1. | Place a drop of well-mixed uncentrifuged fluid on a clean glass slide and cover slip. Examine the preparation using polarized light to detect monosodium urate or calcium pyrophosphate dihydrate crystals. |
| 2. | Swing out the tint plate from the optical path. (The view field becomes dark.) |
| 3. | Place the specimen on stage. Focus on the specimen. (The specimen appears bright on the dark view field.) |
| 4. | Return the tint plate into the optical path. (The background color changes to reddish purple.) |
| 5. | Of the needle-shaped crystals seen in the view field, find the one pointing in the vertical direction and check its color |
| 6. | Rotate the tint plate rotation lever from its utmost right position to its utmost left position (in clockwise direction); check the color change of the crystal. |
| 7. | The crystals can be identified by the changes in their colors. (Refer to the chart below.) |

# CALCULATIONS

Formula for Hemacytometer

| Cells Counted |  10  dilution |
| --- | --- |
| # of squares counted |  |

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required.

10.2	Rounding

Results for cell counts are rounded to whole numbers.

  - Units of Measure

| Parameter | Units |
| --- | --- |
| Total Nucleated Cell Count | Cells/mL |
| Differential Counts | % |

10.4	Clinical Reportable Range

None defined

10.5	Repeat Criteria and Resulting

Any duplicate counts not agreeing within 20%.

All fluids needing a pathology review are to be taken to the pathologist on call for Hematology.  All slides are to be accompanied by report to correlate any abnormal findings with other tests that have been performed on the patient.

# EXPECTED VALUES

11.1	Reference Ranges

| Parameter/Units of Measurement | Reference Range |
| --- | --- |
| Total Nucleated Cell Count/mcL | Note established |
| Crystal | None seen |

  - Priority 1 & 2 Limits

Refer to Critical Call Policy

11.3	Priority 3 Limit(s)

Refer to Critical Call Policy

# CLINICAL SIGNIFICANCE

12.1	Categorization of Arthritides or Joint Diseases

Except for the identification of crystals and culture for microorganisms, synovial fluid examination usually does not elicit a specific diagnosis.  However, examination of the following characteristics is often valuable in categorizing a joint disease and in facilitating the establishment of a diagnosis: volume, clarity, color, viscosity, mucin clot formation, spontaneous glucose, crystals, and microbiologic culture.

By evaluating these characteristics of the fluid, joint disorders can be separated into five disease groups:

| Disease Groups | Joint Disorders |
| --- | --- |
| Group I Non - inflammatory | Degenerative joint disease, Trauma, Osteochondritis dissecans, Osteochondromatosis, Neuropathic osteoarthropathy, Pigmented villonodular synovitis |
| Group II Inflammatory | Rheumatoid arthritis, Reiter’s syndrome, Alkylosing spondylitis, Rheumatic fever, System lupus erythematosis, Scleroderma, Arthritis with Chronic ulcerative colitis or Regional enteritis |
| Group III Infections | Bacterial, Fungal |
| Group IV Crystal - induced | Gout, Pseudogout |
| Group V Hemorrhage | Hemorrhagic diatheses including – Hemophilia, Trauma, Neuropathic osteoarthropathy |

| Synovial Fluid Test Results According to Group of Arthritides |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Test | Normal | Group I Noninflammatory | Group II Inflammatory | Group III Infectious | Group IV Crystal Induced | Group V Hemorrhagic |
| Clarity | Clear | Clear or Cloudy | Cloudy | Very Cloudy | Cloudy | Very Cloudy |
| Color | Yellow | Yellow | Yellow | Gray-white | Opalescent or colorless | Bloody |
| Mucin Clot | Firm | Firm to Friable | Friable | Friable | Friable | Friable |
| Leukocyte Count, per nm3 | <200 | 200-3,000 | 3,000 - >100,000 | 10,000 - >100,000 | 1,000 - 100,000 | >5,000 |
| % PMS (Segs) | <25 | <30 | >50 | >80 | >70 | >25 |
| Crystals | No | No | No | No | Yes | No |

12.2	Crystals are seldom seen except in arthritides Group IV. Urate crystals are seen in gout; calcium pyrophosphate crystals are seen in pseudogout; and corticosteroid crystals may be present following therapeutic intra-articular injection of steroid.  The presence of cholesterol crystals has been described in osteoarthritis, rheumatoid arthritis, and familial hypercholesterolemia.  Oxalate crystals will be seen if the Synovial fluid was collected in tubes containing oxalate anticoagulant.

12.3	Corticosteroid crystals are usually needle-shaped.  They can be present in leukocytes, and have varying birefringence patterns depending on the particular steroid preparation used for therapeutic injection.  Consequently, for correct interpretation of needle-shaped crystals, one must know whether a prior therapeutic injection has been given.  Cholesterol crystals appear as notched plates, are not present in leukocytes, and are strongly birefringent.

12.4	Additional Microscopic Findings:

The microscopic examination of Synovial fluid may show red cells, leukocytes, and crystal-bearing leukocytes, as previously described.  The presence of synoviocytes (Synovial lining cells) in the fluid is associated with pigmented villonodular synovitis, rheumatic fever and osteoarthritis.  Synovial cells are round and much larger than leukocytes.  Cartilage cells, when present in the Synovial fluid, are associated with traumatic arthritis, osteoarthritis, and pseudogout.  Cartilage cells are much larger than leukocytes and irregular in outline.  RA cells, also called ragocytes, are segmented neutrophils containing round inclusions in their cytoplasm.  These inclusions contain immunoglobulin and complement.  As the name implies, RA cells occur in rheumatoid arthritis, but are not specific for the diagnosis.  Wright-stained smears from patients with systemic lupus erythematosis (SLE) may show typical LE cells in the Synovial fluid.

# PROCEDURE NOTES

- FDA Status: Approved

- Validated Test Modifications: N/A

- Perform cell counts as soon as possible since cells deteriorate with time.

- If there is a clot, perform the count on available liquid and make notation in the report.  Counts on partially clotted samples may be affected depending whether or not cells are trapped in the clot.

- Low power scanning should be performed on smear to evaluate cell distribution and evaluate for presence of malignant cells.

- If crystal examination is ordered, perform this test first to help estimate the dilution needed for the cell count.

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range

None defined.

14.2	Precision

Not applicable.

  - Interfering Substances

- Contamination with birefringent talcum powder may interfere with crystal analysis.

- Use of powdered EDTA or oxalate to anticoagulate specimens may interfere with crystal analysis.

14.4	Clinical Sensitivity/Specificity/Predictive Values

None defined.

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

- QC Responsibilities

- Critical/Alert Value Policy and Procedure

- Corporate Safety Manual

- Local EHS Policy

# REFERENCES

1)	Glasser, Lewis M.D. (1980, November/December).  Reading the signs in the synovia.  Diagnostic Medicine, Body Fluids II.

2)	Todd-Sanford, Davidsohn, Clinical Diagnosis and Management by Laboratory Methods. 15th Edition, Vol.  I, W.B. Saunders Company, 1974, pp. 1254-1279.

3)	Kjeldsberg, Carl and Joseph Knight, Body Fluids, American Society of Clinical Pathologist, Chicago, 1982, pp. 48-55, 68.

4)	Strasinger, Susan King, Urinalysis and Body Fluids, Philadelphia, F.A. Davis Company: 1989, pp. 16 6-180.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 7/15/08 | Page 1 | Updated owner. | O Dualeh | R. Levy |
| .001 | 01/19/2010 | 4.1,4.2 | Updated body fluid diluting reagent | O Dualeh | R.Levy |
| .001 | 01/19/2010 | 4.2 | Removed Unopette® WBC for manual method | O Dualeh | R.Levy |
| .001 | 01/19/2010 | 1 | Updated SOP Title | O Dualeh | R.Levy |
| .001 | 01/19/2010 | 8.7 | Crystal Examination procedure | O Dualeh | R.Levy |
| .001 | 01/21/2010 | 8.6 | Updated Diff Procedure and Slide Stainer | O Dualeh | R.Levy |
| .001 | 01/21/2010 | 7.2 | Updated Slide Stainer | O Dualeh | R.Levy |
| .002 | 01/13/2011 | Page 1 | -Change Owner -Updated Lab Approval -Formatted Table of content | A. Foster | R. Levy |
|  |  | 8.4 | Added frequency of manual body fluid count |  |  |
| .003 | 08/02/2011 | 4.1 | Removed reagents no longer in use | A. Foster | R. Levy |
|  |  | 4.2 | Removed reagent preparation for reagents no longer in use |  |  |
|  |  | 7.3 | Updated Supplier information for microhematocrit tubes |  |  |
| 4 | 7/18/14 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 4 | 7/18/14 | Page 1 | Removed “Annual” from the review section | David Rajamonickam | Wendell McMillan |
| 5 | 6/4/15 | 6.3 | Updated frequency of QC | David Rajamonickam | Wendell McMillan |

# ADDENDA

N/A