Technical SOP

| Title | Quality Control For The Dimension Vista® System |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Prepared by | Ruben Layug |  |  | Date: 10/14/2012 |  |
| Owner | David Rajamonickam |  |  | Date: 03/30/2015 |  |
| Laboratory Approval |  | Local Effective Date: |  |  |  |
| Print Name and Title |  |  | Signature |  | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |  |  |  |
| Review |  |  |  |  |  |
| Print Name |  |  | Signature |  | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	PURPOSE & SCOPE	2

3.	RESPONSIBILITY	3

4.	REAGENTS	3

5.	CALIBRATORS/STANDARDS	3

6.	QUALITY CONTROL	3

7.	EQUIPMENT and SUPPLIES	5

8.	PROCEDURE	5

9.	CALCULATIONS	7

10.    REPORTING RESULTS AND REPEAT CRITERIA	7

11.    EXPECTED VALUES	7

12.    CLINICAL SIGNIFICANCE	7

13.    PROCEDURE NOTES	7

14.    LIMITATIONS OF METHOD	7

15.    SAFETY	7

16.    RELATED DOCUMENTS	8

17.    REFERENCES	8

18.    ADDENDA AND APPENDICES	8

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Quality Control For The Dimension Vista System | Dimension Vista® System | N/A |

Synonyms/Abbreviations

N/A

Department

Chemistry

# PURPOSE & SCOPE

To outline the quality control process for the Siemens Dimension instruments. This procedure applies to all Core Laboratory personnel working with the Siemens

Dimension Vista instruments.

# RESPONSIBILITY

Core Laboratory personnel are responsible for performing and complying with this procedure.  The Core Laboratory Supervisor is responsible for content and review of this procedure.

# REAGENTS

Not applicable

# CALIBRATORS/STANDARDS

Not applicable

# QUALITY CONTROL

6.1	Controls Used

N/A

6.2	Control Preparation and Storage

N/A

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| 1 | Rejection criteria: For all rejection criteria see the Quality Control Rules and Problem Resolution SOP. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the quality control rules. |  |
| 2 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. When there is a failed Quality Control run any patient samples tested after the last successful Quality Control run need to be assessed according to Quality Control Rules and Problem Resolution SOP. - Corrective action documentation must include the following: rule, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated prior to reporting of any patient results. Courses of action might include but are not limited to: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| 3 | All Quality Control run on each shift must be attached to the run summary sheet. The shift lead technologist will review and document all corrected reports on corrective report log. |  |
| 4 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CVs are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SDs and CVs are sent to the Quality Control material manufacturer for peer group review. |  |

6.5	Review Patient Data

Technologist must review each result with error messages. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot; utilize published TEA for acceptability criteria.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director or designee for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

- Centrifuge

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

  - Print QC vial labels

    - Click on ADVANCED from the Vista Home Screen.

    - Click on the QUALITY CONTROL icon.

    - Click on the blue bar and select ORDER QC from the dropdown menu

    - Select the desired QC product from the list.

    - On the right side of the screen, click on the box for the level(s) needed

6.	Click on PRINT BARCODE on the left side of the screen.  DO NOT click

SUBMIT unless you want to run the control at that time.

7.	Retrieve the printed barcode labels from the VISTA label printer.

  - Filling the QC vials

- Select a clean brown glass vial and a white pierceable cap for each control 	to    be prepared.

- Wrap the barcode label around the glass vial.

- Set the Eppendorf pipet to 1000 µL and install a clean tip

- Assure that the QC material is thawed, well-mixed, and within it’s open 	vial expiration period.

- Carefully pipet 2000 µL of control material into the corresponding labeled vial.  NOTE:  use 1000 µL if preparing the Diabetes Control.

- Cap the vial with the pierceable cap.

  - Loading vials onto the Vista

    - Select a Vista calibrator/control rack.

    - Place the vial(s) in the rack with the barcode facing out so it will be visible 	to the barcode reader.

    - Be sure to balance the vials in the rack if loading more than one vial.

    - Load the rack onto the analyzer.

  - Siemens Protein 3 control

    - Remove one vial and one pierceable cap from the box.

    - Rehydrate the lyophilized control with 1.0 mL deionized water.

    - Allow control to reconstitute for 30 minutes at room temperature 22-28 °C.

    - Invert gently to mix.

    - When ready to place on analyzer, remove rubber stopper cap and replace it

with a white pierceable cap.

    - Place vial in calibrator/control rack with barcode facing out and load.

  - MORE Diagnostics Immunosuppressive Drug Control

    - Remove vial from freezer and thaw at room temperature 22-28 °C.

    - If control is already thawed, allow it to come to room temperature.

    - DO NOT place QC material on rocker.  Invert gently to mix.

    - Print barcode labels for sample cups and label 3 Vista sample cups and

place them in a GRAY Vista sample rack.

    - Program QC as follows:

a.	Go to ADVANCED—QC—ORDER QC

b.	Select the CSA control from the QC list

c.	Check the appropriate box for each QC level to be run

d.	Click on the box for the sample rack ID

e.	Scan the sample rack barcode you have selected

f.	Click on SUBMIT on the left side of the QC order screen

g.	Pipet approximately 300 µL of QC material into the cups

h.	Place the rack on analyzer to run

  - Running QC in cups may be helpful in troubleshooting QC problems.  If the 	integrity of the on-board QC material is questionable, a quick way to check is to run

the control using a cup.  Follow these steps to load QC material in a cup:

    - Go to ADVANCED—QC—ORDER QC

    - Select the desired control from the list

      - Click on QC panel if more than one test is to be run on the QC

      - Click on single method if only one test is to be run

    - Select the tests to be run and the level needed

    - In the lower right corner, select CUPS rather than VIALS, which is the 	default for all QC except CSA

    - Scan the sample rack that the QC is to be run in

    - Click on SUBMIT

    - Pipet the QC material into the cups and load onto the instrument

# CALCULATIONS

Not applicable

# REPORTING RESULTS AND REPEAT CRITERIA

Not applicable

# EXPECTED VALUES

Not applicable

# CLINICAL SIGNIFICANCE

Not applicable

# PROCEDURE NOTES

Not applicable

# LIMITATIONS OF METHOD

Not applicable

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

# REFERENCES

Not applicable

# REVISION HISTORY

| Version | Date | Reason for Revision | Revised By | Approved By |
| --- | --- | --- | --- | --- |
| 000 | 03-30-15 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 000 | 03-30-15 | Updated SOP template | David Rajamonickam | Wendell McMillan |
| 000 | 03-30-15 | Updated the Tolerance Limits | David Rajamonickam | Wendell McMillan |

# ADDENDA AND APPENDICES

Not applicable