Technical SOP

| Title | Microalbumin by Dimension Vista® System |  |
| --- | --- | --- |
| Prepared by | Ruben Layug | Date: 9/20/2012 |
| Owner | David Rajamonickam | Date: 3/17/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	7

7.	Equipment And Supplies	10

8.	Procedure	11

9.	Calculations	12

10.	Reporting Results And Repeat Criteria	12

11.	Expected Values	12

12.	Clinical Significance	12

13.	Procedure Notes	13

14.	Limitations Of Method	13

15.	Safety	14

16.	Related Documents	15

17.	References	15

18.	Revision History	15

19.	Addenda	15

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Urine 24 hr Microalbumin | Dimension Vista® System | MA24 |

Synonyms/Abbreviations

MALB

Department

Chemistry

# ANALYTICAL PRINCIPLE

The Microalbumin (MALB) method is based on a particle-enhanced turbidimetric inhibition immunoassay (PETINIA) adapted to the Dimension® clinical chemistry system which allows direct quantitation of albumin in urine samples. The MALB Flex® reagent cartridge contains a particle reagent (PR) consisting of synthetic particles with human albumin bound to the surface. Aggregates of these particles are formed when a monoclonal antibody (Ab) to human albumin is introduced. Albumin (ALB) present in the sample competes with the particles for the antibody, thereby decreasing the rate of aggregation. Hence, the rate of aggregation is inversely proportional to the concentration of albumin in the sample. The rate of aggregation is measured using bichromatic turbidimetric reading at 340 and 700 nm.

PR + Ab + ALB (sample)	                            PR + Ab + ALB (sample)

(scatters light at 340 nm)

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | 24-hour urine: Do not include first morning specimen in 24-hour collection. |
| Special Collection Procedures | Samples should not be collected after exertion, in the presence of urinary tract infection, during acute illness, immediately after surgery or after an acute fluid load. Specimens should be collected without preservatives |
| Other | None |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Urine: 24 hour, unpreserved and well mixed |  |
| -Other Acceptable | None |  |
| Collection Container | Clean, unbreakable, and leak proof container |  |
| Volume - Optimum | 1.0 mL |  |
| - Minimum | 0.5 mL |  |
| Transport Container and Temperature | Leak-proof urine container Room temp |  |
| Stability & Storage Requirements | Room Temperature 22-28 °C: | 2 days |
|  | Refrigerated: 2-8˚C | 14 days |
|  | Frozen: -20C to -70C | Not Recommended |
| Timing Considerations | None |  |
| Random Urine | Stable at Room temperature: 22-28 °C for 2 days Stable at Refrigerated : 2-8˚C for 14 days |  |
| Unacceptable Specimens & Actions to Take | Leaking container, frozen specimens, specimens collected in Boric acid re-submit specimen. |  |
| Compromising Physical Characteristics | Specimens visibly contaminated with blood are not suitable for analysis of albumin concentration. |  |
| Other Considerations | All specimens should be at room temperature before analysis |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| Microalbumin | Siemens, Flex® cartridge, Cat. No. DF114 |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagent boxes upon receipt and enter in reagent log.  Each container is be labeled with  (1) substance name,  (2) lot number,  (3) expiration date, and (5) checked for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

Corrosive. Contains sodium hydroxide. Causes burns. In case of contact with eyes, rinse immediately with plenty of water and seek medical advice. In case of accident or if you feel unwell, seek medical advice immediately (show label where possible).

| Reagent | Microalbumin |
| --- | --- |
| Container | Reagent cartridge |
| Storage | Store at 2-8° C |
| Stability | - Reagent is stable until expiration date stamped on the reagent cartridges. - Sealed wells on the instrument are stable for 30 days. - Once wells 1-8 have been entered by the instrument, they are stable for 3 days. |
| Preparation | Reagents are supplied ready for use. No additional preparation is required. |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| Microalbumin Calibrator | Siemens Dimension®, Cat. No. DC114 |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container is labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | Microalbumin Calibrator |
| --- | --- |
| Preparation | - Remove vials from refrigerator and allow to stand at room temperature (22-28 °C) for 10 to 15 minutes. - Swirl vials gently for 30 seconds, then gently invert 10 times. - Let vials stand for 10 minutes, then gently invert 10 times. - Use immediately or refrigerate at 2-8°C for future use. |
| Storage/Stability | - Store at 2 – 8° C. - The unopened reagents are stable until the expiration date printed on the label. - Once opened assigned values are stable for 24 hours when vials are closed tightly and stored at 2–8°C. |

5.3	Calibration Parameter

| Assay Range | 0.5 – 680 mg/dL |
| --- | --- |
| Calibration Material | Microalbumin Calibrator, Cat. No. DC114 |
| Calibration Scheme | 5 levels, n = 2 |
| Units | mg/L |
| Typical Calibration Levels | 0, 12.5, 25.0, 50.0, 110.0 mg/L |
| Calibration Frequency | Every 30 days for any one lot. |
| A new calibration is required | • For each new lot of Flex® reagent cartridges • After major maintenance or service, if indicated by quality control results • As indicated in laboratory quality control procedures • When required by government regulations |
| Starting coefficients | C0 590 C1 -605 C2 -1.6 C3 37.4 C4 0.5 |

5.4	Calibration Procedure

Auto Calibration:

- Place the required calibrator vials in a carrier.  Make sure the barcode labels are          entirely visible through the slots.

- Place the carrier in the loading area.

- Position the carrier with the labels facing away from the user.

- Press the Load button.

- Automatic calibration requires that calibrators be on the instrument. As the time for processing approaches, the instrument reviews onboard inventory for the appropriate calibrators.

Manual Calibration:

- Verify that calibrators and reagents are in inventory on the instrument.

- Press System > Method Summary > Calibration.

- Select a method from the sidebar menu. Press the Order Calibration button on the screen.

- Verify that the information on the screen is correct. Verify that the calibrator lot is correct using the drop-down menu.

  - When calibrating using Vials press OK.

  - When calibrating using Cups, check the Use Cups box. This displays the rack and cup position fields. For additional cups use the positions in ascending order. Be sure to use the number of calibration levels and cups as specified in the method IFU. Scan the rack barcode and place calibrator cups in an adapter in position 1 on a rack. Press OK and load the rack on the instrument.

- The status field in the calibration screen changes sequentially to Awaiting Scheduling, Preparing Calibrators and Processing.

5.5	Tolerance Limits

| IF……… | THEN…… |
| --- | --- |
| If result fall within assay-specific specification, and QC values are within acceptable limits, | proceed with analysis |
| If result falls outside assay-specific specification, and QC values are out of Acceptable limits, | troubleshoot the assay and/or instrument and repeat calibration |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| LiquichekTM Unassayed Chemistry Control | Bio-Rad Laboratories: Cat: 691 Level 1 Cat: 692 Level 2 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | LiquichekTM Unassayed Chemistry Control, Levels 1 and 2 |
| --- | --- |
| Preparation | - Best if QC is thawed for 24 hours at 2-8C. - Alternatively, allow the frozen control to stand at room temperature (22-28 °C) until completely thawed. Do not use a warming device. - Swirl gently to ensure homogeneity. Do not use a mechanical mixing device. - Use IMMEDIATELY. After each use, promptly replace the stopper and return to 2-8ºC storage. |
| Storage/Stability | - Unopened controls are stable until the expiration date when stored at -20 to -70oC. Avoid a frost-free freezer. - Thawed controls are stable for 15 days when tightly capped and stored at 2-8C. DO NOT REFREEZE. - Discard the vial if there is evidence of microbial contamination or excessive turbidity. - Do not use after expiration date indicated on the container. |

6.3	Frequency

Analyze all levels of QC material after every calibration and each day of testing (notated on the QC frequency sheets posted on the instruments).

Refer to the Dimension Vista® QC Schedule in the Laboratory policy Quality Control Program and in the Dimension Vista® Quick Reference Guide.

6.4	Tolerance Limits

| Step | Action |  |
| --- | --- | --- |
| 1 | Acceptable ranges for QC are posted near each instrument and are programmed into the instrument software system. |  |
| 2 | Rejection criteria: For all rejection criteria see the Quality Control Rules and Problem Resolution SOP. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the quality control rules. |  |
| 3 | Corrective Action: - All rejected runs must be effectively addressed through corrective action. When there is a failed Quality Control run any patient samples tested after the last successful Quality Control run need to be assessed according to Quality Control Rules and Problem Resolution SOP. - Corrective action documentation must include the following: rule, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, the resolution, and the date and initials of the person recording the information. - All failures must be investigated prior to reporting of any patient results. Courses of action might include but are not limited to: - rerun the control - reconstitute a new control and rerun - change reagents - recalibrate instrument |  |
| 4 | All Quality Control run on each shift must be attached to the run summary sheet. The shift lead technologist will review and document all corrected reports on corrective report log. |  |
| 5 | Review of QC - Weekly and Monthly QC must be reviewed by the supervisor or designee on a weekly and monthly basis. If the SDs or CVs are greater than established ranges, investigate the cause for the imprecision and document implementation of corrective actions. - Monthly QC must be presented to Medical Director or designee for review and signature. - Monthly QC SDs and CVs are sent to the Quality Control material manufacturer for peer group review. |  |

6.5	Review Patient Data

Technologist must review each result print-out for error messages. Refer to the Dimension® system manual “Error messages” section for troubleshooting. Check for unusual patterns, trends, or distributions in patient results (such as an unusually high percentage of abnormal results). Resolve any problems noted before issuing patient reports.

6.6	Documentation

- QC tolerance limits are programmed into the instrument via the EasyLink. The instrument calculates cumulative mean, SD and CV and stores all information for easy retrieval.

- Quality control records are reviewed daily at the bench, weekly by the Group Lead or designee, and monthly by the Supervisor/Manager or designee.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot.

- Training must be successfully completed and documented prior to performing this test.

- This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Monthly QC must be presented to the Medical Director for review and signature.

- Monthly QC mean and SD are sent to Bio-Rad Laboratories for peer group comparison.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Dimension Vista® System

7.2	Equipment

- Refrigerator capable of sustaining 2–8C.

- Freezer capable of sustaining range not to exceed -20 to -70C.

7.3	Supplies

- Aliquot Plates

- System Fluids

- Assorted calibrated pipettes (MLA or equivalent) and disposable tips

# PROCEDURE

MALB® Flex® reagent cartridge Cat. No. DF114 is required to perform this test. Microalbumin is performed on the Dimension® clinical chemistry system after the method is calibrated (see Reference Material in Calibration section) and Quality Controls are acceptable.

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection (when appropriate to negate risk of splash) are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Sample Processing |
| --- | --- |
| 1. | A sample rack holding tubes or cups is placed on the rack input lane. |
| 2. | The sample shuttle moves the rack to the barcode reader which identifies the rack and samples to the system. |
| 3. | The rack moves into the sample server and to the rack positioner. |
| 4. | At the same time, aliquot plates move from the aliquot loader into position. |
| 5. | The aliquot probe aspirates the sample from the tubes or cups and dispenses it into the wells of the aliquot plates. |
| 6. | After each aspirate-dispense action, the probe is thoroughly rinsed inside and out to prevent sample carryover. |
| 7. | When sample aspiration is completed, the sample server moves the rack back to the sample shuttle, where it is placed on the output lane and can be removed by the operator. |

| 8.2 | Specimen Testing |
| --- | --- |
| 1. | For QC placement and frequency, refer to the Dimension Vista® QC Schedule in the Laboratory QC Program. |
| 2. | Follow the instructions, outlined in the Dimension Vista® Operator’s Manual |
| 3. | The instrument reporting system contains error messages to warn the user of specific malfunctions. Results followed by such error messages should be held for follow-up. Refer to the Dimension Vista® system manual “Error messages” section for troubleshooting. |
| 4. | Follow protocol in Section 10.5 “Repeat criteria and resulting” for samples with results above or below the Analytical Measurement Range (AMR). Investigate any failed delta result and repeat, if necessary. |
| 5. | Append the appropriate English text code qualifier messages to any samples requiring a comment regarding sample quality and/or any other pertinent factors. |

  - Test Conditions

| Sample Size | 17 µL |
| --- | --- |
| Particle Reagent Volume | 76 µL |
| Antibody Reagent Volume | 76 µL |
| Assay Buffer Volume | 288 µL |
| Temperature | 37 °C |
| Wavelengths | 340 and 700 nm |
| Type of Measurement | Turbidimetric rate |

# CALCULATIONS

- The instrument automatically calculates and prints the concentration of Microalbumin in mg/dL.

- MALB Creatinine Ratio in mg/g = (MALB mg/dL/Urine Creat mg/dL) x 1000

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

No rounding is necessary.

  - Units of Measure

Microalbumin: Mg/dL, mg/24h or µg/min

Microalbumin/Creatinine Ratio: mg/g

10.4	Clinically Reportable Range (CRR)

Microalbumin:  0.5 – 13,600 mg/dL

Microalbumin/Creatinine Ratio: <0.1 – >100.0 mg albumin /g Creatinine

  - Repeat Criteria and Resulting

All repeats must replicate the original result within the total allowable error (TEa) of the assay.  Refer to TEa listing for specific information.

| IF the result is … | THEN… |
| --- | --- |
| <0.5 mg/dL | Check sample for bubbles, cellular debris. Review previous results. If Delta failure flags or other reason to suspect result then repeat test otherwise report as: <0.5 mg/dL |
| >34.0 mg/dL | On Board Automated Dilution: Samples with result is in excess of 34.0 mg/dL will automatically be diluted and repeated on the Dimension® system with the dilution factors of 20 and 400. Do not release the results until the dilution is completed. |
| >13,600 mg/dL | If the recommended dilution does not give results within the clinically reportable range, report as: “> 13,600 mg/dL” |

# EXPECTED VALUES

11.1	Reference Ranges

| Microalbumin |  |
| --- | --- |
| 24hours Urine | <30 mg/24h or <20µg/min |
| Microalbumin/Creatinine Ratio (Excretion Rate) | < 30 mg albumin/g Creatinine |

11.2	Critical Values

None established.

11.3	Priority 3 Limit(s)

None established.

# CLINICAL SIGNIFICANCE

Microalbuminuria is the term given to the condition of increased excretion of albumin in urine. High urinary Microalbumin is an early marker of renal disease. Measurements of Microalbumin levels in urine are used to predict the development of diabetic nephropathy. A measurement of Microalbumin is therefore considered a standard for the detection of diabetic complications.

Diabetic nephropathy occurs in 20-40% of patients with diabetes an is the single leading

cause of end-stage renal disease. A measurement of albumin in urine is useful in monitoring the development of incipient diabetic nephropathy and for the diagnosis and treatment of a variety of diseases characterized by albuminuria.

# PROCEDURE NOTES

- FDA Status:  FDA Approved/cleared

- Validated Test Modifications:  None

The instrument reporting system contains error messages to warn the user of specific malfunctions. MALB results followed by such error messages should be held for follow-up. Refer to your Dimension® Operator’s Guide.

A system malfunction may exist if the following 5-test precision is observed:

Concentration				      SD

12.5 mg/L				> 1.9 mg/L

50 mg/L				> 2.5 mg/L

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0.5 - 34 mg/dL

14.2	Precision

| Material Bio-Rad LiquichekTM Urine Control | Mean mg/L | Standard Deviation (%CV) |  |
| --- | --- | --- | --- |
|  |  | Within-run | Total |
| Level 1 | 10 | 0.65 (6.2) | 1.06 (10.2) |
| Level 2 | 68 | 0.69 (1.0) | 1.17 (1.7) |
| Urine Pool 1 | 20 | 0.64 (3.2) | 1.57 (7.8) |
| Urine Pool 2 | 50 | 0.65 (1.3) | 2.07 (4.2) |

14.3	Interfering Substances

The MALB method was evaluated for interference according to CLSI EP7-A29. Bias is the difference in the results between the control sample (without the interferent) and the test sample (contains the

interferent) expressed in percent. Bias exceeding 10 % is considered interference.

Substance Tested 		Substance 		Albumin 		Bias*

Concentration 	mg/dL [mg/L] 	%

Hemoglobin (hemolysate) 	115 mg/dL 		2.205 		 	-1

Bilirubin (unconjugated) 	2 mg/dL 		1.763 			-6

Bilirubin (conjugated) 	2 mg/dL 		1.763		 	+8

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not available.

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needle sticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- If you drive on company business (in a company car, a rental, or your own personal vehicle) you must comply with the requirements in the "Driver Safety Manual".  Please remember that driving requires your FULL attention and multi-tasking while driving (e.g., talking on a cell phone) is an unacceptable risk that you must not take.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Dimension Vista® Clinical Chemistry System Operator’s Manual

- Dimension Vista® Calibration/Verification Procedure

- Dimension Vista® Cal Accept Guidelines

- Dimension Vista® Calibration summary

- Sample Processing, Siemens Dimension Vista®  procedure

- Start up and Maintenance, Siemens Dimension Vista® procedure

- Laboratory Quality Control Program

- QC Schedule for Siemens Dimension Vista®

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Siemens Dimension Vista® Limits Chart

- Quest Diagnostics Records Management Procedure

- Dimension Vista® System Error Messages Chart

- Centrifuge Use, Maintenance and Functions Checks (Lab policy)

- Repeat Testing Requirement (Lab policy)

- Current Allowable Total Error Specifications at http://questnet1.qdx.com/Business_Groups/Medical/qc/docs/qc_bpt_tea.xls

- Current package insert URCA Flex® Reagent Cartridge K1077

# REFERENCES

  - Ghoshal, Amit K. and Soldin, Steven J., Evaluation of the Dade Behring Dimension® RxL: Integrated chemistry system-pediatric reference ranges. Clinica Chimica Acta 2003; 331:144.

  - Package Insert, URCA Flex® Reagent Cartridge K1077, Siemens Healthcare Diagnostics Inc., 05/23/2011.

  - Package Insert, CHEM I CAL, Siemens Healthcare Diagnostics Inc., 03/2010.

  - Package Insert, Unassayed Liquichek Chemistry Controls, Bio-Rad Laboratories, 10/2010.

  - Clinical guide to Laboratory Tests by Norbert W. Tietz 3rd edition, page 24.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 3-4-13 | 9 | Added formula for calculating Microalbumin/Creatinine Ratio. | R. Layug | R. Levy |
| 000 | 3-4-13 | 10.3 | Updated Units of measure. | R. Layug | R. Levy |
| 000 | 3-4-13 | 10.4 | Updated CRR. | R. Layug | R. Levy |
| 000 | 3-4-13 | 11.1 | Updated Reference range. | R. Layug | R. Levy |
| 000 | 3-4-13 | 14.3 | Updated interfering substances. | R. Layug | R. Levy |
| 1 | 3-17-14 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| 1 | 3-17-14 | 6.4 | Updated Tolerance Limits | David Rajamonickam | Wendell McMillan |
| 1 | 3-17-14 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |
| 2 | 4-15-15 | 6.4 | Updated Tolerance Limits | David Rajamonickam | Wendell McMillan |
| 2 | 4-15-15 | 10.5 | Updated Repeat Criteria & Resulting | David Rajamonickam | Wendell McMillan |
| 2 | 4-15-15 | 3.2 | Included random urine requirements. | David Rajamonickam | Wendell McMillan |
| 2 | 4-15-15 | 17 | Added reference to random urine | David Rajamonickam | Wendell McMillan |

# ADDENDA

None