Technical SOP

| Title | Activated Partial Thromboplastin Time (APTT) by STA Compact |  |
| --- | --- | --- |
| Prepared by | Omar Dualeh | Date: 09/15/2008 |
| Owner | David Rajamonickam | Date: 09/15/2014 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	5

6.	Quality Control	5

7.	Equipment And Supplies	8

8.	Procedure	8

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	11

13.	Procedure Notes	11

14.	Limitations Of Method	11

15.	Safety	12

16.	Related Documents	13

17.	References	13

18.	Revision History	13

19.	Addenda	14

# TEST INFORMATION

| Assay | Method/Instrument | Order Code |
| --- | --- | --- |
| Activated Partial Thromboplastin Time | STA Compact | APT |

Synonyms/Abbreviations

APTT

Department

Providence Hospital Coagulation

# ANALYTICAL PRINCIPLE

The APTT involves the recalcification of plasma in the presence of a standard amount of cephalin (platelet substitute) and a particulate activator (silica).  The time of clot formation is determined.  The instrument uses an electromagnetic clot detection system. The oscillation of a steel ball within the thromboplastin and plasma is monitored.  When the oscillation of the steel ball is stopped by clot formation, the sensor registers the time in seconds.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None |
| Specimen Collection and/or Timing | If multiple tests are being drawn, the coagulation specimen must be done first unless blood cultures are being collected at the same time. Vacutainer tubes must be filled to the blue line to ensure the proper ratio of blood to anticoagulant. |
| Special Collection Procedures | N/A |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |  |
| --- | --- | --- | --- |
| Type -Preferred | Plasma (LB) |  |  |
| -Other Acceptable | None |  |  |
| Collection Container | Light blue top tube (3.2% sodium citrate) |  |  |
| Volume - Optimum | 4.5 mL Plasma |  |  |
| - Minimum | 3.0 mL Plasma |  |  |
| Transport Container and Temperature | Light blue top tube (3.2% sodium citrate). Room temperature. |  |  |
| Stability & Storage Requirements | Room Temperature: | Plasma: 8 hours. 20 ± 5 ºC |  |
|  | Refrigerated: | Do not refrigerate |  |
|  | Frozen: | Do not freeze |  |
| Timing Considerations | After collection, whole blood for coagulation studies should be transported to the laboratory and centrifuged ASAP. |  |  |
| Unacceptable Specimens & Actions to Take | Clotted: |  | Test not performed, request a recollection. |
|  | Gross hemolysis: |  | Test not performed, request a recollection. |
|  | QNS: |  | Test not performed, request a recollection. |
|  | Short draw: |  | Whole blood not drawn to the blue line. Test not performed, request a recollection. |
|  | Old specimen: |  | Whole blood received greater than 2 hours after collection. Test not performed, request a recollection. |
| Compromising Physical Characteristics | Hemolysis: Avoid hemolysis Lipemia: Acceptable Icterus: Acceptable |  |  |
| Other Considerations | None |  |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| STA PTT A 5 | Diagnostica Stago- 00595 | 12 X 5mL |
| STA CaCl2 0.025M | Diagnostica Stago- 00367 | 24 X 15 mL |
| STA Desorb U | Diagnostica Stago- 00995 | 12 per box |
| NERL reagent grade H2O | NERL Diagnostics 3650 | 6 X 480 mL |

4.2	Reagent Preparation and Storage

NOTES:  Date, time and initial all reagents upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date and time, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | STA PTT A 5 |
| --- | --- |
| Container | Manufacturer Supplied Vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date on the vial Once reconstituted, remains stable for 24 hours on STA |
| Preparation | Reconstitute the vial with 5 mL of distilled water. Allow the reconstituted reagent to stand at room temperature for 30 minutes. Swirl the reconstituted reagent gently to obtain a homogenous suspension, and place white perforated cap on the vial. |

| Reagent | STA CaCl2 0.025M |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | 2-25 degrees C |
| Stability | Unopened: Stable until expiration date on the vial After opening, stable for 24 hours at 37 degrees C; and 72 hours on the STA compact. |
| Preparation | No reconstitution required. Allow the reagent to stand at room temperature for 30 minutes. Do not install either a STA-reducer or a perforated cap on the solution bottle. |
| Reagent | STA Desorb U |
| Container | Manufacturer supplied vial |
| Storage | 2-8 degrees C |
| Stability | Stable until expiration date indicated on the box label |
| Preparation | Ready for use |

| Reagent | NERL reagent grade H2O |
| --- | --- |
| Container | Manufacturer supplied bottle |
| Storage | Room temperature |
| Stability | Stable for 30 days after opening |
| Preparation | Ready for use |

# CALIBRATORS/STANDARDS

The instrument automatically calibrates new lot numbers of reagent.

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| STA Coag Control N + ABN | Diagnostica Stago- 00676 (N + ABN kit) |

6.2	Control Preparation and Storage

NOTE:    Date, Time and initial all controls upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date and time, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | STA Coag Control N + ABN |
| --- | --- |
| Preparation | Reconstitute each vial with 1.0 mL NERL reagent grade water. Allow the reconstituted to stand at room temperature for 30 minutes. Swirl gently. |
| Storage/Stability | The reagents in unopened vials are stable until the expiration date indicated on the box label, when stored at 2-8 degrees C. Once reconstituted, the controls are stable for 8 hours on the instrument. |

  - Frequency

Both control N and control ABN are run once every 8 hours, at the beginning of every shift.

  - Tolerance Limits

    - The established QC ranges are lot specific, and are in the STA QC file.

    - The STA Compact monitors all control ranges automatically.  If any control exceeds + 2 SD, the instrument will audibly and visually alarm the operators. The 1-2s rule will be used for “rejection”.  If one QC result in the run exceeds 2 SD, run will be considered unacceptable.  In this case, the technologist must rerun the QC before running any patient sample and document any potential patient impact.

    - Rejected runs must be effectively addressed by corrective action.  Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed.

      - Troubleshooting QC.  If QC fails, take the following steps:

- Ensure that you are using the correct control lot number and the correct reagent lot number.

- Rerun the control using freshly prepared control material.

- If these steps fail to resolve the problem, attempt to run the assay on another analyzer.

- If QC is still unacceptable, immediately contact your supervisor.

    - Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, and the date and initials of the person recording the information.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

  - Documentation

    - Control results are automatically filed in the STA Compact QC file. All results for a 24 hour period are reduced to a mean value, which is used in statistical analysis, and plotted on Levy-Jennings charts.

    - To print all QC data points, perform the following procedure prior to midnight. From the main menu under CALIB/CONTROl select QUALITY CONTROL.  Cursor to the APTT test and press enter to view the Levy-Jennings chart. Press F1 to view the results in tabular form.  Press F6 then enter to print the individual values under current controls.  Press ESC to exit back to the graph.  Press F2 or F3 to view other levels and continue with F1 to view the result list.

    - Refer to Record Retention SOP for record retention requirements, found in the Administrative Manual.

  - Review of QC

    - Weekly and Monthly QC must be reviewed by the supervisor on weekly and monthly basis, if the SD’s or CV’s are greater than the established ranges, investigate the cause for the imprecision and document implementation of corrective actions.

    - Monthly QC must be presented to Medical Director for review and signature.

6.8	Quality Assurance Program

6.8.1	Refer to corporate and local polices and procedures for other quality assurance activities applicable to this procedure.

6.8.2	Each new lot number of reagent must be tested with external control materials and/or previously analyzed samples.  Performance of the new lot must be equivalent to the previous lot and must be document in the parallel check book.

6.8.3	All persons performing this assay must complete training and reviewed annually for competency, the result are documented on file.

6.8.4	This procedure must be incorporated into the departmental competency assessment program.

6.8.5	The laboratory participates in CAP proficiency testing.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

STA Compact

7.2	Equipment

Centrifuge capable of producing platelet poor plasma

7.3	Supplies

Calibrated Pipettes

Pipette tips

Cuvette Rolls

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | Perform Daily Maintenance: - From the Main Menu, press STATUS, then SYSTEM and examine the screen for the daily maintenance checks. - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. - Record Maintenance in the maintenance log as you perform it. - If any temperature check exceeds the tolerance limit, then check the peltier fluid, and the fluid circulation. If the fluid is low, replenish. If this does not solve the problem, call the STAGO service number on the instrument. |
| 2. | Perform Weekly, Monthly, or Quarterly Maintenance (if required) - Refer to the STAGO #1 and #2 maintenance log and the operator’s manual for maintenance procedures. |
| 3. | Load cuvettes and cleaner/wash solution on the analyzer if needed. To load Desorb, request the reagent drawer to open through the main menu under Loading. Scan the barcode of the reagent. Scan the barcode of the reagent. Place the reagent into a position 40 or 41 in the product drawer |
| 4. | For more details on Instrument set-up consult the operator’s manual |

| 8.2 | Specimen / Reagent Preparation |
| --- | --- |
| 1. | Prepare the PTT reagent as stated in the reagent preparation section |
| 2. | Prepare controls as stated in the control preparation section |
| 3. | Load reagents. Request the reagent drawer to open through the main menu under Loading. Scan the barcode of the PTT reagent. Place the PTT reagent into any position 19-24 in the product drawer. Scan the barcode of the CaCl2 reagent Load the CaCl2 into position 39. |
| 4. | Load controls. Place the N control into well 1 and the ABN control into well 2. |
| 5. | When all reagents and controls are loaded, press ESC. When the Quit box comes up, chose YES to exit the reagent-loading menu. The loading drawer will close. |
| 6. | Prepare samples by centrifuging citrated whole blood for 1500 G for 10 min or using (Stat Spin Express 4) at 5100/4100 RPM/ RCF for 5 min to produce a platelet poor sample. |
| 7. | For more details on specimen/reagent preparation, consult the operator’s manual |

| 8.3 | Test Run |
| --- | --- |
| 1. | Request quality control. Through Main Menu under CALIB/CONTROL, select QUALITY CONTROL and press enter. Cursor to APTT, and select by pressing F1 and then F10 to run QC. |
| 2. | If QC is acceptable, load patient’s samples. Access the sample drawer through the Main Menu under LOADING. After the drawer opens, identify the type of specimen. If micro sample, press F8. If STAT, press F12. Identify the sample by bar coding or manually entering the patient identification number on the keyboard. Place the sample into the sample drawer. Move the cursor to the appropriate RECORDED PROFILES to select the profile needed. If a single test is needed, move the cursor to selection and select the test needed. |
| 3. | Once samples are loaded, press ESC, and then Enter on the Quit menu. The sample drawer will close. Once you exit out of the test menu, the instrument will immediately start processing samples. |
| 4. | The TEST STATUS screen will appear. If there is not enough reagents to run the test, the reagent will appear in read with the amount of deficiency. This deficiency will BLOCK the sample pipettor. Add the deficient reagents to run samples, per the directions above. Go to the SYSTEM menu, and cursor down to STOP SAMPLE PIPETTING YES. Press Enter to change to NO. Press ESC and the instrument will begin to run. |
| 5. | All patient results are displayed on the TEST PANEL screen and automatically print out and transmit to the LIS if selected. |

| 8.4 | Special Handling |
| --- | --- |
| 1. | For results that require operation intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will bring up the FILE PROCESSING screen. Follow the options in the left-hand corner of the screen. |

# CALCULATIONS

None.

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

Not applicable

10.2	Rounding

No rounding is needed.

10.3	Units of Measure

The APTT is reported in seconds, to the nearest tenth.

10.4	Clinically Reportable Range (CRR)

20-250 seconds

  - Repeat Criteria and Resulting

| IF the APTT result is … | THEN… |
| --- | --- |
| <20 seconds | Check the sample for clots, spin the sample again. Request a redraw. If the retest on the redrawn sample matches, then release the result as <20 and enter an HCOM that the result has been verified. |
| > 90 seconds and <250 seconds | Check the sample for clots, spin the sample again. If clotted request a redraw. If not, rerun. If the rerun matches within the instrument’s flagging limits (+ 5%), then release the original results. This is a critical result, and must be called following the Panic Value SOP. |
| >250 seconds | Check the sample for clots, spin the sample again. If clotted request a redraw. If not, rerun. If the rerun matches within the instrument’s flagging limits, then release the result as “>250 seconds”. This is a critical result, and must be called following the Panic Value SOP. |

# EXPECTED VALUES

11.1	Reference Ranges

21.3 – 35.8 seconds

11.2	Priority 1

> 90 seconds.

11.3	Priority 2 & 3 Results

Refer to the Critical Value Policy and Procedure.

# CLINICAL SIGNIFICANCE

The APTT is a basic coagulation screening test, useful in the assessment of congenital and or acquired deficiencies of the intrinsic coagulation pathway (factors I, II, V, VIII, IX, XI, and XII).  A prolonged PTT has been observed in the following clinical states:

Congenital Deficiencies:

- If the PT is normal, the following factors may be deficient:  VIII, IX, XI, and XII.

- If all of these factors are normal, a consider deficiency in prekalikrein (Fletcher factor) or HMW Kininogen (Fitzgerald factor).

Acquired Deficiencies and Abnormal Conditions

- Liver disease

- DIC

- Circulating anticoagulants, such as Lupus Anticoagulant

Therapeutic Monitoring:

- During heparin or oral anticoagulant therapy

- Treatment with thrombin inhibitors

- Note:  Platelet counts should be followed to ensure rapid diagnosis if heparin-induced thrombocytopenia develops.

# PROCEDURE NOTES

# FDA Status: Approved

Validated Test Modifications: None

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

20-250 seconds.

14.2	Precision

|  | Intra-Assay Reproducibility |  | Inter-Assay Reproducibility |  |
| --- | --- | --- | --- | --- |
|  | Sample 1 | Sample 2 | Sample 3 | Sample 4 |
| N | 21 | 21 | 10 | 10 |
| Mean | 30.8 | 52.7 | 31.6 | 54.7 |
| SD | 0.20 | 0.53 | 0.44 | 1.46 |
| CV | 0.7 | 1.0 | 1.4 | 2.7 |

14.3	Interfering Substances

- Samples suspected of being clotted should not be tested.

- Interferences when monitoring heparin therapy:

- When monitoring heparin therapy, any release of platelet factor 4 (PF4), which is a potent inhibitor of heparin, represents a major source of error.. To minimize the risk of contamination with PF4:

- Do not collect blood in glass, which might cause the release of PF4.

- Perform centrifugation within one hour.

- Lupus anticoagulants and deficiencies of certain factors (e.g., factor XII) may prolong the baseline PTT and/or accentuate the prolongation of the PTT when heparin is added. In these situations an alternative assay, such as the Heparin anti-Xa assay, should be used rather than the APTT to monitor heparin therapy.   For further information on laboratory tests used for monitoring heparin, low molecular weight, direct thrombin inhibitors and/or oral anticoagulant therapy, and the therapeutic range for the tests, consult the Quest Diagnostics Directory of Service, Coagulation Section, or contact QDNI's coagulation specialist Dr. Jeff Dlott at 703-802-7259.

- With very high doses of heparin, the APTT is unclottable, and therefore not useful.

14.4	Clinical Sensitivity/Specificity/Predictive Values

The APTT is statistically lengthened in young subjects, and shortened in older populations.   With single factor deficiencies, the deficient factor has to be below 15% to 45% before the APTT becomes prolonged. The APTT is more sensitive to intrinsic pathway deficiencies than to common pathway deficiencies.

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- Warnings of other specific hazards are noted in this procedure.  Please comply with the requirements to reduce your risk of injury."

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

Providence Hospital Laboratory Panic Value Standard Operating Procedure

STA Compact Operator’s Manual

Current Package Insert

# REFERENCES

- Jacobs, D.S.; De Mott, W.R; and Oxley D.K.; Laboratory Test Handbook, 5th edition,. Lexi-Comp Inc., Cleveland, OH. 2001.

- STA-PTT A 5, Determination of Activated Partial Thromboplastin Time (APTT) by STA Analyzers; Package Insert REF 0000595, 2004.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| .000 | 01/19/10 | 19 | Updated Addendum (deleted package insert) | O Dualeh | R.Levy |
| .001 | 10/14/10 | Page 1 | Update Owner | A Foster | R. Levy |
|  |  | 8.0 | Updated Procedure to reflect removal of package inserts |  |  |
|  |  | 16.0 | Updated related documents to include current product insert |  |  |
| .002 | 12/13/2011 | Page 1 | Updated Owner | W.McMillan | R. Levy |
|  |  | 1 | Changed Department to Providence Hospital Coagulation |  |  |
|  |  | 3.1 | Removed Special Collection Procedures |  |  |
|  |  | 3.2 | Removed specific requirements for centrifuge |  |  |
|  |  | 6.2 | Removed “exactly” from Control preparation directions |  |  |
|  |  | 6.3 | Changed frequency to every 6 hours |  |  |
|  |  | 6.4 | Added patient impact statement to 6.4.2 |  |  |
|  |  | 7.2 | Added need for centrifuge to produce Platelet Poor Plasma |  |  |
|  |  | 7.3 | Added need for calibrated pipettes |  |  |
|  |  | 8.1 | Changed “Four Month” to “Quarterly” |  |  |
|  |  | 8.2 | Removed specific instructions for centrifuge and added need for centrifuge to produce Platelet Poor Plasma |  |  |
|  |  | 14.3 | Changed numbering |  |  |
| .003 | 11/6/2012 | Page 1 | Updated owner. | R Layug | R Levy |
|  |  | 6.3 | Changed frequency to every 8 hours |  |  |
| .004 | 02-04-2015 | Page 1 | Updated Owner, BU Name in Header, Added PRO in SOP Number in footer and removed leading zeros to the Version | David Rajamonickam | Wendell McMillan |
| .004 | 02-04-2015 | 3.1 | Updated Patient Preparation | David Rajamonickam | Wendell McMillan |
| .004 | 02-04-2015 | 3.2 | Updated Specimen Type & Handling | David Rajamonickam | Wendell McMillan |
| .004 | 02-04-2015 | 10.5 | Updated Repeat Criteria and Resulting | David Rajamonickam | Wendell McMillan |
| .004 | 02-04-2015 | 12 | Updated Clinical Significance | David Rajamonickam | Wendell McMillan |
| 5 | 1/21/2016 | 4.2 | Updated Reagent preparation and storage | David Rajamonickam | Wendell McMillan |
| 5 | 1/21/2016 | 6.2 | Updated Control preparation and storage | David Rajamonickam | Wendell McMillan |
| 5 | 1/21/2016 |  | Added Addendum A and B | David Rajamonickam | Wendell McMillan |

# ADDENDA

| Addendum | Title |
| --- | --- |
| Addendum A | Instructions for Preparing Collection Tube for Hematocrit >55% |
| Addendum B | Phlebotomist Instructions for Blood Collection |

Addendum A

Instructions for Preparing Collection Tube for Hematocrit >55%

Explanation:

Polycythemia is a disease state in which the proportion of blood volume that is occupied by red blood cells increases - basically when Hematocrit (HCT) is greater than 55%.  It can cause prolonged coagulation results.

When a prolonged coagulation result is obtained, check the specimen for a clot first.

If the specimen is not clotted, be sure the specimen is not under-filled or over-filled, then check the HCT result.

If a HCT result of greater than 55% is obtained, immediately notify the doctor or attending nurse and ask for a redraw using a special tube prepared by the lab.

To prepare a special tube in the lab use the following instructions and formula:

The anticoagulant volume in the collection tube must be adjusted to obtain a 9:1 ratio of blood to Sodium Citrate.  Under or over-filling of the specially prepared collection tube is not acceptable.

The vacuum in the collection tube will be broken to adjust the volume of collection anticoagulant.  Because of this special collection technique, the stability for these whole blood specimens is reduced to four (4) hours after collection.

Formula to calculate the anticoagulant volume is:

Anticoagulant in mL = [(100 – HCT) / (595 – HCT)] x Volume of blood

Example 1:  Specimen with a 70% HCT in a 2.7 mL tube:

Patient with HCT of 70%

Using a 2.7 mL tube

Anticoagulant in mL = [(100 – 70) / (595 – 70)] x 2.7 = 0.15 mL or 150 uL

Pipette a 2.7 mL tube in a way to leave only 150 uL of anticoagulant in there.

A 2.7 mL tube contains 0.3mL anticoagulant; therefore remove 0.15mL

Example 2:  Specimen with a 70% HCT in a 1.8 mL tube:

Patient with HCT of 70%

Using a 1.8 mL tube

Anticoagulant in mL = [(100 – 70) / (595 – 70)] x 1.8 = 0.1 mL or 100 uL

Pipette a 1.8 mL tube in a way to leave only 100 uL of anticoagulant in there.

A 1.8 mL tube contains 0.2mL anticoagulant; therefore remove 0.1mL

Example 3:  Specimen with a 60% HCT in a 2.7 mL tube:

Patient with HCT of 60%

Using a 2.7 mL tube

Anticoagulant in mL = [(100 – 60) / (595 – 60)] x 2.7 = 0.2 mL or 200 uL

Pipette a 2.7 mL tube in a way to leave only 200 uL of anticoagulant in there.

A 2.7 mL tube contains 0.3mL anticoagulant; therefore remove 0.1mL

Example 4:  Specimen with a 60% HCT in a 1.8 mL tube:

Patient with HCT of 60%

Using a 1.8 mL tube

Anticoagulant in mL = [(100 – 60) / (595 – 60)] x 1.8 = 0.13 mL or 130 uL

Pipette a 1.8 mL tube in a way to leave only 130 uL of anticoagulant in there.

A 1.8 mL tube contains 0.2mL anticoagulant; therefore remove 0.07mL

Addendum B

Phlebotomist Instructions for Blood Collection

The technologist will prepare a special tube in which the anticoagulant has been adjusted, therefore the tube is not vacuumed.  The technologist will inform the phlebotomist of the exact amount of blood needed to fill the tube.

Equipment and Supplies

Latex gloves

Latex free tourniquet

Latex free Band Aid or Tape

Alcohol Prep (70% alcohol)

2x2 sterile gauze

Special Collection tube

Blood Collection Set 21 or 23 gauge winged set

Blood Transfer Device

3mL syringe

Biohazard bag

Biohazard sharps container

LIS collection list and label/Lab requisition

Collection Steps

- Introduce yourself to the patient by stating your first and last name.

- Positively identify the patient according to the SOP ‘Patient Identification’, Phlebotomy procedure manual.

- Wash hands. Apply gloves.

- Explain the procedure to the patient and obtain patient’s consent to draw blood.

- Collect equipment and correct technologist-provided collection tube.

- Assemble equipment and break needle and syringe seals in the presence of the patient.

- Apply tourniquet about midway between the elbow and the shoulder 3-4 inches above the venipuncture site).  Place patient’s arm in a downward position to prevent reflux of ‘backflow’ of blood from the tube into the venous system.  Ask the patient to close hand gently.

- Palpate/feel for vein locating a vein that will flow fast (reducing the possibility of the blood clotting).

- Clean the area for venipuncture with a 70% alcohol pad decontaminating the special collection site.

- Allow the area to air-dry completely.

- Assemble the 21 or 23 gauge winged set to the 3mL syringe.  Pull back the plunger to dispel all the air out of the syringe.

- With the bevel up, align the needle with the vein while holding the skin taut.  Insert the needle at a 15-30 degree angle with the skin.  Remove your hand from drawing the skin taut.  Grasp the syringe and draw back bringing the plunger tip to the exact amount of blood requested by the technologist.

- Release the tourniquet. Ask the patient to open hand.

- Place gauze above the puncture site and remove the needle while simultaneously applying pressure on the puncture sit.  Firmly activate needle safety shield, a click must be heard to ensure that the safety shield is secure.

- Remove 21 or 23 gauge winged set from syringe.

- Attach the blood-filled syringe to the Blood Transfer Device.

- Connect the Blood Transfer Device to the un-vacuumed tube, provided by the technologist, and slow and gently fill the collection tube.  DO NOT FORCE blood into tube.  Pressure can lead to tube explosion and blood exposure.

- Place the cap on the tube and invert a few times to make sure the anticoagulant is mixed with blood.

- Dispose of all blood collection equipment into the nearest sharps container.  DO NOT disassemble the syringe from the Blood Transfer Device.

- Dispose of all other used materials in appropriate container and wash hands.

- Label the sample with the LIS collection label and write the time, date, and your tech code.

Transport specimen to the Lab.